## Slide 1: PIX-based detectors
F.Murtas
INFN - Laboratori Nazionali di Frascati and CERN
Timepix as user friendly silicon detector
Installation of Timepixes and Timepix3 in Ua9-SPS at CERN
Instrumentation for extracted beams
Diamondpix : higly pixelated diamond detector
GEMpix : a gas chamber for hadron therapy quality controls
Future programs

## Slide 2: Hybrid Silicon Pixel Detectors
Standard CMOS can be used allowing on-pixel signal processing
Sensor material can be changed according to the application(Si, GaAs, CdTe..)
Notes: Mention signal generation Add X section?

## Slide 3: Timepix family
Timepix :
Time frame oriented
Counting mode or
Time of arrival or
Charge
Dead Time
> 5 keV
10 ns resolution
kB/s
Timepix3
Pixel oriented
Counting mode and
Time of arrival and
Charge
Continous measurement
> 3 keV
1.5 ns resoution
MB/s
High Temperature
Timepix2
Pixel oriented
Time frame oriented
?
?

## Slide 4: Timepix on ISS
Image of the astronaut Chris Cassidy working near the Timepix USB on the International Space Station (Courtesy of NASA, photo ref. no. iss036e006175)
University of Houston, IEAP Prague, NASA
5.5 mGy/d
0.3 mGy/d

## Slide 5: The UA9 SPS experimental setup 2017
Primary Beam
Collimator
Graphite 1m
Beam Pipe
Collimator
Graphite 1m
B
L
M
Absorber
W 60cm
2 Timepixes
BLM
BLM
BLM
BLM
BLM
BLM
RP1
QF 52010
QD 52110
QF 52210
QD 52310
RP1
RP3
RP3
Crystal
1 2
IHEP
GONIO1
IHEP
GONIO2
Crystal
3 4
CpFM
Absorber
W 100cm
1 new Timepix
BLM
3 Timepix installed inside the beam pipe in secondary vacuum Roman Pot
~ 43 m
~ 85 m
~ 50 m

## Slide 6: The UA9 SPS experimental setup 2018
Primary Beam
Collimator
Graphite 1m
Beam Pipe
Collimator
Graphite 1m
~ 43 m
~ 85 m
~ 44 m
B
L
M
Absorber
W 60cm
2 Timepixes
1 Timepix3
We doubled the number of Timepix detectors
A third Roman Pot was installed on
January 2018 with 2 Timepixes
A new Timepix3 has been installed in June
on Roman Pot 1 external side
BLM
BLM
BLM
BLM
BLM
BLM
RP1
QF 52010
QD 52110
QF 52210
QD 52310
RP1
RP3
RP3
2 Timepixes
RP0
RP0
Crystal
1 2
IHEP
GONIO1
Absorber
IHEP
GONIO2
Crystal
3 4
CpFM
Absorber
W 100cm
1 Timepixes
BLM
Katrine readout of Timepix3
(P. Burian)
~ 6 m

## Slide 7: The UA9 detectors : Timepix
Based on silicon sensor 300 mm thick
About 20 readouts per seconds
  200 ms with 20 ns resolution or
  2 ms with 200 ns resolution
Cluster reconstruction offline (was online in 2016)
Synchronized readout with SPS Radiofrequency
We have five assemblies installed with 14x14 mm2 active area (2 cm2)
2 Timpixes on Roman pot #0 (after the first Gonio crystals 1 and 2 image of the channeled beam)
1 Timepix on Roman pot #1 (after the second Gonio crystals 3 and 4 image of the double channeled beam)
2 Timpixes on Roman pot #3 (dispersive area for extracted beam image)
The detector is read in ToA (Time of Arrival) to measure the particle flux versus time
with a memory of 10 or 100 beam turns and Medipix mode for a high statistic images.
Rate of 10-20 fames per second with dead time (can be improved )

## Slide 8: The UA9 detectors : Timepix3
Based on silicon sensor 300 mm thick 14x14 mm2 active area (2 cm2)
Chips provided by Medipix Collaboration
Timepix3 board designed and built by University of West Boemia and IEAP in Prague
Also the holder was realized by the same group
designed also for the chip heat extraction
Data acquisition with time frame and continuous
  1.56 ns resolution
  charge and time at the same time
Possibility to read the beam without dead time
(Data Driven Mode)
holder
Timepix3 data acquisition system
Petr Burian (Technical Coordinator) Pavel Broulím (HW designer) Lukáš Pušman, (mechanics designer) Benedikt Ludwig Bergmann (data evaluation)

## Slide 9: Double Channelling seen by Timepix in UA9
Vertical
14mm
Horizontal 14mm
On October 17th double crystal experiment was performed :
a second crystal was inserted in the channeled beam
Horizontal 14mm
Time
Horizontal profile vs time
2nd crystal shadowduring insertion
2nd crystal shadow
on channeled beam
2nd crystal shadow
set for channeling
2nd channeled beam
Beam Characteristics :
Single bunch
Intensity 1.2 1011 p/bunch
Momentum 270 GeV/c
User LHCMD3 Qh =20.13 Qv= 20.18
Main
Beam
F.Murtas, A. Natochi, W.Scandale

## Slide 10: Online double channeling monitoring
F.Murtas, A. Natochi, W.Scandale

## Slide 11: Timepix telescope for crystal tests
We have used two quads and two timepix as telescope
for measurements with track reconstruction using the ToA mode.
The telescope was 100 meters from the crystals
The result are comparable with the CMS tracker used up to now
This is the image of main and the channeled beam
produced by the crystal on test.

## Slide 12: TimepixQuad detector for test on beam extraction
We have used two quads as beam monitor near the crystals
in counting mode for studies on beam extraction
subtraction
Without any object
on the beam
Crystal in VR
+ septum
Image subtraction
Chanelling
Volume
Reflection
F.Murtas, A. Natochi, W.Scandale

## Slide 13: Measurements on neutron beam (nTOF@CERN)
There is a big interest in the Neutron Spallation Source
for a good monitoring of beam position and intensity
The Timepix-quad has been recently
used for the beam alignment at nTOF faclity at CERN

## Slide 14
DIAMONDPix
Higly pixelated diamond detector

## Slide 15: Diamondpix : a radiation hardness detector
Recently two Diamondpix have been tested on
charged and neutron beam for future applications.
Pads on diamond
downward surface
Gold cathode on diamond
upward surface
Timepix3 Pads
Made by Element6 and Fraunhofer IZM

## Slide 16: Diamondpix : a radiation hardness detector
Details on :
Diamondpix: a CVD diamond detector with Timepix3 chip interface TNS 00581-2018-R1
Neutron / Gamma discrimination
Neutrons
Gammas
Beam image
G.Claps, F, Cordella, C.Digiulio, L. Foggetta

## Slide 17
GEMPix
Higly pixelated 3GEM detector

## Slide 18: GEMPix: triple-GEM + Timepix ASIC
3 GEMs:
Kapton foil with thin copper layers
Gas amplification in holes (large field)
Total gain of max 105
4 Timepix chips:
512 x 512 pixels
each 55µm x 55 µm
detection threshold per pixel of 1000 electrons

## Slide 19: GEMPix: triple-GEM + Timepix ASIC
Head-on
Side-on
Gas flux AR CO2
Triple GEM
Mylar window
Particles to be analysed
Gas flux
Triple GEM
Particles
to be analysed
Length analysed

## Slide 20: 3D Track Reconstruction (TOA)
3x3 cm2
Z coordinate (1 cm)
Z coordinate (1 cm)
Possibility to study single particles
→ lateral measurements
→ secondaries
Stuart P.George, F.Murtas, M. Silari

## Slide 21: Charge Measurement (TOT)
X-ray detection: 6 keV from 55Fe (1s frame)
20% energy resolution
20% energy resolution for 6 keV
Single 6 keV X-Ray
Stuart P.George, F.Murtas, M. Silari

## Slide 22: GEMPix in CNAO water phantom, 12C-ion beam
Beam
Water phantom
GEMPix
Background: ion chambers in a water phantom are commonly used for quality assurance in radiation therapy, good reliability but bad spatial resolution (state of the art: array of ICs with some mm spatial resolution)
GEMPix inserted in water phantom with 3D positioning system
3D energy deposition obtained (2D images at different depths)
J Leidner, F.Murtas, M. Silari

## Slide 23: Depth Scan: Bragg Curve
GEMPix measurement normalized to dose delivery system of CNAO (DDS), integral of the 2D images (previous slide) summed per depth position -> Bragg curve
In general: good agreement with FLUKA simulations and measurements with PTW peakfinder (commercial device to find Bragg peak position, based on ion chambers)
Our results: Underestimation in plateau, fluctuations in peak, underestimation in tail
Improvements: TP correction for GEMPix and new, integrated system
Zoom to Bragg peak
Beam
J Leidner, F.Murtas, M. Silari

## Slide 24: 2D beam images
X/Y axes: position on detector
Z axis (color): deposited energy
Beam spot taken on Plateau, Bragg Peak and Tail
Frame length: 20 ms / 100 ms (before / after the Bragg Peak)
Beam halo: single particle reconstruction
Plateau
Peak
Tail

## Slide 25: 3D energy deposition
Reconstructed 3D energy deposition (smoothed)
Unique imaging technology
Beam
Bragg Peak
Fragmentation Tail

## Slide 26: Summary
Timepixes (Timepix, Timepix3, Quads, Diamondpix, GEMPix )
show good performances in high intensity beam monitoring.
They have been installed successfully in different facility
(UA9, NToF @CERN and CNAO and BTF in Italy …. but also in ISS)
and they provide useful information on beam characteristics
and they become important for future project .
Improvement on acquisition software are fundamental
Waiting Timepix2 test and results on performances