## Slide 1
COMPASS experience with
HAMAMATSU MAPMTs
S. Dalla Torre

## Slide 2: references
P. Abbon et al., NIMA 587 (2008) 371-387
P. Abbon et al., NIMA 616 (2010) 21-37
[ P. Abbon et al., NIMA 631 (2011) 26-39 ]
Collaborators to the COMPASS RICH upgrade project from :
Alessandria, Bonn, Erlangen, Freiburg, Liberec, LIP, Mainz, Munich,
Prague, Saclay, Torino, Trieste
underlined: specific contribution to MAPMT sector
in colour: project leaders & characterisation

## Slide 3: OUTLINE
Introduction
MAPMTs : characterization – main chapter
MAPMTs : quality control – only a flavour
MAPMTs : yields and robustness

## Slide 4: OUTLINE
Introduction
MAPMTs : characterization – main chapter
MAPMTs : quality control – only a flavour
MAPMTs : yields and robustness

## Slide 5: COMPASS RICH-1
hadron PID from 3 to 60 GeV/c
acceptance: H: 500 mrad V: 400 mrad
trigger rates up to 20 KHz, beam rates up to 4∙107 Hz
radiator:
C4F10
5 m
6 m
3 m
MWPC’s + CsI
UV mirror
wall
Al vessel
radiator gas: C4F10
large gaseous RICH
two kind of photon detectors
detector designed in 1996
in operation since 2002
upgraded in 2006 (MAPMTs)
PMT’s
beam pipe

## Slide 6: UPGRADE WITH MAPMT
photons
MAPMT
concentrator
field lens
PRINCIPLE
Surface ratio: ~ 7

## Slide 7: UPGRADE WITH MAPMT
shielding against
B ≤ 200 G
home made voltage divider
IMPLEMENTATION

## Slide 8: UPGRADE WITH MAPMT
DETECTOR
ASSEMBLED
outside the RICH vessel
inside the RICH vessel

## Slide 9: AN INTERESTING ASPECT
image centroide
Theoretical
reference
He-Na laser
632 nm
optical system
under test
Lens focus
Detected
Image
(CCD camera)
Hartmann plate
for parallel light
1. the principle
2. the images
Zernike
polynomials
3. the analysis
tolerance
(= center points +/- 0.1 mm)
Deviations from
ideal surface
OPTICS QC: THE HARTMANN METHOD

## Slide 10: OPTICS QC: THE RESULTS
576 TELESCOPES:
A) ~70% within 50 µm tolerance
B) ~20% within 100 µm tolerance
C) ~10% within 150 µm tolerances
A
C
B

## Slide 11: OUTLINE
Introduction
MAPMTs : characterization – main chapter
MAPMTs : quality control – only a flavour
MAPMTs : yields and robustness

## Slide 12: WHICH MAPMT ?
R7600-03-M16 by Hamamatsu
16 anode channels
bi-alkali photo-cathode
UV extended glass window
Photo-cathode surface 18:1 18:1mm2
Multiplicative chain 12 dynodes
Gain 3.5 106 (typical)
> 1 106 at 800V
Anode DC per channel 0.8 nA (typical)
< 2 nA
Total DC 12 nA (typical),
60 nA (maximum)
Anode signal rise time 0.83 ns (typical)
Signal transit time 10.9 ns (typical)
Uniformity between anodes 1:2.5 (typical)
1:3 (maximum)
QE at 420nm >20%
QE at 250nm >5%

## Slide 13: 12 vs 11 dynodes 1/2
Typical mean gains (900 V)
First peak : 8 x 105
Second peak: 4 x 106
Population of the first peak:
30%-40%
contributions to the first peak:
Photoelectrons skipping the
first dynode
Conversions on the first dynode
900 V
Single photon mode, single anode
Already observed:
I. Chirikov-Zorin et al., NIMA 456 (2001) 310; I. Chirikov-Zorin et al., NIMA 461 (2001) 587.
BOTH PEAKS
GOOD FOR US !

## Slide 14: 12 vs 11 dynodes 2/2
Fractional integral
zoom
Crucial: the possibility to apply electronics thresholds low enough
to guarantee good single pe detection efficiency

## Slide 15: CROSS-TALK
10 fC threshold
cross-talk rate < 0.1 %
cross-talk rate
reference anode
WORNING:
the cross-talk rate dramatically
depends on the coupling
between anodes and FE

## Slide 16: EFFECTIVE PIXEL SIZE
Scanning with a light source
with a dimension of 0.1 mm r.m.s.

## Slide 17: MAPMT in moderate B
for each pixel
Shielding: B < 1 mT @MAPMT
All pixels
External
column
Internal
column
zoom
pe detection efficiency in B field
B = 0
B = 18 mT

## Slide 18: RATE CAPABILITY
mean signal amplitude versus rate/pixel, standard voltage divider scheme
pulsed light source synchronous to the trigger + random background from a lamp
IMPORTANT: MIP flux, if non negligible, can modify this picture
measured for single photoelectron
HV range in operation

## Slide 19: TIME RESPONSE
Folded with the time resolution
of the read-out chain ( 75 ps )
= 320 ps
The tail in the time distribution
is due to photons impinging
at the edges
a
0 1 2 3 4 5 6 t (ns)
at pixel centre
0.5 mm from pixel centre
0.5 mm from pixel border
at pixel border

## Slide 20: OUTLINE
Introduction
MAPMTs : characterization – main chapter
MAPMTs : quality control – only a flavour
MAPMTs : yields and robustness

## Slide 21: QUALITY CONTROL
measured @ 360 nm
PROTOCOL – 2 h / MAPMT
(largely automatized)
Visual Inspection (Pixel grid)
Noise rates and dark current
Scope Image stored
Signal shape, uniformity, gain
  Amplitude distribution @ several Voltages
Quantum efficiency
  measurement with blue LED and UV LED
Individual ADC
spectra for each
channel, voltage,
wavelenght
gain vs HV for each
MAPMT anode

## Slide 22: OUTLINE
Introduction
MAPMTs : characterization – main chapter
MAPMTs : quality control – only a flavour
MAPMTs : yields and robustness

## Slide 23: YIELDS & ROBUSTNESS
tested: 642 MAPMTs, i.e > 10 k anode channels
rejected:
  1 MAPMT dead
  1 MAPMT a few degree misalignment of the metallic dynode set
  20 MAPMTs too high dark current for 1 or 2 anodes
  (dark I < 2 nA/ch after ½ h in dark environment)
  rejection rate: ~ 2.5 %
after 5 years, running ~ 6 months/year
1 MAPMT dead over 576 in operation  less than 1‰ / y

## Slide 24: THANK YOU