## Slide 1: X-ray generator: an application of micro pattern gas detector
Toru Tamagawa (RIKEN, Japan)
4th Micro Pattern Gaseous Detector Conference (Trieste, Italy)
Kazuki Nishida(*) (TUS/RIKEN), Wataru Iwakiri (RIKEN),
Ryoichi Suzuki (AIST), Hidetoshi Kato (AIST), Shigetomo Shiki (AIST),
Takao Kitaguchi (Hiroshima Univ.), Yoko Takeuchi (TUS/RIKEN),
Teruaki Enoto (Kyoto Univ.), Asami Hayato (RIKEN)
(*) graduate student, heavily contribute to this work.
Who we are?
=> Original developer of SciEnergy GEM since 2003!
=> Applying them to astrophysical satellite mission.
Abnormal use of GEM for X-ray generator
Notes: 15分＋5分＝20分 スライドは全部で10枚程度かな。今はタイトルと目次を除いて12枚。少し多い。。。

## Slide 2
Outline
Introduction of "triggerable X-ray generator"
Fabricate and test the x-ray generator
Timing capability
A problem - breakdown
Possible industrial applications
Summary and Outlook

## Slide 3
Ordinal X-ray generators at lab.
Ordinary X-ray Tube
Pyroelectric X-ray Generator
High V created by heat with pyroelectric crystal
Small and handy
easy to operate
Intermittent operation
Low intensity, short life
Amptek COOL-X
electron source: filament (thermal)
High intensity
Robust in operation
Cooling system needed
= Large in size

## Slide 4
Purpose
Onboard calibration device for X-ray detectors in space.
Drift velocity measurement of GEM-TPC polarimeter (Kitaguchi, #54)
Continuous energy calibration of micro-calorimeter (compensate energy scale variation.)
Requirements
Small enough for flight
Triggerable (ns~ms timing)
Final goal
On-chip X-ray gen.
A part of circuit.
Triggerable X-ray generator
Electron
source
X-ray
electrons
target
10-100 kV
(1) HV ON/OFF
(1) Switching high voltage
Need specific electronics for swith
Expensive for high speed electronics robust against HV.
How to realize triggerable X-ray?
What we need?
(3)
(2) shutter
(2) Mechanical shutter
Limited speed (~1 ms)
Need large mechanics.

## Slide 5
Conceptual design
Selection of electron source
Thermal emission (filament) => NG
Field emission (FE) => promising
Combining CCNS with GEM
CCNS substrate
GEM
Coniferous Carbon Nano Structure (CCNS)
Very sharp at tip
Robust at root
Already used for
portable X-ray generator at AIST, Japan
Carbon nanotube (CNT)
Very sharp tip (good FE device)
FE threshold ~ 10 kV/cm
Connection between CNT and substrate is very fragile. easy to be destroyed by strong E.
Push
on
Moreover, unwanted part of CCNS is covered by GEM substrate (reduce electron absorption at mesh)
CCNS substrate
Easy to apply FE-threshold-field
100V per 100um thick => 10 kV/cm
Switching the electron source V~100 V
  Low voltage (<100V) switching make everything easy.

## Slide 6
Outline
Introduction
Fabricate and test the x-ray generator
Timing capability
A problem – breakdown
Possible industrial applications
Summary and Outlook

## Slide 7
GEM
6mm
In dia.
CCNS + GEM
Electron source: CCNS+GEM
Electron source
CCNS substrate
made at AIST, Japan
2mm
in dia.
*) Thanks Arai-san (Fujikura) for GEM.
300 um
x30
SEM images
magnified
x104
1um

## Slide 8
X-ray emission test
Test Setup
X-ray Spectrum
CCNS
GEM
Ti target
Be
window
X-ray
~100V
+10kV
electrons
Si det.
(XR-100)
X-ray emitted at 75 V (75um thick GEM)
Typical X-ray flux: 107~108 c/s/sr
The X-ray generator worked perfectly as we expected !
CCNS+
GEM
target
Ti Ka (4.5keV)
Ti Kb (4.9keV)
ADC (ch)
Count/ch
e
X-rays

## Slide 9
I-V curve of electron source
Current (A)
GEM voltage (V)
I-V plot of CCNS+GEM
Electron
tunneling
(FE)
CCNS
Vacuum
w/ E-field
E
0
Fermi
level
work
func f
potential
Simple model of FE
I/E2 (uA m2/V2)
2.2 2.3 2.4 2.5
1/E (10-7 m/V)
Faular-Nordenheim relation
g = 1.01x103
(cf. g=1000-2000 for CNT)
Our electron source obeys the low of field emission of a typical CNT. (efficient electron emissions)
E: external electric field
g: field amplification factor
(sharper tip has larger g.)

## Slide 10
Outline
Introduction
Fabricate and test the x-ray generator
Timing capability
A problem - breakdown
Possible industrial applications
Summary and Outlook

## Slide 11
Timing capability test setup
Gate driver (pulses)
Simple gate driver for switching GEM V
About 20ns rise/fall-time with 4700pF load. (cf. 10ns with zero-load)
Half-bridge
MOSFET
DC 100V
TTL Pulse
Pulsed
100V
Gate drive IC
IC
~100V
HI
LO
TTL
pulse
CCNS
GEM
Ti target
Be
window
X-ray
+10kV
electrons
Si det.
20ns
Rise-time
HV out (200V)

## Slide 12
Timing capability
Input pulse
X-rays
X-rays when pulse-on, no X-ray when pulse-off
100 usec
0.6 usec
500 usec
NOTE: The minimum width was limited by the gate driver IC. We can improve more.
Pulsed operation successful.
CCNS-GEM X-ray gen. working as a triggerable X-ray generator.

## Slide 13
Outline
Introduction
Fabricate and test the x-ray generator
Timing capability
A problem - breakdown
Possible applications
Summary and Outlook

## Slide 14
SEM image of broken GEM
One problem: breakdown
GEM voltage breakdown at higher applied voltage (>300V).
Resistance of GEM down to a few kW.
This drops the reliability of our X-ray generator.
Target current (uA)
Elapsed time (sec)
GEM V
Tripped
GEM 330 V
aging started
Aging ongoing
Harmless fragment
ramp-up

## Slide 15
SEM image of a smoking gun
Breakdown mechanism
Brand-new CCNS substrate contains carbon fragments in the surrounding area.
Coming off the fragments and attaching to the GEM cause breakdown.
X-ray energy (keV)
EDS spectrum
CKa
Cu L
Cu Ka
Cu Kb
Clear evidence of carbon fragments
GEM electrode
carbon fragment
carbon fragment
Covered by electrode
active area
We are testing how it works.
How to prevent?
Mask the fragments with electrode.

## Slide 16
Outline
Introduction
Fabricate and test the x-ray generator
Timing capability
A problem - breakdown
Possible industrial applications
Summary and Outlook

## Slide 17
Take snap shots coincident with rotation of turbine.
Functionality is needed not in imager but in X-ray generator.
Imaging fast rotating devices
Portable gas turbine
(c) IHI
400,000rpm=6.7kHz
400W generator
X-ray generator
X-ray imager
Possible industrial applications
Flat panel X-ray generator
GEM electrode matrix circuit: Any portion of generator can be turned on separately.
trigger
Flat panel X-ray gen
object
X imager
3D imaging with parallax
We are developing highly-functional
X-ray generators with MPGD technique.
Superpose images with displacement
to make focus on a layer.
No movable portion

## Slide 18
Summary and outlook
A new triggerable X-ray generator combining CCNS with GEM are fabricated.
It is demonstrated that the CCNS+GEM X-ray generator worked perfectly as we expected.
It is demonstrated that the X-ray generator had at least 600ns switching capability. (of course, can improve more.)
There are a lot of industrial applications. More ideas welcome!
Thank you for your attention!