## Slide 1: Laser Raman spectroscopy for the KATRIN experiment ASPERA Technology Forum “Mirrors and lasers in astroparticle physics infrastructures”
The KATRIN experiment
Laser Raman
Tritium resistant coatings

## Slide 2: Astroparticle physics
Connecting elementary particles and the universe
ν
  Neutrino

## Slide 3: Impact of the neutrino mass on cosmology
mn = 4 eV/c2
mn = 0 eV/c2
mn = 1 eV/c2
mn = 7 eV/c2
Neutrino mass is an important input parameter for cosmology
mν < 2.3 eV/c2 (Mainz)
  Even tiny neutrino mass has impact on energy content of the universe
  Simulation of evolution of universe
   Massive neutrinos smear out large scale structure
WMAP
mν = 0 eV/c2
mν = 0.95 eV/c2
mν = 1.9 eV/c2
Agarwal, Feldman 2010
1 eV/c2 ≙ 1.8 · 10-36 kg

## Slide 4: The KATRIN experiment
KArlsruhe TRItium Neutrino experiment
t1/2 = 12.3 years

## Slide 5: Main spectrometer transport
Donau
250 km
Notes: 5

## Slide 6: The KATRIN experiment
Determination of anti-neutrino mass with 200 meV/c2 sensitivity (90 % C.L.)
Tritium source
Transport section
Spectrometer
1 meV/c2 ≙ 1.8 · 10-33 kg
KArlsruhe TRItium Neutrino experiment
t1/2 = 12.3 years

## Slide 7: Inner Loop system of KATRIN
Buffer vessel
Laser Raman
cell
Pressure controlled buffer vessel
Permeator
Pump
Source
conductance L
Δp
T2
Waste
p = 100-200 mbar
Objectives of Inner Loop
Stable gas injection into source qin = L · Δp = 1.8 mbar l s-1
 Pressure stabilisation < 10-3
Expected gas composition:
Nonstop monitoring of tritium purity with 10-3 precision needed for > 60 days
> 91 % T2 < 9 % DT < 1 % D2 |  | Tritium purity εT > 95 %
Laser Raman system for monitoring of gas composition
Notes: 7

## Slide 8: Hydrogen isotopologues
T2
DT
HT
H2
D2
HD
6 hydrogen isotopologues
β-spectrum depends on gas composition
Gas composition needed for data analysis of KATRIN
3 hydrogen isotopes
Hydrogen
Deuterium
Tritium
Laser Raman spectroscopy

## Slide 9: The Raman Effect
Raman scattering
Photon loses energy to molecule  Change of wavelength  Excitation of molecule
Rotation
Vibration
Simulated spectrum (laser line 532 nm)
Q1 branch
Analysis
Line position  Qualitative analysis
Line intensity  Quantitative analsis

## Slide 10: Experimental setup
LARA cell
DPSS Nd:YAG laser (532nm, ≤ 5W)
CCD
Filter
Spectrograph
Optical fibre
3 LARA systems at TLK

## Slide 11: LARA, tritium loops and the appendix
Successful commissioning in 11/2009
Cell
Appendix
Glove box with tritium loops
Appendix: Connection between LARA and sample cell inside glovebox
LARA setup

## Slide 12: Long-term monitoring
Non-stop monitoring for > 21 days
Relative intensity Irel (%)
Time (days)
ptotal = 200 ± 0.3 mbar
Laser power: 5 W
Acquisition time: 250 s
Tritium purity > 95 %
Relative intensity = Absolute intensity normalized to sum of all absolute peak intensities
S. Fischer, et al., Fusion Sci. Technol., 60, 3 2011
10-3 precision reached  KATRIN requirements reached

## Slide 13: Damage of coatings
Damage of inner coatings of LARA cell after 3 months filled with ~200 mbar T2 gas
Before: Operation for years with < 13.7 mbar T2 without problem
Laser windows: 532nm anti-reflection coating
Raman windows: Broadband anti-reflection coating
Dielectric coatings, made by electron beam depositing

## Slide 14: Possible reasons for damage
Radiation damage
  Densification (~ 3 - 4%)  Changing of refraction index
“Low” pressure conditions
  Arising substoichiometric SiO2
Tritium atmosphere
  Diffusion of hydrogen (isotopes)
Radiochemical processes
No completely suitable literature available  COAting Test EXperiment needed  COATEX
Shi-Zhen et al., Chinese Phys. Lett., 25 (1) 2008
T.M. Stephan et al., SPIE Conf. 1848 1993
R. A.B. Devine, Nucl. Inst. Meth. in Phys. Res. B, 91 378 1994
Pitts, Czanderna, Nucl. Inst. Meth. in Phys. Res. B, 13 245 1986
A. Ibarra et a., J. Nucl. Mater. 367-370, B 1003 2007 J. D. Fowler et al., J. Am. Ceram. Soc. 60 155 1977
Notes: Radiochemical process: Ion He3

## Slide 15: Coating samples
Anti-reflection / high reflectivity coatings for 532nm, 593nm, 1064nm
  Electron beam depositing (EBM)
  Ion assisted beam depositing (IAM)
  Magnetron sputtering (MS)
  Ion beam sputtering (IBS)
1 sol-gel coating from Lawrence Livermore National Laboratory (US)
1 “naked” window (pure SiO2) as reference
Already damaged + new LARA windows (EBD)
I.M. Thomas, Applied Optics, 25 9 (1986)
Decreasing porosity

## Slide 16: Overview of COATEX
Laser
Photo diode
1) Storage of samples in tritium (1-3 weeks)
2) Transmission – Reflection measurement
3) Differential interference contrast microscopy
~200 mbar T2

## Slide 17: Storage of samples in tritium
4 storage vessels in CAPER glove box
Actions in glove box
  Filling, evacuation, flushing with H2 + D2
  Gas chromatography analysis possible
Storage vessel
T2 supply
Coating samples in holder

## Slide 18: T-R measurement principle
Mirror
Laser diode
Beam splitter
50:50
Sample on movable stage
Observable:
Ratio of PD voltages
532nm, 593nm, 1064nm

## Slide 19: T-R measurement setup
Laser diode
50:50 beam splitter
Monitor photo diode
Reflection photo diode
Coating sample on motorised stage
Transmission photo diode

## Slide 20: Installation of T-R setup
Installation in glove box
T-R setup before installation

## Slide 21: Status of COATEX
COATEX has just started
  T-R measurements of all samples before commissioning with T2
  DIC images of already damaged windows
  Storage vessels filled with tritium for 10 – 28 days
  First T-R and DIC measurements performed with samples from storage vessel 1  First results
Experience gained
  Handling of 1 storage vessel per week
  Surface contamination of samples ≤ 107 Bq / cm2
Measurement campaign until ~ spring 2012

## Slide 22: T-R-setup: Commissioning and first results
Preliminary
Preliminary
Preliminary
Sample 6 (uncoated)
Rotated by 90°

## Slide 23: Differential interference contrast microscopy
Sample 5
Old, already damaged, windows before T2 treatment (EBD)
Sample 1 (IAD)
Sample 5 (porous SiO2)
Sample 4 (EBD)
Storage vessel 1: images after ~18 days of tritium
More data needed  Continue tritium loading
All pictures taken at fusion material laboratory (KIT)
Growing / multiplication of damages?

## Slide 24: Summary
KATRIN will measure neutrino mass with 200 meV/c2 sensitivity (90% C.L.)
Laser Raman (LARA) system for monitoring of gas composition
LARA has reached 10-3 precision over 21 days
First results of COATEX after 18 days of tritium exposure
  EBD and sol-gel coatings: First indications for damage visible
  Other coatings: No indications yet

## Slide 26: How is the tritium removed?
Tritium free
injection rate = 5 × 1019 molecules/s
R>107
R>107
p(T2) < 10-20 mbar
< 300 000 Tritium molecules
Transport section
p = 10-11 mbar
Tritium retention

## Slide 27: Consequences of coating damage
Diffuse scattering + (partly) reflection of laser beam
Distortion of laser polarisation  Decrease of LARA accuracy
Damping of scattered light  Decrease of signal/noise ratio and precision
Change of transmission coefficient T(λ) of Raman window  Decrease of accuracy

## Slide 28: The method works!
pcell = 154.8 mbar Laser power = 5 W
Acq. Time = 1000 s
Gas composition from gas chromatography
(5-10% uncertainty)
Isotop. | Conc.
T2 | 0.27 %
DT | 0.49 %
HT | 8.5 %
D2 | not rec.
HD | 8.3 %
H2 | 82 %
All isotopologues can be detected simultaneously
Level of detection
pmin = 0.015 mbar
M. Sturm et al, Laser Physics, 2010, Vol. 20, 2
Notes: 28
Komma runter”

## Slide 29: Intensity variations
Stainless steel vessel wall with (H2, D2)
Time (days)
Permeation into stainless steel
Exchange reaction with T2 of next filling
Evacuation of vessel
Relative intensity Irel (%)
Filling with T2 gas
Isotope exchange reactions in gas
Gas - wall interactions
S. Fischer, et al., Fusion Sci. Technol., 60, 3 2011
Notes: 29

## Slide 30
KATRIN Collaboration 2010