## Slide 1: Low energy and high resolution performance of the MÖNCH hybrid pixel detector
Roberto Dinapoli :: SLS Detectors Group :: Paul Scherrer Institut
Pixel 2016
Notes: MÖNCH is a 25 μm pitch charge integrating detector aimed at exploring the limits of current hybrid silicon detector technology.
It opens new perspectives for many synchrotron applications where currently the detector is the limiting factor e.g. inelastic x-ray scattering (IXS), Laue diffraction and soft X-ray or high resolution imaging.
The reduced input capacitance allows to obtain an impressive low noise of less than 35 electrons RMS and the high segmentation provides high resolution images.
However, the small pixels present many challenges in terms of bump bond yield, electronics design and data throughput.
At this pitch, the charge produced by single X-rays is almost always shared between neighboring pixels.
This allows to interpolate between channels obtaining ~ micron spatial resolution, but on the other hand deteriorates the energy resolving power of the detector.
The possibility of exploiting the charge sharing to improve the spatial resolution requires detailed modeling and characterization, which can also help to retrieve the energy information.
The results obtained using the two existing fully functional prototypes of size 4x4 mm2 (25.6 kpixel) and 10x10 mm2 (160 kpixel) will be presented and the challenges to be faced to develop the planned 4x3 cm2 (1.9 Mpixel) module will be discussed.

## Slide 2: Hybrid Detectors of the SLS Detector group
Pixels
Single
Photon
Counting
Charge
Integrating
172 µm
AGIPD
PILATUS
200 µm
75 µm
EIGER
JUNGFRAU
Microstrips
50 um
GOTTHARD
MYTHEN
Poster by D. Mezza
Talk by G.Tinti

## Slide 3: Challenges and solutions
High resolution
Low noise
Large dynamic range

## Slide 4: Applications
High
spatial
resolution
Large dynamic range
Low
noise
Soft X-rays
XRF
Laue
diffraction
Color imaging
HI-res imaging
(R)IXS

## Slide 5: Applications
High
spatial
resolution
Large dynamic range
Low
noise
Soft X-rays
XRF
Laue
diffraction
Color imaging
HI-res imaging
(R)IXS
MFZ (Maximum Fun Zone)

## Slide 6: Challenges and solutions
High resolution
Low noise
Large dynamic range
25 µm pitch pixels
High gain
Adaptive gain
Bump bonding yield
100% charge sharing
Sensitive to leakage current
Many pixels per unit area, i.e. limited resources per pixel, huge data throughput

## Slide 7: Challenges and solutions
High resolution
Low noise
Large dynamic range
25 µm pitch pixels
High gain
Adaptive gain
Bump bonding yield
100% charge sharing
Sensitive to leakage current
Many pixels per unit area, i.e. limited resources per pixel, huge data throughput

## Slide 8: Prototypes
UMC 110nm AE (Al only, shrunk technology, great quality/price)
Area 4x4 mm2, 160x160 = 25.6kpixels
Five different pixel architectures
  Different preamp
  Dynamic gain switching
  Input overvoltage control
4 analog outputs
2 kHz frame rate (1Gb port)
MÖNCH 02
MÖNCH 01
JUNGFRAU01
Multipitch strip sensor
MÖNCH 01
“Guest” of Jungfrau01 MPW
Test structures
Seven single pixel channels
Wirebonded to a strip detector for calibration

## Slide 9: Basic pixel architecture
CDS
Stage
Cfb
Detector
Test
Calibrate
CS
Preamp
Off-pixel
buffer
Column bus
Column
receiver
PIXEL
CDS
reset
Preamp
reset
Storage
Pads
Single ended
to differential
output driver
(1x/40cols)
MÖNCH 02:
Five different sub-blocks
Column selection
MUX

## Slide 10: Variations
CDS
Stage
Cfb
Detector
Test
Calibrate
CS
Preamp
Off-pixel
buffer
Column bus
Column
receiver
PIXEL
CDS
reset
Preamp
reset
Storage
Caps
Pads
Five different preamp versions (including dynamic gain switching)
Selectable CDS Gain
CDS can be bypassed
two storage caps
Column selection
MUX
Single ended
to differential
output driver
(1x/40cols)

## Slide 11: Variations
Single ended
to differential
output driver
(1x/40cols)
CDS
Stage
Cfb
Detector
Test
Calibrate
CS
Preamp
Off-pixel
buffer
Column bus
Column
receiver
PIXEL
CDS
reset
Preamp
reset
Storage
Caps
Pads
Vth
S
R
DGS
Q
reset
Dynamic gain selection
Cfb2
Cfb
Column selection
MUX

## Slide 12: MÖNCH 03
Area 10x10 mm2, 400x400 = 160k pixels
Architecture optimized for very low noise (single photon)
32 analog outputs
Currently 2 kHz frame rate (10 Gb port)
Up to 6 kHz by design
Region of interest possible (frame rate > 6kHz)

## Slide 13: Bump bonding
Very first prototypes
Single chip processing
4x1 mm2
After optimization
Single chip processing
4x1 mm2
2013 >95%
2014 >99.95%
2015 >99.95%
Wafer processing
1x1 cm2
Standard Indium process
Performed in-house at PSI
Bump diameter ~ 15 μm
Notes: 1SC=40x160=6400 pixels

## Slide 14: Charge sharing and interpolation
Radiation
conversion
Charge collection
after drift and diffusion
Low resolution
analog image
High resolution
digital image
Charge sharing depends on:
  Charge mobility
  Applied bias voltage
  Sensor thickness
  X-ray energy
  Max. 100 photons/mm2/frame to separate single photons
  @10 kHz frame rate =>
    106 photons/mm2/s
    =>Long measurement time

## Slide 15: Clustering
In single photon regime, clustering is possible
    Preserves spectral information
    Suppresses charge sharing
    Improves with single-pixel gain calibration
Raw data
Clusters
Measured
MÖNCH
JUNGFRAU

## Slide 16: Gain, noise and dynamic range
Gain
Adjustable gain for CDS (static) and preamp (dynamic in MÖNCH02)
Single photon sensitivity down to 600 eV (5σ from noise)
  700 eV for MÖNCH03, probably due to readout system
  Dynamic range up to ≈1800 ∙ 1keV γ, ≈150 ∙ 12keV γ
MÖNCH02
-800 -600 -400 -200 0 200 400 600 800 e-
-2.9 -2.1 -1.45 -0.72 0 0.72 1.45 2.1 2.9 keV
36
MÖNCH03
MÖNCH03
IR-laser and ND filters
Noise
Dynamic range
Mo
Ge
Cu

## Slide 17: Moench 03 at low energies
Energies between 1.75 keV and 3.5 keV.
Sensor entrance window ~ 1.5 µm.
Photon-finding algorithm works properly, and gain calibration seems to be portable at low energies within a few percent.
2×2 subcluster spectra

## Slide 18: Iterative correction method
The bins of the η-distribution histogram are assigned to position bin within the pixel
  η flat distribution gives flat flatfield image
Iterative algorithm defines η bins map to obtain flat distribution
Algorithm converges once the flat field image is flat
Position dependent resolution
η bin map
Flat field
1
10
20
30
40

## Slide 19: Interpolation
25 µm pixels
Interpolated
MÖNCH02
TOMCAT 16.7 keV
Kidney stone
2 µm bins

## Slide 20: Interpolation
25 µm pixels
Interpolated
MÖNCH03
W-tube 55kV
EIGER HDI
5 µm bins
Grating-based,G2-less, phase contrast imaging

## Slide 21: Interpolation at low energy
Absorption image of metal grid
PHOENIX beamline at PSI
2 keV energy photon beam
No position interpolation
25 µm bins
With position interpolation
1 µm bins
63 μm pitch
12 μm bar width
51 μm hole width

## Slide 22: Towards larger area
Large area MÖNCH (end 2017?)
  Low noise, dynamic gain switching
  ASIC 2 x 3 cm2 (768 x 1024 channels ~800k pixel)
  Modules up to 4 x 3 cm2 (~ 1.6M pixel)
  Almost 1G pixel after interpolation
DAQ challenge
  Currently 32x40MHz ADCs, 2X10Gb interfaces
    -Already limited by backend (file writing)
  -On the fly (in firmware?) data compression necessary to avoid data throughput overflow -Online clustering
Sensors optimized for soft X-rays
  20-200 nm backplane by FBK

## Slide 23: Conclusions and outlook
We are learning how to tame a 25 µm pixel pitch hybrid detector
  Good bump-bond yield, Low noise, Large dynamic range
  Some imaging applications tested…
  …many more still to be tried e.g. Laue diffraction, IXS, soft X-rays < 2 keV
We have to work on final detector system
    Define specifications (XFEL?)
    Large area ASIC
    Thin backplane sensors (FBK)
    Capability to handle huge data amount
  MÖNCH can be a tool to study charge collection
    Different sensor designs, thickness, e-/h collection
    High-Z sensors

## Slide 24: The SLS detector group at PSI
Aldo Mozzanica
Sophie
Redford
Dhanya
Thattil
Gemma Tinti
Marco
Ramilli
Sebastian
Cartier
Davide
Mezza
Erik
Fröjdh
Dominic
Greiffenberg
Xintian
Shi
Jiaguo
Zhang
Anna Bergamaschi
Roberto
Dinapoli
Martin
Brückner
Christian
Ruder
Lukas
Schädler
Davit
Mayilyan
Th. Huthwelker @PHOENIX, A. Diaz @cSAXS, Zh. Wang, M.Kagias, M. Stampanoni @TOMCAT
Bernd
Schmitt

## Slide 25: Thank you!

## Slide 26: MÖNCH 03 additioal features
Readout on both sides
Improved column buffer
Improved s. ended to diff. off-chip buffer
Storage caps can be kept disconnected from pixel buffer
One storage cap is now a MimCap
Region of interest possible (frame rate > 6kHz)
Additional debugging features

## Slide 27: Nanoscope imaging with Moench
Magnification: ~80x
No Zernike phase dots
Si111 and 40 μm Cu
Siemens star test images
MÖNCH02
25μm pixel
interpolation
5 x stitching
not cooled
Photonic Science CCD
4μm pixel
2 x stitching
cooled
80μm

## Slide 28: Charge sharing
The charge diffuses while drifting to the electrodes
The diffusion length depends on the drift time:
  Charge mobility/diffusivity
  Applied bias voltage
  Absorption depth
    Sensor thickness
    X-ray energy
The charge can be collected by several strips/pixels
  For 320 um thick sensors, 120V bias voltage we have measured a charge sharing region of 15-20 um
MÖNCH
PILATUS
MYTHEN
GOTTHARD
X-ray
JUNGFRAU
EIGER

## Slide 29: Linear charge collection model
d
p
Full charge collection
Charge sharing
Sides
Corners
Noise convolution
α : dchargeSharing /pitch
Q0 : total charge
Q : collected charge
σ : electronic noise ENCrms
A : number of photons

## Slide 30: Dynamic range for MÖNCH 03
Highest Gain Setting
Highest Gain Setting
Dynamic range:
600 3 keV photons at lowest gain setting
18 3 keV photons at highest gain setting
  non-linearity due output buffer saturation
  optimize chip settings
IR laser
ND filters
Lowest Gain Setting
3 keV photons
Kapton thickness

## Slide 31: Energy resolution
Compton
Edge
1.45 keV
Main beam
20 keV
Si escape
18.2 keV
In
3.3 keV
Ti
4.5 keV
Cu
8.0 keV
Au
9.7 keV
Rb?
13.3 keV
No spectral information from single pixel with no cuts
Clustering increases the noise
Cuts to charge sharing as “offline collimation”
  99% rejection rate
Use larger pixels
MÖNCH02 20keV monochromatic data.
The fluorescence lines come from the detector itself.
Ni
7.4 keV
Br
11.9 keV

## Slide 32: Position dependent resolution
ERF(σ)
Resolution=σ
Position dependent resolution
Average better than 2µm
Ultimate lower than 1µm
250 μm
No interpolation
Interpolation
Charge collection efficiency at 16 keV
Selective
radiation damage
Resolution
in X
< 1um RMS
In the corner
Pencil beam scan at cSAXS with nano-focused beam using FZP
Analysis in progress

## Slide 33: Grating-based phase contrast imaging
Phase shift
Creates periodic
intensity
modulations
G2 Analyzer grating
Makes fringes
visible on large
detector pixels
Darkfield
Absorption
No sample
Sample
A detector with sufficient spatial resolution and SNR allows to remove G2
Lower dose
Shorter exposure
No stepping
Larger area
Higher energies
~ 5 um
G1 Phase grating

## Slide 34: Phase retrieval
Absorption
Performed at TOMCAT beamline at PSI
16 keV beam
G2-less works in 2D and matches simulations
Hilbert-transform based algorithm developed to deal with position dependent resolution
Fringes visible only at the boundary between pixels
Si wafer target
Pyramids of different size

## Slide 35: Imaging
After clustering
  Photon counting behavior
  25 µm spatial resolution
…but we can exploit charge sharing within cluster to interpolate
Images acquired MÖNCH03 with W-anode X-ray tube, 55kV 30mA

## Slide 36: Position dependent charge sharing
Looking at the charge collection distribution adds information to the absorption position
Charge collection
Charge collection
Charge collection
Charge collection

## Slide 37: η distribution algorithm
Flat field image acquired previous to the measurement to measure the η correction map
Each (ηx, ηy) pair can be assigned to a position in the pixel after linearization
Works well in 1D but…
…in 2D ηx and ηy are strongly correlated and linearization does not work properly
  Non-flat flat field image and artefacts
Iterative method developed to obtain artefact-free images
ηy = T/(B+T)
ηx = R/(L+R)

## Slide 38: Energy calibration
Developed linear model to deal with 100% charge sharing
Gain used to convert pedestal distribution into noise in e-
Width of the photon peak given by
ENC and Fano Noise
Dynamic range and linearity still to be fully characterized

## Slide 39: Pencil beam scan
Charge collection efficiency at 16 keV
Resolution in X
< 1um RMS
Resolution in Y
̴̴ ≈1.5um RMS
Selective
radiation damage
Beam asymmetry
or
electronic noise
Experiment at cSAXS with nanofocused beam using FZP
8, 16 keV data acquired
Analysis in progress

## Slide 40: Non uniform charge sharing seems the most limiting factor
Absorption image of metal grid
  63 μm pitch, 12 μm bar, 51 μm hole
near the pixel edges
near the pixel center
No position interpolation
25 µm bins
With position interpolation
1 µm bins
PHOENIX beamline at PSI, 2 keV
  In vacuum

## Slide 41: Interpolation limitations
  Higher Signal to Noise Ratio
    High photon energy
    Low electronic noise
  More uniform charge sharing
    Lower photon energy
    Lower sensor bias
    Thicker sensor
    Smaller pitch
Position resolution improves with: