## Slide 1: Study of CsI(tl) eltx
Pre-amp Pulse processing
Kuzim ditribution study
Eltx scheme

## Slide 2: Eltx General layout
Shaper
To DAQ
It’s t is
the main
feature of
this part of the circuit
N times
M times

## Slide 3: Some theory
The output of this filter to a q(t) is:
If t1 different from t2
or
If t1 = t2

## Slide 4
Pre-amp output
Pulsing with a led
Channel 2
Channel 3
The two channel are very similar if not identical this is the output given to an approx Dirca delta (By D. Tagnani& R. De Sangro

## Slide 5: Fit to the shape
t=40 ms

## Slide 6: Kuzmin parametrization 1
He assumes the scintillation charateristc energy realease time
are small with respect to pre-amp stage output (physics driven assumption)
The shaping circuit shaping time are small with respect to the pre-amp integration time. We think this assumption is correct.
Non ne sono sicuro ma:
a) 32 tic della funzione dovrebbero corrispondere a un tic dell' ADC.
b) 1 tic dell’ADC dura DT = 0.567 microsecondi (questo è sicuro)
Quindi direi che un tic della funzione dura DT / 32 = 17.7 nanosecondi
By G. De Nardo

## Slide 7: Kuzmin parametrization 2
Under Guglielmo
assumption we get:
2.2*t=300*17.7 ns
t = 2.6 ms

## Slide 8: Fit of the output shapes
2 hypothesys
1) cr-rc with different ts: fun = [2]*(x**[0])*exp(-(x)*[1])+[5]*(x**[3])*exp(-(x)*[4])
NO. NAME VALUE ERROR NEGATIVE POSITIVE
1 p0 4.25733e+00 6.10326e-06
2 p1 5.14089e-02 2.49127e-07
3 p2 -4.94500e-01 1.26221e-05
4 p3 4.25175e+00 6.10543e-06
5 p4 5.13272e-02 2.47983e-07
6 p5 5.03530e-01 1.28593e-05
EXTERNAL ERROR MATRIX. NDIM= 6 NPAR= 6 ERR DEF=1
3.725e-11 1.511e-13 2.383e-11 9.657e-12 -1.478e-13 2.418e-11
1.511e-13 6.206e-14 1.489e-13 -1.540e-13 5.325e-14 1.448e-13
2.383e-11 1.489e-13 1.593e-10 -2.372e-11 -1.530e-13 -6.866e-11
9.657e-12 -1.540e-13 -2.372e-11 3.728e-11 1.513e-13 -2.411e-11
-1.478e-13 5.325e-14 -1.530e-13 1.513e-13 6.150e-14 -1.493e-13
2.418e-11 1.448e-13 -6.866e-11 -2.411e-11 -1.493e-13 1.654e-10

## Slide 9: Fit result

## Slide 10: Fit Kuzmin
2) Simpler hypothesis:
fun= [2]*((x)**[0])*exp(-(x)*[1])
EXTERNAL ERROR MATRIX. NDIM= 3 NPAR= 3 ERR DEF=1
7.256e-03 4.680e-05 -1.190e-16
4.680e-05 3.059e-07 -7.653e-19
-1.190e-16 -7.653e-19 1.955e-30
FCN=327.797 FROM MIGRAD STATUS=CONVERGED 61 CALLS 62 TOTAL
EDM=7.25556e-09 STRATEGY= 1 ERROR MATRIX ACCURATE
EXT PARAMETER PARABOLIC MINOS ERRORS
NO. NAME VALUE ERROR NEGATIVE POSITIVE
1 p0 1.00777e+01 8.51800e-02
2 p1 6.98135e-02 5.53098e-04
3 p2 4.10362e-15 1.39827e-15

## Slide 11: Results

## Slide 12: Choice
the second Hypothesis is simpler:
The circuit is:
10
If R1C1 << t pre-amp it derives the signal

## Slide 13: The output for that circuit is
Where the input is assumed to be a q(t) the pre-amp output is close to it (t=40 ms)
In our case n=10
P0=n=10 fit result
P1=1/t t= 14.32*17.7= 250 ns fit result
P2=Vin (1/t)n= 4.1*10-15
Abbiamo un piccolo problema:
Se assumiamo che Guglielmo abbia ragione ne segue:
= 250 ns (tick da 17.7 ns).
Altrimenti:
t = 250*32 ns = 8 ms (tick da 17.7*32 ns).

## Slide 14: Pile up simulated studies
Input signal to pre-amp
Pre-amp output signal

## Slide 15: Pile –up 2
After the derivative of
the pre output
After convolution with kuzmin distribution

## Slide 16: Codes
void Computederiv(const double* const input,
double* output,
int len)
{
int j;
for(j=0;j<len-1;j++)
{
output[j-1]=input[j+1]-input[j];
}
}
void Convolve(const double* const input,
unsigned int inputLength,
const double* const filter,
unsigned int filterLength,
double* output)
{
unsigned int i;
unsigned int j;
unsigned int lengthOfOutput = inputLength
+ filterLength - 1;
for(i = 0; i < lengthOfOutput; ++i)
{
output[i] = 0.f;
}
for(i = 0; i < inputLength; ++i)
{
for(j = 0; j < filterLength; ++j)
{
output[i + j] += input[i] * filter[j];
}
}
}