## Slide 1: Radio-Guided-Surgery: status and perspectives
R. Faccini
University “La Sapienza” and INFN Rome
Incontro Nazionale Fisica Nucleare 16/11/2016
Chirone Chir2
http://arpg-serv.ing2.uniroma1.it/arpg-site/index.php

## Slide 2: Overview
Radio-Guided Surgery(RGS): basic concepts
g probes – briefly
b+ probes
Novel development: b- RGS

## Slide 3: Radioguided surgery

## Slide 4: Radio Guided Surgery
Each tumor requires its own tracer
Probe adjustable to needs
PET/SPECT scan to estimate receptivity and background
Administration of radio-tracer
During surgery a probe is used to detect residuals/lymphnodes

## Slide 5: Intra-operative probes
Needed when
  Complete Tumor resection is mandatory
    example: brain, pediatric tumors
  Identification of lymph-node is needed
    Example: breast tumor, melanoma
  Tumor identification can reduce invasiveness
    Example: parathyroid adenoma, insulinoma,…
Notes: PANCREAS

## Slide 6: Critical Aspects
Identification of clinical applications
Choice of the radio-nuclide
Identification of tracer and its match with the radio-nuclide
Probe design
Tests

## Slide 7: Possible radio-nuclides
Already tested nuclides include
SPECT
PET
Molecular RT

## Slide 8: Radioguided surgery: g probes

## Slide 9: Existing: gamma probes
Choice of radio-nuclide: gamma emitting, in particular 99mTc
- well known detectors (camera or probes)
- used for SPECT: large number of known radio-tracers
Applications of interest:
  Search for tumor residuals (colon cancer, parathyroid adenoma, osteoid osteoma)
  Complete sentinel-node mapping (malignant melanoma and breast cancer)

## Slide 10: Gamma probes applications
TARGET
TUMOR
UTILITY
Local non-specific tracers
Mariani, Giuliano, Strauss 2004
Systemic administration

## Slide 11: LIMITS OF g-RGS
140 keV photons
 attenuation in body ~8cm
probe
Long range of gamma’s involve:
  exposure of medical personnel
  Background from healthy organs
Difficult to apply in:
Brain tumors
Abdominal tumors
Pediatric tumors

## Slide 12: Radioguided surgery: b+ probes

## Slide 13: b+ RGS
Use of b+ decaying isotopes (same as PET)  detect positron (little penetration)
Dual detectors for gamma background subtraction
probe
e+ (signal)
g
Two channels:
one sensitive to both positrons and gamma (14) (N1=P+G)
one sensitive only to gamma (12) (N2=G)
signal=N1-N2
Daghighial et al 1994
Raylman et al 2001

## Slide 14: b+ RGS: developments
1994 Daghighian: two different detectors for beta and g
2001 Raylman: integrated dual-detector with semi-conductor devices
2005 Yamamoto: plastic+BGO phoswitch
Tests with point-like source
No shielding
Shielded against b
ratio

## Slide 15: b+ RGS: clinical applications
Brain tumors: difficult use of pre-operative information
2007Bonzom et al :
  attach it to the surgical device
  Different tracers considered to carry 18F
2009 Bogalhas et al:
  Test on phantom
Conclusions

## Slide 16: b+ RGS
Use of b+ tracers(positrons): pros
  Detect positron that travels ~100 times less
  PET tracers can be used
Use of b+ tracers: cons
  Large irradiation of medical personnel (~20 mSv/hr)
  Positrons annihilate and produce photons  background  need to subtract background
        Longer time to have a response
        More encumbrant detector

## Slide 17: Radioguided surgery: b- probes

## Slide 18
Use of b- tracers (electrons): pros
  Detect electrons that travel ~100 times less than γ
  Tracers with 90Y can be used (already used for Molecular RT)
  No background from gamma
      Shorter time to have a response
        Smaller administered activity
      Smaller and more versatile detector
      Very reduced effect of nearby healthy tissues
      Reduced dose to medical staff
A CHANGE IN PARADIGM
EXTEND RGS TO MORE CLINICAL CASES
e- (signal)
E. Solfaroli Camillocci et al, Sci. Repts. 4,4401 (2014)
probe
Notes: brevetto!?

## Slide 19: RESEARCH PATH
Probe Prototypes

## Slide 20: Low energy e- detector
p-terphenyl as scintillator
  High signal:
    LY(pterf)=3LY(Anthracene)
  Low Z
    Low sensitivity to photons
  Small attenuation length
    l=(4.73 ± 0.06) mm
R. Faccini et al, Properties of P-Terphenyl as detector for a, b, and g radiation, IEEE Trans. on Nucl. Sci. 2014; 61: 1483-7
Light sensor
scintillator
g
e-
Usually unused because signal attenuates if detector too thick
 Not for low energy electrons
2mm
Note: also in case of pure b- emitters gamma rejection is important because of brehmsstrahlung

## Slide 21: Prototypes
Core: cylindrical scintillator of p-terphenyl d=2.1mm, h=1.7mm
encapsulated into a PVC ring to shield it against radiation coming from the sides;
inserted as a tip inside an easy handling aluminum body.
A thin black PVC cap makes the enclosure light tight.
Two options for light collection:
  Scintillating fiber and PMT
  SiPM (SensL B/C-series)
14cm
1cm
0.4mm thick PVC layer in front
PVC
7mm
2.1mm
CORE
Optical fiber
PMT
Constraints on medical devices apply

## Slide 22: Electronics Read-out
Electronics read-out is portable and customized to match the surgeon needs
acoustic and visual alarm;
wireless data transfer;
no connection with electrical line (bacteries)
user interface available both for PC or tablet.
No Danger of electric discharge on patient
ARDUSiPM
LABE-INFN RM1

## Slide 23: RESEARCH PATH
Probe Prototypes
23
Lab tests (phantom factory)
Measure spatial sensitivity
Gamma rejection
Estimate performances on phantoms
Estimate dose on surgeon
Estimated dose on surgeon administering 3MBq/kg:
  1mSv/hr on surgeon’s hands
  0.13 mSv/hr on medical personnel
E. Solfaroli Camillocci et al, J. Phys.: Conf. Ser. 620 012009(2015)

## Slide 24: Sensitivity to Electrons
Detection efficiency on 90Sr point source
  Rate 3.8·105 cps/MBq.
  εβ=40%
Scan in water
  Eβ>500keV.
  Detection efficiency εβ>80% in the β- 90Y energy range.
Scan with different thicknesses of water

## Slide 25: Background Rejection
90Y in water
Almost transparent to Bremsstrahlung photons.
ɛɣ<10-3
ɛɣ<10-2
60Co Eɣ=1170, 1330KeV
137Cs Eɣ=662KeV
133Ba Eɣ=80-350KeV
Background is mainly due to photons coming from Bremsstrahlung.
Sensitivity to photons
Bremsstrahlung Eɣ spectrum

## Slide 26: “Ad-hoc” Phantoms
Tumor residual V=0.05ml
embedded in tissue with A/10
Motorized scans with S4-Probe
All the possible configurations of tumor residual embedded in healthy tissue.
To simulate tumor remnant embedded in healthy tissue.

## Slide 27: Active Spot Identification
False positive probability
TNR=10

## Slide 28: Human Factor
To include the human factor in the test colleagues were asked to simulate the surgeon:
All “surgeons” required at least 4-5 seconds per position to take a decision.
Phantoms simulating tumor remnants embedded in healthy tissue with different TNRs

## Slide 29: RESEARCH PATH
Probe Prototypes
29
Lab tests (phantoms)
Identification of proof of principle
Start with existing radiotracers:
Somatostatine analogues marked 90Y
Meningioma is avid of 90Y-DOTATOC
 Able to detect 0.1 ml residuals administering only 1MBq/kg

## Slide 30: RESEARCH PATH
Probe Prototypes
30
Lab tests (phantoms)
Identification of proof of principle
Identification of first clinical cases
Start with existing radiotracers:
Somatostatine analogues marked 90Y
Estimate sensitivity from PET with 68Ga
First clinical cases:
Glioma (low and high grade)
GEP-NET (start from small bowel)

## Slide 31: DOTATOC uptake in glioma
DOTATOC is a somatostatin analog  known receptivity from NET .. but glioma?
Even if TNR too low for therapy, can it be used for RGS?
Use 68Ga-DOTATOC PET scans to estimate signal and background
ROI definition
meningioma
HG glioma
Non-tumor
F. Collamati et al,. J Nucl Med 56 (2015) 3-8

## Slide 32: Glioma vs Meningioma
meningioma
glioma
Meningioma has a definitely larger uptake, but is the glioma one acceptable?
Use FLUKA simulation to translate from activities to false positive (FP) and false negative (FN) rates
Consider a residual identified if FP<1% and FN<5%
TNR of glioma is acceptable if probe can take >~6s to give answer
TNR
0 40 80 120
TNR
0 10 20
F. Collamati et al,. J Nucl Med 56 (2015) 3-8

## Slide 33: RGS for meningioma
Very large uptake
Can inject as low as 0.5 MBq/kg
*
**
* Time needed to detect 0.1 ml residual if 3MBq/kg are administered
** Activity that needs to be administered to achieve 1s response time

## Slide 34: RGS for glioma
Needs to wait for ~6s, but it works
Margins to improve probe
*
**
* Time needed to detect 0.1 ml residual if 3MBq/kg are administered
** Activity that needs to be administered to achieve 1s response time

## Slide 35: Time Evolution of uptake
35
Lab tests (phantoms)
Identification of proof of principle
Values estimated from dosimetric SPECTs with 177Lu-DOTATOC
F. Collamati et al,. J Nucl Med 56 (2015) 1501-6
TNR for liver NET vs time
CONCLUSIONS:
GEP-NETs (small bowel, insulinoma, …) are a good candidate and similar among each-other
the best SUV and TNR are achieved if surgery is 24hrs after injection
Notes: A similar clinical study was performed on SPECT images of patients with liver NET, allowing us to conclude that the TNR is good enough to perform B- RGS within 1 second per spot by injecting an activity that is one third of the one used for a PET.
In this study we also evaluated the variation in time post injection of the TNR of liver NET, finding it to reach a maximum at about 24h p.i., that thus represents in view of a possible protocol the best time lapse between the injection and the performing of RGS

## Slide 36: RESEARCH PATH
Probe Prototypes
36
Lab tests (phantoms)
Identification of proof of principle
Identification of first clinical cases
Ex-Vivo tests on Meningioma
We are here

## Slide 37: Ex-vivo test on meningioma
PET with Ga68 on Sep 14th
  Tumor SUV ~2g/ml (relatively low, but enough)
  TNR ~ 14 (good)
E. Solfaroli et al., submitted to JNM
8mCi Y90—DOTATOC on Oct 9th
Surgery on Oct 10th

## Slide 38: The Samples
A
C
B
E
F
G

## Slide 39: Evaluating the samples rate

## Slide 40: Results
Residuals as small as 0.2ml are visible
Predictions with simulation are reliable (115 cps predicted, 105 observed)
Healthy brain ~1cps (simulation) also infiltrated dure can be identified
+ Confirmed very
low exposure of
medical personnell
E. Solfaroli et al, Physica Medica 32 (2016), pp. 1139-1144

## Slide 41
IRRADIATION MEASUREMENTS

## Slide 42: Probe developments
Develop prototype for endoscopic (laparoscopy/Da Vinci …) use
  Multichannel design with p-terphenil
Use CMOS technology to lower energy threshold (with INFN& Uni PG)  allow use of other isotopes
  Matrix design for basic “imaging”
CHIR2

## Slide 43: Radiotracers for Radioguided surgery

## Slide 44: g-RGS: the Tc crysis
99mTc is produced from 99Mo decays
Current production method
  Canada 50%, Netherland 40%
  + …
Need alternative methods
  100Mo(p,2n)99mTc [direct 99mTc production]  Canada
  100Mo(,n)99Mo, 238U(,f)99Mo USA, Canada
  98Mo(n,)99Mo  testing parent 99Mo production
  100Mo(n,2n)99Mo  testing (easier to implement)
Nuclear plants
being dismissed
Developments in accelerator physics required

## Slide 45: b+ RGS: different isotopes
18F not the only tracer
Mean b+ energy critical parameter
Each tracer has different clinical applications
(87.9%)

## Slide 46: b- probe on b+ isotopes
Positron/g separation with copper layers
Results scaled to 10kBq/ml
10 kBq/ml | 18F | 68Ga
Counts no shield (e+g) | 11.9±0.3 | 51.2±0.8
Counts with Cu shield (g) | 3.0±0.1 | 1.7±0.1
Difference (e) | 8.9±0.3 | 49.5±0.8

## Slide 47: b- RGS: Other isotopes (I)
Isotope | T1/2(h) | Eg(keV) | Ig(%) | EPb(keV) | Ib(%) | Use
18F | 1.8 | 511 | 200 | 633.5 | 97 | FDG, most tumors
21Si | 2.6 |  |  | 1491 | 100 | Same family as C (used in PET)
32P | 343 |  |  | 1710 | 100 | Brain tumors [1]
67Cu | 62 | 93/184 | 16/48 | 377/468/561 | 22/20/99 | Cu-64 in immuno-PET[3]
83Br | 2.4 |  |  | 935 | 99 | Same family as F
90Y | 64 |  |  | 2280 | 100 | NET & Brain tumors[1]
97Zr | 17 | 743 | 93 | 759 | 88 | 89Zr used in immuno-PET [2]
131I | 192 | 365/637 | 82/7 | 334/696 | 7/90 | Thyroid[1]
133I | 20.8 | 530 | 87 | 1227 | 83.4 | Thyroid
153Sm | 46 | 103 | 29 | 635/704/808 | 31/49/18 | Bone Cancer[1]
177Lu | 160 | 112/208 | 6/10 | 500 | 79 | NET & Brain tumors[1]
188Re | 17 | 155 | 15 | 1962/ 2118 | 25/72 | Bone & Liver [2] Cancer
used in Nucl. Med.
Isotopes of those used in Nucl. Med.
[1] Review in Yeaong C.H. et al. J. Zhejiang Univ. SCIENCE B Vol.15 No.10 P.845-863 (2014)
[2] van de Watering F.C. et al. Biomed Res Int. 2014;2014:203601.
[3] Asabella A.N. et al. BioMed Res. Intl. 786463 (2014)
Same chem family of those used in Nucl. Med.
C. Mancini-Terracciano et al., arXiv:1610.09246

## Slide 48: b- RGS: Other isotopes (I)
Isotope | Eg(keV) | Ig(%) | EPb(keV) | Ib(%) | T*min(s)
18F | 511 | 200 | 633.5 | 97 | >25
21Si |  |  | 1491 | 100 | 0.4
32P |  |  | 1710 | 100 | 0.3
67Cu | 93/184 | 16/48 | 377/468/561 | 22/20/99 | >25
83Br |  |  | 935 | 99 | 0.9
90Y |  |  | 2280 | 100 | 0.5
97Zr | 743 | 93 | 759 | 88 | 0.8
131I | 365/637 | 82/7 | 334/696 | 7/90 | >25
133I | 530 | 87 | 1227 | 83.4 | 2.8
153Sm | 103 | 29 | 635/704/808 | 31/49/18 | 3.1
177Lu | 112/208 | 6/10 | 500 | 79 | >25
188Re | 155 | 15 | 1962/ 2118 | 25/72 | 0.4
* Time needed to detect a 0.1 ml residual with FN<5% FP<1% if 3MBq/kg are administered and SUV=4, TNR=8
OK with current probe
High SUV and TNR required: improvements in g rejection and energy threhold useful
Significant improvements in g rejection and energy threhold required

## Slide 49: b- RGS: alternative tracers with 90Y
Synthesis of new radio-
tradiotracers (with 90Y)
Monoclonal antibodies
(NIMOTUZUMAB) for EGFR receptors
MIBG
Nucl . Medicine and Chemistry
Chemistry and CTF
Development of nano-scale carriers composed of polymers, antibody and ittrium
Y
AB

## Slide 50
Nuclear Physics
Nuclear Medicine
Particle accelerators
Radiation detection
Oncology
Radio-Chemistry
Radio-Guided-Surgery

## Slide 51: BACKUP

## Slide 52: FOOT in pills
Sections/Labs: Bologna, Frascati, Milano, Napoli, Perugia, (Pavia), Pisa, Roma1, Roma2, Torino, Trento
People: ~50 researcher, ~24 FTE
DATA taking foreseen @ CNAO, TIFPA, LNS, BTF
Experiment with translational approach: focus on nuclear physics, physics applied to medicine and radioprotection in space

## Slide 53
Generation of calibration database for deconvolution of material absorption using FLUKA MC trained with experimental data.
Development of lookup table to connect the emission point distribution of the secondary charged particle with the BP, in different, non homogeneous material
Dosimetric imaging
BEAM
Profiler
brain
True
emission
profile
Reconstructed emission profile
After deconvolution

## Slide 54: The INSIDE PET system
DAQ sustains annihilation and prompt photon rates during the beam irradiation
Two planar panels each 10 cm x 20 cm wide. Each panel will be made by 2 x 4 detection modules
Each module is composed of a pixelated LYSO scintillator matrix 16 x 16 pixels, 3x3 mm2 crystals, 3.1 mm pitch, for a total sensitive area of 5x5 cm2
One SiPM array ( 16x16 pixels) is coupled to each LYSO matrix. 200 ps FWHM TOF capability
Notes: Two planar panels each 10 cm x 20 cm wide. Each panel will be made by 2 x 4 detection modules.
Each module is composed of a pixelated LYSO scintillator matrix 16 x 16 pixels, 3 mm x 3 mm crystals, pitch 3,1mm, for a total sensitive area of 5 cmx 5 cm.
One array of SiPM ( 16x16 pixels) are coupled one-to-one to each LYSO matrix.

## Slide 55: Current Efforts (II)
Setup the whole chain for radio-tracer development:
- cold synthesis
- hot synthesis
- in vitro tests
- animal tests
 equip a facility for
animal tests with radioactivity
 Need imaging capable to measure biodistribution of b- emitting tracers  SPECT with Brehmsstrahlung
Nucl . Medicine and Chemistry
Chemistry and CTF