## Slide 1: WP 7 (INFN-RDH) Detector for high intensity beam M 6 (INFN-IRPT) New TERA chip development
Camere Monitor per fasci di alta intensità
R. Cirio, F. Fausti, L. Fanola Guarachi, S. Giordanengo, F. Marchetto, G. Mazza, V. Monaco, R. Sacchi, A. Vignati
INFN e Università di Torino
Notes: I will describe the progress of the work project 7 and 6 that involves

## Slide 2: Fasci di elevata intensità
Typical Characteristics for high flux pulsed charged particle beams | 
Pulse frequency (kHz) | 0.2 – 1
Pulse Length (μs) | 5 – 20
Number of particles per pulse (prot/pulse) | 107 -108
Instantaneous Intensity (prot/s) | 1012 -1014 (~µA)
Acceleratori di nuova generazione a fasci pulsati
  Sincrociclotroni, Accel. lineari, Cyclinacs, FFAGs, ...
  Acceleratori Laser Driven
Nuove modalità di trattamento ipofrazionato
Notes: Nowadays new compact accelerators for charged particles beam therapy have been proposed where high beam intensities occur in short pulses. The typical characteristics of high flux beam are:

## Slide 3
Fasci di elevata intensità
Chart: Ebeam (MeV)
Intensità di corrente raccolta nella camera vs. Ebeam
altà densità di ionizzazione → inefficienza di raccolta
saturazione dell'elettronica di FE
~ 2-3 ordini di grandezza
Notes: Nowadays new compact accelerators for charged particles beam therapy have been proposed where high beam intensities occur in short pulses. The typical characteristics of high flux beam are:

## Slide 4
Efficienza di raccolta
Ricombinazione iniziale e colonnare
ricombinazione tra cariche generate lungo la traccia
non dipende dal dose rate
corretta attraverso calibrazione dosimetrica delle camere
teoria di Jaffe
Ricombinazione di volume
ricombinazione tra cariche generate da tracce vicine
dipende dal dose rate, quantità che si vuole misurare !
diverse parametrizzazioni in letteratura, (Boag, Wilson, Townsend…)
In generale
  cresce all’aumentare della densità di ionizzazione nel gas
  diminusce al crescere del rapporto E/d=V/d2

## Slide 5
Rivelatore a doppio gap
Intensità di corrente del fascio
Intensità di corrente raccolta agli elettrodi
L’efficienza di raccolta depende da:
tensione V tra gli elettrodi
distanza d tra gli elettrodi
densità di ionizzazione n0
saturazione
Con V e d fissati il rapporto tra le correnti misurate nelle due camere depende solo da n0
d2=2d1
Notes: For High Flux Beams the recombination phenomenon increases and therefore the total charge collection is inefficient unless a very high voltage between the electrodes is applied. The solution is a detector with different gaps….

## Slide 6
f1 = Efficiienza di raccolta della camera 1
f2/f1 = Rapporto tra le efficienze
E’ dato dal rapporto (normalizzato per le distanze) tra le correnti delle 2 camere
Si determina l’efficienza della camera 1
Rivelatore a doppio gap
→ relazione univoca tra efficienza f1 e rapporto f2/f1
→ determinabile sperimentalmente

## Slide 7
Rivelatore Multi-Gap
3 IC con anodi e catodi indipendenti, gap da 0,5 - 1 - 1,5 cm
Elettronica di lettura basata su front-end TERA08 modificato

## Slide 8
Caratterizzazione del rivelatore
Carbon Ion Beam 120 MeV/u @ CNAO
Proton Beam at Bern University Hospital (Cyclotron of 15MeV)
Fascio pulsato @ CATANA (62 MeV)
Nota:
Efficienza determinata variando la tensione e fittando i dati con modelli di ricombinazione

## Slide 9
Misure a CNAO
Fit globali delle tre camere con modello di Boag-Wilson, a tre flussi differenti

## Slide 10
Densità di ionizzazione vs flusso del fascio
Misure a CNAO

## Slide 11
f2/f1
f3/f1
Misure a CNAO
caratteristica del rivelatore

## Slide 12
Caratteristica del rivelatore
Bern Cyclotron (protons 15 MeV)
LNS
(pulsed protons 62 MeV)
Conclusioni
verificato la validità del metodo proposto
caratteristica indipendente dall'intensità/pulse width
caratteristica dipende dalla qualità del fascio

## Slide 13
Risultati
Tesi di Dottorato di L. Fanola Guarachi
NIM A 798 (2015) 107- 110 (caratterizzazione del readout) + articolo in preparazione per risultati si fascio
Camera doppio gap realizzata da De.Tec.Tor per ELIMED

## Slide 14
TERA09 Design
Basato sull'architettura di TERA08:
Same process technology AMS CMOS 0.35 μm
Increase clock from 100 to ~ 300 MHz
Max count frequency ~ 80 MHz (20 MHz in TERA08)
Counters sums are integrated in the chip
Cooperation agreement signed with De.Tec.Tor S.r.l for the development of the new chip → PhD Federico Fausti;
A joint patent INFN/UniTo/De.Tec.Tor has been submitted;
Premio Marconi della SIF 2015 a Federico Fausti (De.Tec.Tor) e Simona Giordanengo (INFN)

## Slide 15
Block diagram of TERA09
38 bit
sums triggered by load signal
half-full register warning signal provided

## Slide 16: The current to frequency converter
feedback capacitor Cint 600fF → 1,2 pF
OTA bias current 200 μA → 800 μA
clock cycle saved in the FSM, single 200 fF Csub
post-layout simulations 100 MHz → 320 MHz clock

## Slide 17
ASIC submission : July - Sep. 2015
Test board design: Oct.-Nov. 2015
Chips from foundry: Dec. 2015
ASIC characterization: beginning 2016
Test radiation tolerance: 2016
Integration with detector and beam test: 2016
Project status and plans
as presented at RDH meeting
(Rome, Dec. 17st 2014)

## Slide 18
Submission date to Europractice: Sep. 14th 2015
Submission date to AMS: Sep. 21st 2015
Samples out from AMS: Nov. 6th 2015
Packaged samples from Europractice: Nov. 23rd 2015
Test board PCB design: Oct. 2015
Test board PCB manufacturing: Nov. 2015
DAQ system with LabVIEW FPGA: Nov.-Dec. 2015
ASIC characterization: Dec.-March 2016
Integration with detectors and beam tests
Project status and plans:
Current state

## Slide 19
TERA09 (23/11/2015)

## Slide 20
TERA09 Test board

## Slide 21
DAQ based in NI FPGA

## Slide 22
Screenshots: pedestals
+
=

## Slide 23
+
=
i=1μA, clk=250Mhz
Screenshots: 1 ch signal

## Slide 24
Results 100 MHz clk
Chart: 
| 7.1316999999999995 | 6.6619999999999999 | 6.1476000000000006 | 5.6849999999999996 | 5.1612 | 4.6890000000000001 | 4.1627000000000001 | 3.6949999999999998 | 3.1656 | 2.7 | 2.1673 | 1.7050000000000001 | 1.1689000000000001 | 0.57299999999999995 | 0 | -0.56999999999999995 | -1.1345999999999998 | -1.7170000000000001 | -2.1335000000000002 | -2.714 | -3.1318999999999999 | -3.7109999999999999 | -4.1301999999999994 | -4.7089999999999996 | -5.1298999999999992 | -5.7069999999999999 | -6.1171000000000006 | -6.6849999999999996 | -7.1014999999999997
Linearity | 25003045.6 | 24999892.6 | 25002255.3 | 25002860.2 | 24036687.3 | 21860135.8 | 19446152.5 | 17255081.2 | 14798745.9 | 12614885.4 | 10138385.3 | 7962840.9 | 5471028.9 | 2660709.6 | 0 | -2697961 | -5306423.4 | -8038515 | -9957188.7 | -12669301.1 | -14589031.7 | -17291632.3 | -19223738.5 | -21915430.8 | -23854828.7 | -25002990.5 | -25002785 | -24999957.3 | -24999956.5
deviation from linearity ± 0,5 %
channel saturation in TERA08

## Slide 25
Chart: 
| 12.477 | 11.493 | 10.494 | 9.4939999999999998 | 8.4949999999999992 | 7.4950000000000001 | 6.4960000000000004 | 5.4989999999999997 | 4.5 | 3.5009999999999999 | 2.5030000000000001 | 1.504 | 0.50600000000000001 | 0 | -0.502 | -1.5009999999999999 | -2.4990000000000001 | -3.4969999999999999 | -4.4960000000000004 | -5.4950000000000001 | -6.4930000000000003 | -7.4939999999999998 | -8.4939999999999998 | -9.4930000000000003 | -10.491 | -11.481999999999999 | -12.492000000000001
Linearity | 50005724.9 | 50006362.3 | 50007371.6 | 45695205.6 | 40632941.2 | 35781448 | 31005651.6 | 26241279.8 | 21494691.8 | 16712744.8 | 11951249.8 | 7172113.9 | 2394043.8 | 0 | -2487828.6 | -7375985.8 | -12251706.6 | -17118335.1 | -21997344 | -26861245.1 | -31744897.2 | -36630444.4 | -41542922.1 | -46442127.2 | -50006156.8 | -49999854 | -50006434.7
Results 200 MHz clk

## Slide 26
Chart: 
Results 265 MHz clk
deviation from linearity at large positive currents ~ 4% under investigation
strange features above 270 MHz frequency

## Slide 27: Conclusioni
27
Parte digitale testata fino a 265 MHz
Caraterizzazione della parte analogia in progress
Alcune cose da capire/testare, in particolare ad alte frequenze
Da fare:
Disegnare front-end board (Riunione con De.Tec.Tor a breve)
Test di danneggiamento da radiazione
  necessari ?
Convenzione con De.Tec.Tor per sfruttamento commerciale TERA09
pensare al prossimo chip

## Slide 28
Thanks for your attention

## Slide 29
64 channels
registers
Sum of 4 channels
Sum of 16 channels
Sum of 64 channels
Digital part of the chip finalized

## Slide 30
TERA09 Design
Several attempts to modify the analog part to increase the frequency of the recycling integrator.
Two examples of OTA design:
Mirrored Configuration
Folded Cascode
Submission of the first TERA09 prototype scheduled for July 27th

## Slide 31: Camera Multi-Gap
Material | Thickness (mm) | 
Mylar | 12,00 E-3 | Cathode
Aluminum |  | 
Nitrogen/Air | Gap | Frame
Aluminum | 17,00 E-3 | Anode
Kapton | 25,00 E-3 | 
Notes: The beam passes through 3 parallel ionization chambers with independent anodes and cathodes separated by a different gaps with the characteristics described in the table…

## Slide 32
Detector Assembly
Notes: Every anode or cathode is glued onto a fiberglass frame to assure rigidity and flatness uniformity, this is done in a series of steps. First special stretching procedure using small clamps and springs . Afterwards the anode or cathodes are glued onto the frame and finally after one day the clamps are removed.

## Slide 33
4
3
2
1
8
7
6
5
12
11
10
9
Detector Assembly
BEAM
anode
frame
cathode
anode
anode
cathode
cathode
frame
frame
frame
frame
frame
Notes: Each frame containing an anode or cathode is placed into the box taking into account its position along the beam line. And when the package is completely assembled the readout electronics is placed.

## Slide 34
Test with Carbon Ions at CNAO
Synchrotron
Beam lateral FWHM: 4 mm
Beam Flux:
≈ 3×107 (part/s)
≈ 2×107 (part/s)
≈ 6×106 (part/s)
Beam Energy (120MeV/u)
Voltage between electrodes: 13-400 V
108 particles
Notes: For each voltage and each flux the monitor was irradiated with 10 to the 8 particles

## Slide 35
Continous beam of carbon ion at CNAO - Synchrotron, 120 MeV/u
Notes: The number of counts, corresponding to the delivery of 10 to the 8 carbon ions, as a function of the voltage applied between the electrodes for the different beam fluxes is shown.
for all the uxes, the saturation
curves at lower potential of the three ionization chambers are overlapped.
Such behavior is due to the fact that, each ionization chamber has dierent gap and thus different collection effciencies, giving rise a different slopes for
the three saturation curves.

## Slide 36
Boag’s Theory
Boag – Wilson Theory
Townsend Theory
1

## Slide 37
where
By the minimization of χ2 it is possible to find Csat, and the nO and K:
Boag-Wilson Theory
m

## Slide 38
Normalizing the data to chamber 2:
Global Fit

## Slide 39
Nuovo chip TERA09
Per le applicazioni previste è necesario aumentare il range dinamico > 100X rispetto al chip TERA08
Soluzione studiata e testata su TERA08;
modifiche alla parte analogica e digitale del chip
finanziamento tramite premiale IRPT
   120 k€ + AdR biennale
sottomissione primo prototipo a fine luglio
secondo prototipo nel 2016
1 canale: saturazione a  4 A
64 canali in parallelo: saturazione a  270 A
Trasferimento Tecnologico
co-finanziamento della ditta De.Tec.Tor , interessata allo sfruttamento commerciale, Cooperation Agreement con INFN
Sottomessa domanda di brevetto INFN-Unito-De.Tec.Tor

## Slide 40
To do 2015 - 2016
Testare la camera con fasci pulsati con le caratteristiche descritte
Integrare il nuovo chip TERA09 nella camera

## Slide 41
The PCB design
schematic
3D representation
gerber

## Slide 42
The PCB functional blocks diagram

## Slide 43
A. La Rosa et Al., Nucl. Instr. and Meth. A 583 (2007) 461-468.
First linearity study: the CAD prediction
The single channel saturation point is moved from 4μA to 16 μA
TERA08
TERA09
Using a Qc of 200fC
(most used in therapy)
The digital architecture (with an integrated channels parallel connection) permits to extend the current input range up to 1mA (and over, with a Qc that can reach 660fC)
measurements
post layout simulations

## Slide 44: Addressing scheme (followed also for the digital logic tests)
Channel address via 7 bit bus
  0b5b4b3b2b1b0: select register channel b5b4b3b2b1b0
  1b5b4b3b2x0 : select 1st level sum registers
  1b5b4xx01 : select 2nd level sum registers
  1xxxx11 : select 3rd level sum register
44

## Slide 45
On-line control of the sums
Reset interaction and correct settings (rstA, rstD, rst_sync)
Half_full warning
Digital logic tests
These functionalities have been tested with a
LVDS clock up to 265MHz  66.25 MHz of converter frequency
 13 μA of input current
(single channel current injection)
Clock generator Agilent 8113A

## Slide 46
TERA09 Characterization

## Slide 47
Conversion linearity
&
Quantum charge fluctuations
Changing the clock frequency
On board 350MHz vco
SMA for external clk source

## Slide 48
Chart: Charge quantum vs clk frequency (ΔV= 1V, 2V, 3V)
| 50 | 50 | 50 | 100 | 100 | 100 | 150 | 150 | 150 | 200 | 200 | 200 | 250 | 250 | 250 | 265 | 265 | 265
 | 2.0819e-13 | 4.1814e-13 | 6.1666e-13 | 2.1028e-13 | 4.1941e-13 | 6.1773e-13 | 2.0058e-13 | 4.0927e-13 | 6.1272e-13 | 2.1061e-13 | 4.1338e-13 | 5.9863e-13 | 1.9023e-13 | 3.848e-13 | 5.8066e-13 | 1.8697e-13 | 3.7903e-13 | 5.732e-13

## Slide 49
100 MHz clk
Chart: 
| 5.1612 | 4.6890000000000001 | 4.1627000000000001 | 3.6949999999999998 | 3.1656 | 2.7 | 2.1673 | 1.7050000000000001 | 1.1689000000000001 | 0.57299999999999995 | -0.56999999999999995 | -1.1345999999999998 | -1.7170000000000001 | -2.1335000000000002 | -2.714 | -3.1318999999999999 | -3.7109999999999999 | -4.1301999999999994 | -4.7089999999999996 | -5.1298999999999992
Linearity deviations % | 0.17145 | 0.074395 | -0.12116 | -0.075611 | -0.16884 | -0.094183 | -0.18775 | 0.014386 | -0.11937 | 0.94957 | -0.44669 | 0.058742 | -0.28295 | -0.059942 | -0.16297 | 0.010624 | -0.058402 | 0.029833 | 0.016461 | 0.08396

## Slide 50
265 MHz clk
Chart: