## Slide 1: Noise studies pre-amp
Breve descrizione del setup
Descrizione misure
Risultati

## Slide 2
Setup to measure the noise
Bacteries turn out to be the most
silent source of power.
To measure 1/f noise we need to ampl

## Slide 3
Setup
E5052b
Pre-amp and faraday cage

## Slide 4: Pre-amp layout
The pre-amp has to be loaded with different capacitance mimicking the effect of different detectors (APD about 250 pF).

## Slide 5: Example output
1/f E5052Bnoise floor
AC decoupling input filter low
Frequency cut
High frequency cut
due to RC in feedback
We need to amplify the signal
To measure 1/f pre-amp noise figure

## Slide 6: Output (after 40 Db amplification
1/f corner noise at 100 Hz

## Slide 7: What do we see on the scope?
RMS about 1.4 mV

## Slide 8: What does it happen when we load the pre-amp input?

## Slide 9
Digital scope mesurement
1/f noise spectrum after 40 dB amlification
Summary

## Slide 10: Problem: how to compare scope measurement with spectrum analyser ones?
First of all: scope measurement are «integrated» in the all frequency
range.
Noise power distribution affects a broader frequency range.
Nevertheless if we integrate the noise spectrum in the frequency range the amplifier normaly works we should find similar values.

## Slide 11
V/
Noise distribution
Noise spectra studies
Here we completely drop the 1/f contribution from
data. We’ll integrate every spectrum in different
Range from 1kHz to 1 MHz and compare with the
scope measured RMS.

## Slide 12: Integral noise R.M.S.
Chart: 
| 5 | 10 | 100 | 220 | 1000
1kHz cut | 1.06 | 1.09 | 1.51 | 1.75 | 2.1
25 kHz cut | 1.03 | 1.07 | 1.49 | 1.73 | 2.06
250 kHz cut | 0.85 | 0.88 | 1.31 | 1.54 | 1.89
1 MHz cut | 0.61 | 0.64 | 1 | 1.3 | 1.57
Scope Measures | 0.68 | 1.1 | 1.4 | 1.6

## Slide 13: Conclusion
Noise spectra have been measured in all the interesting frequency range.
We now know the noise power distribution of the amplifier and its shape as a function of the «load».
There is agreement between scope meassured integral noise distribution and it’s «integral» as derived from the measured distribution.