## Slide 1
Claudia Polito
PhD Student - Morphogenesis and Tissue Engineering– Curriculum CELL and TISSUE BIOPHYSICS
Sapienza University of Rome, Italy
DUAL MODALITY IMAGING WITH A ULTRASOUND-GAMMA PROBE FOR ONCOLOGY

## Slide 2
1/15
STATE of THE ART
Introduction
Equipment and Methods
Results
Conclusion
State of the Art
Importance of Multi Modality Imaging
Why UltraSounds?
Background (ECORAD Project)
MOLECULAR
IMAGING
Modern tools of molecular and cell biology
Advanced technology for non-invasive imaging
These approaches should help to lead better methods for studying biological processes, diagnosing and managing diseases as well as in the translation of preclinical research to clinical applications
US
CT
MRI
PET
SPECT

## Slide 3
IMPORTANCE OF MULTI MODALITY IMAGING
Introduction
Equipment and Methods
Results
Conclusion
State of the Art
Importance of Multi Modality Imaging
Why UltraSounds?
Background (ECORAD Project)
Morphological Imaging
Functional Imaging
2/15
claudia.polito@uniroma1.it
look inside the body with amazing accuracy
no trauma and risk related to the exploratory surgery
define with high precision the anatomical details.
reveal physiological activities within a certain tissue or organ employing tracers and reflecting their spatial distribution within the body.

## Slide 4
IMPORTANCE OF MULTI MODALITY IMAGING
Introduction
Equipment and Methods
Results
Conclusion
State of the Art
Importance of Multi Modality Imaging
Why UltraSounds?
Background (ECORAD Project)
Morphological Imaging
Functional Imaging
2/15
claudia.polito@uniroma1.it
The great size of these scanners does not satisfy the growing request of handheld detectors or dedicated devices (es. preclinical trials or small animal imaging).

## Slide 5
WHY ULTRASOUNDS?
Introduction
Equipment and Methods
Results
Conclusion
State of the Art
Importance of Multi Modality Imaging
Why UltraSound?
Background (ECORAD Project)
Real time acquisition
High Spatial Resolution
Moderate costs
Absence of dose to the patient
Possibility to realize handheld detectors
Ultrasound
Scintigraphy
Scintigraphy (High Resolution Camera)
Example of suspected breast lesion
An accurate diagnosis is necessary to define an efficient therapeutic strategy
3/15
claudia.polito@uniroma1.it

## Slide 6
BACKGROUND (ECORAD PROJECT)
Introduction
Equipment and Methods
Results
Conclusion
State of the Art
Importance of Multi Modality Imaging
Why UltraSound?
Background (ECORAD Project)
4/15
claudia.polito@uniroma1.it
Small FoV gamma camera based on LaBr3:Ce scintillation crystal
US probe lodged on a shifting bad moving across the gamma device FoV (the arrow indicates the direction of the US probe movement on the gamma detector FoV)
Probe initially developed under INFN ‘ECORAD’ Project.
Patent AP/A17495
International Extension
PCT/IT2015/000130

## Slide 7
BACKGROUND (ECORAD PROJECT)
Introduction
Equipment and Methods
Results
Conclusion
State of the Art
Importance of Multi Modality Imaging
Why UltraSound?
Background (ECORAD Project)
4/15
claudia.polito@uniroma1.it
Quail egg
“Biological Experiment”

## Slide 8
INTEGRATED PROBE
Introduction
Equipment and Methods
Results
Conclusion
Integrated Probe
US Probe
Gamma Probe
US-Gamma Probe Features
5/15
claudia.polito@uniroma1.it
A small size device can easily handheld by the medical staff for clinical and surgical applications, providing a very effective localization method!
APPLICATIONS
Small organs located at short depths
(as thyroid)
*Pani et al., “Integrated ultrasound and gamma imaging probe for medical diagnosis” (2016). JOURNAL OF INSTRUMENTATION, doi: 10.1088/1748-0221/11/03/C03037.

## Slide 9
US PROBE
Introduction
Equipment and Methods
Results
Conclusion
Integrated Probe
US Probe
Gamma Probe
US-Gamma Probe Features
6/15
Operating Frequency:7.5-10 MHz
The thickness of Back-Suppressor is minimized (about 6 mm) to reduce the gamma attenuation.
Size 70x10x12 mm x mm x mm
Cables system designed to reduce obstruction, realized with 200 shielded cables.
Y Axix
X Axix
PMT
Slit
Collimator
crystal
Ultrasound
Probe
CUSTOMIZED ULTRASOUND LINEAR PROBE B-MODE
(192 piezoelectric crystals)
claudia.polito@uniroma1.it

## Slide 10
GAMMA PROBE
Introduction
Equipment and Methods
Results
Conclusion
Integrated Probe
US Probe
Gamma Probe
US-Gamma Probe Features
7/15
2 PSPMT MULTI ANODE R8900-M16 Hamamatsu (1.5mm glass window, number of anodes=4x8 array, 6.08mm pitch)
Readout Electronics (32 independent channels)
Planar Scintillation crystal: NaI:Tl (58x32x4.2 mm3), glass window thickness 2 mm.
≈3mm
Dead zone
claudia.polito@uniroma1.it
Slit Collimator (0.2 mm septum, 1.5 mm pitch)
Overall dimensions chosen in order to optimize SR and Position Linearity and to maximize the efficiency performing a gamma image in few seconds!

## Slide 11
US-GAMMA PROBE FEATURES (I)
Introduction
Equipment and Methods
Results
Conclusion
Integrated Probe
US Probe
Gamma Probe
US-Gamma Probe Features
8/15
claudia.polito@uniroma1.it
Orthogonal planes (gamma image profile overlapped on the structure highlighted by US)
Slit Efficiency is 10 times more than one from general purpose parallel hole at 5 cm SCD
The US and gamma images have to be compared on the same spatial resolution and position linearity for the right overlap

## Slide 12
US-GAMMA PROBE FEATURES (I)
Introduction
Equipment and Methods
Results
Conclusion
Integrated Probe
US Probe
Gamma Probe
US-Gamma Probe Features
8/15
claudia.polito@uniroma1.it
Orthogonal planes (gamma image profile overlapped on the structure highlighted by US)
Slit Efficiency is 10 times more than one from general purpose parallel hole at 5 cm SCD
The US and gamma images have to be compared on the same spatial resolution and position linearity for the right overlap

## Slide 13
US-GAMMA PROBE FEATURES (I)
Introduction
Equipment and Methods
Results
Conclusion
Integrated Probe
US Probe
Gamma Probe
US-Gamma Probe Features
8/15
claudia.polito@uniroma1.it
Orthogonal planes (gamma image profile overlapped on the structure highlighted by US)
Slit Efficiency is 10 times more than one from general purpose parallel hole at 5 cm SCD
The US and gamma images have to be compared on the same spatial resolution and position linearity for the right overlap

## Slide 14
US-GAMMA PROBE FEATURES (II)
Introduction
Equipment and Methods
Results
Conclusion
Integrated Probe
US Probe
Gamma Probe
US-Gamma Probe Features
9/15
Attenuation evaluated with flood field = 0,55
0.4 mm Tc99m Collimated Source
PMT
PMT
US probe
0mm
39mm
6
12
18
24
31.5
Counts
Channels
with US probe
without US probe
Position (mm) | 0 | 6 | 12 | 18 | 24 | 31.5 | 39
Probe/ no probe (%) | 0.60 | 0.58 | 0.62 | 0.59 | 0.59 | 0.59 | 0.73
claudia.polito@uniroma1.it

## Slide 15
INTEGRATED IMAGES (I)
Introduction
Equipment and Methods
Results
Conclusion
Integrated Images (I)
Integrated Images (II)
10/15
Chart: Image Pixel
US
Gamma
A
B
C
Spot | SR UltraSound (mm) | SR Gamma (mm) | Position US (mm) | Position Gamma (mm)
A | 1.6 | 1.6 | 20.5 | 19.7
B | 1.8 | 1.4 | 31.6 | 32.2
C | 1.8 | 1.8 | 40.9 | 40.5
claudia.polito@uniroma1.it

## Slide 16
INTEGRATED IMAGES (II)
Introduction
Equipment and Methods
Results
Conclusion
Integrated Images (I)
Integrated Images (II)
11/15
FOLEY (500 μCi of Tc-99m in 1c.c.)
Size ch. 2.0 mm
Length 31 cm
1.5 ml
7 mm
2 cm
Chart: pixel
US
Gamma
FWHM
SR UltraSound (mm) | SR Gamma (mm) | Position US (mm) | Position Gamma (mm)
6 | 6.4 | 3 | 2.8
claudia.polito@uniroma1.it

## Slide 17
CONCLUSION
Introduction
Equipment and Methods
Results
Conclusion
12/15
Conclusion
Near Terms Objective
claudia.polito@uniroma1.it
Compact design
Accurate diagnosis and possible applications for moving organs
Simultaneous acquisition in short times
Perfectly superimposed information
(both in position and in resolution)
Integration of US and Gamma devices allows an excellent co-registration of images providing accurate diagnosis

## Slide 18
NEAR TERM OBJECTIVE
Introduction
Equipment and Methods
Results
Conclusion
13/15
Conclusion
Near Terms Objective
Probe Miniaturization
claudia.polito@uniroma1.it
8 mm
SiPM
Active area for channel: 3mm
50μm pitch
Gain: 1.2 * 106
PDE 35%

## Slide 19
NEAR TERM OBJECTIVE
Introduction
Equipment and Methods
Results
Conclusion
14/15
Conclusion
Near Terms Objective
Probe Miniaturization
claudia.polito@uniroma1.it
High gain
Linearity
High photo-detection efficiency
Low dark count level
Excellent timing resolution
Lower voltage: 30V
Small Size Readout Electronics

## Slide 20
NEAR TERM OBJECTIVE
Introduction
Equipment and Methods
Results
Conclusion
15/15
Conclusion
Near Terms Objective
Probe Miniaturization
claudia.polito@uniroma1.it
Fast decay time
Great Resolution
Non hygroscopic
Scintillation Crystals
 | NaI:Tl | CRY019
Density(g/cm3) | 3,67 | 7,1
Photon Yield (kph/MeV) | 40 | 28
Decay Time (ns) | 230 | 41
Emission (nm) | 415 | 420
Index of Refraction | 1,85 | 1,81
Energy Resolution@622 keV (%) | 5,6 | 8,5 (661 kev)
Hygroscopic | yes | no

## Slide 21
THANK YOU FOR LISTENING
Claudia Polito
XIV Seminar on Software for Nuclear, Subnuclear and Applied Physics, Alghero 04 – 09 June 2017