## Slide 1
Strange and non-strange mesons induced processes studies at DAFNE, J-PARC and RIKEN:
present and future
Istituto Nazionale Fisica Nucleare - Laboratori Nazionali di Frascati
Frascati, 10-11 / 07 / 2017

## Slide 2
Spectroscopic characterization of Silicon Drift Detectors for SIDDHARTA-2 experiment
Marco Miliucci
On behalf of SIDDHARTA-2 collaboration
Marco.Miliucci@lnf.infn.it
Frascati, 11/07/2017

## Slide 3
Outline
  SIDDHARTA
  DAQ
  Spectroscopic characterization
  Conclusions
Frascati, 11/07/2017

## Slide 4
Outline
  SIDDHARTA
  DAQ
  Spectroscopic Characterization
  Conclusions
Frascati, 11/07/2017

## Slide 5: Silicon Drift Detectors for Hadronic Atom Research by Timing Application
Frascati, 11/07/2017

## Slide 6: Silicon Drift Detectors for Hadronic Atom Research by Timing Application
Frascati, 11/07/2017
Region Of Interest
4500 – 12000 eV
Monte Carlo simulation of K- d for SIDDHARTA-2 for 800 pb-1

## Slide 7
Outline
  SIDDHARTA
  DAQ
  Spectroscopic Characterization
  Conclusions
Frascati, 11/07/2017

## Slide 8
P. Rehak at al, Nucl. Instr. Meth. Phys. Res. A235 (1985) 224-234
Frascati, 11/07/2017
Silicon Drift Detectors
IEEE transactions on nuclear science, vol. 47, no. 4, august 2000 1691
RN
Back
Anode

## Slide 9
Frascati, 11/07/2017
Silicon Drift Detectors

## Slide 10
Frascati, 11/07/2017
Electronic Chain

## Slide 11
Frascati, 11/07/2017
Experimental setup
CAMERA
P < 10-5 mbar
T = R.T. -> 100 K

## Slide 12
Outline
  SIDDHARTA
  DAQ
  Spectroscopic characterization
  Conclusions
Frascati, 11/07/2017

## Slide 13
Frascati, 11/07/2017
X RAY spectrum
Foto targhetta
Ti Fe Cu Br

## Slide 14
Frascati, 11/07/2017
X RAY spectrum
T SDDs: 120 K
Pressure: 8.2 * 10-6 mbar
R1 Voltage: -15,2 V
High Voltage: -150 V
Xray Tube: 23 KV @ 10μA
Mean Rate on SDD: 1,12 KHz
Ti Kα
Cu Kα
Fe Kα
Ti Kβ
Cu Kβ
Fe Kβ
Br Kα
Br Kβ
ROI
Energy (eV)

## Slide 15
Frascati, 11/07/2017
Linearity
SDD_1
SDD_7
Energy (adc)
Energy (eV)
Calibration
Calibration
y = p1*x + p0
p0 = 1592.4 ± 0.2 [adc]
p1 = 0.59472 ± 0.00004 [adc/eV]

## Slide 16
Frascati, 11/07/2017
Linearity vs HV
SDD_1
SDD_7

## Slide 17
Frascati, 11/07/2017
Temperature Stability
SDD_1
SDD_7
Channel | ΔT SDD [°C] | mMAX [adc/eV] | mMIN [adc/eV] | (Δm/m )/ΔT [°C-1]
1 | 9,6 | 1,68301 ± 0,00002 | 1,68225 ± 0,00003 | 5*10-5
7 | 9,6 | 1,6212 ± 0,0002 | 1,62047 ± 0,00008 | 7*10-5

## Slide 18
Frascati, 11/07/2017
Resolution vs HV
FWHM Mn_CH1 = 161 ± 1 eV
FWHM Mn_CH7 = 158 ± 1 eV

## Slide 19
Frascati, 11/07/2017
Resolution
@ 120K
Energy (adc)
Energy (adc)
Iron Kα,Kβ
Titanium Kα,Kβ
Ti FWHM = 116.8 ± 0.3 eV
Fe FWHM = 133.9 ± 0.4 eV
Mn FWHM = 128 ± 3 eV

## Slide 20
Frascati, 11/07/2017
Resolution vs Temperature
SDD_4
SDD_2
SDD_3
*Paper in preparation

## Slide 21
Frascati, 11/07/2017
Timing at BTF
t0
e-
t1
LINAC
e-, e+ beam
Drift time = Te = t1-t0
Energy Beam: 450 MeV
Multeplicity e-= 200-300
Multeplicity e+= 30-40
BEAM CONVERTER
anode
Time Resolution = ΔTen

## Slide 22
Frascati, 11/07/2017
Drift Time vs Temperature
*Paper in preparation
T SDDs: 110 – 210 K
Pressure: 10-4 mbar
First Ring Voltage: -15,2 V
Last Ring Voltage: -150 V
Scintillator + FADC

## Slide 23
Outline
  SIDDHARTA
  DAQ
  Spectroscopic Characterization
  Conclusions
Frascati, 11/07/2017

## Slide 24
Frascati, 11/07/2017
NUOVA CAMERA
NEW SETUP
Dr. J. Zmeskal, Stefan Meyer Institute, Vienna

## Slide 25
Frascati, 11/07/2017
NEW SETUP
Dr. J. Zmeskal, Stefan Meyer Institute, Vienna

## Slide 26
Frascati, 11/07/2017
Electronic conversion : Δ = 10-4
              HV independent (100 – 170) V
Thermal Stability (SDD+CUBE) : Δ < 10-4 °C-1
Resolution: 128 ± 3 eV (@ 120 K)
            HV independent (100 – 170) V
            T dependent (further investigation)
Drift time : 372 ± 8 ns (@ 110 K)
            T dependent (further investigation)
Conclusions

## Slide 27
THANK YOU
Marco.Miliucci@lnf.infn.it