## Slide 1: Recent results of cavity infusion at DESY
Outline
Nitrogen infusion
Cavity & sample results
Cross-section of samples
Analysis of precipitates size
Christopher Bate
on behalf of DESY SRF Team
TTC Meeting @ Milan
2018/02/08

## Slide 2: Nitrogen infusion – DESY furnace
800C – 3h - vacuum
Cool down to 120C
48h @ 120C + 25 mTorr N2
Cool down
Ofen(Bem. Alter Ofen mit.. Wesentliche Punkte (druck,Ölpumpen)),
Druck/T Kurve von Joern
2 turbo molecular pumps (Varian, 6000 l/s each , max. intake pressure ~ 10-2 mbar)
Tmax ~ 1100°C
V = 1800 x 625 x 660 mm³
EPDM sealed
Roughing pumps (oil sealed):
  Rotary vane & roots pump

## Slide 3: Cavity assembly in furnace

## Slide 4: First infusion try at DESY – Q vs E and sample SEM
Infusion recipe | Measured
800°C @ 3h, p10-6 mbar, cool down | 800°C @ 2h, p10-5 mbar, cool down (slower?)
120°C @ 48h | 120°C @ 48h
pN210-2 mbar | pN210-4 mbar
Infusion recipe | Measured
120°C @ 48h | 120°C @ 48h
Precipitates on sample (BCP)
Although very few
1DE18
E3
@ 2K

## Slide 5: Second run - bakeout same as infusion w/o Nitrogen
Furnace parameters
800°C @ 2h, p10-5 mbar
120°C @ 48h, p10-7 mbar
Furnace parameters
1DE17
More precipitates than before (BCP sample)
But same Q-Slope
No correlation between precipitates and cavity performance?
D4
@ 2K

## Slide 6: Third try w/o Nitrogen
Furnace parameters
800°C @ 3h, p7.510-6 mbar
120°C @ 48h, p10-7 mbar
Furnace parameters
D5
1DE16
Still observing precipitates on sample (BCP)
More or less than before?
What are they?
@ 2K

## Slide 7: Cross-section EDX map
Possible phase: Nb2C
D4
(1DE17)
N
D5
(1DE16)
C
Pt
SE

## Slide 8: Samples - precipitates size comparison
D5 (1DE16) :
D4 (1DE17):
One sample shows deeper and bigger precipitates

## Slide 9: Samples - precipitates size statistics
Chart: Size in µm²
| 3.0000000000000001E-3 | 4.7647058823529417E-3 | 6.5294117647058825E-3 | 8.2941176470588233E-3 | 1.0058823529411766E-2 | 1.1823529411764705E-2 | 1.3588235294117647E-2 | 1.535294117647059E-2 | 1.7117647058823529E-2 | 1.8882352941176472E-2 | 2.0647058823529411E-2 | 2.2411764705882353E-2 | 2.4176470588235296E-2 | 2.5941176470588235E-2 | 2.7705882352941177E-2 | 2.9470588235294116E-2 | 3.1235294117647059E-2
1DE17 (D4) | 40 | 42 | 84 | 42 | 28 | 12 | 17 | 13 | 6 | 3 | 4 | 2 | 1 | 0 | 0 | 1 | 1
1DE16 (D5) | 11 | 138 | 42 | 21 | 14 | 6 | 3 | 2 | 3 | 2 | 0 | 1 | 2 | 1 | 0
Average size:
Count: 247
Total area:
Average size:
Count: 299
Total area:
1DE17
1DE16
D4
D5

## Slide 10: Samples - precipitates size statistics
Average size:
Average size:
D4
D5
1DE17
1DE16

## Slide 11: Second run of 1DE16 (w/o Nitrogen)
Furnace parameters
800°C @ 3h, p5.610-6 mbar
120°C @ 48h, p10-7 mbar
Furnace parameters
D6
1DE16
Observing precipitates on sample (BCP)
Only on certain grains – grow only on certain crystal orientation!
Dense and numerous
@ 2K

## Slide 12: Fifth run with another single cell cavity (w/o Nitrogen)
Furnace parameters
800°C @ 3h, p510-6 mbar
120°C @ 48h, p10-7 mbar
Furnace parameters
Electropolished fine-grain sample
No precipitates at all!
F1
F1
1DE09
@ 2K

## Slide 13: Cavity and sample summary
Cavity (all FG) | Material | Surface chemistry | Furnace cycle | p_max [mbar] @ 800°C |  | Sample | Material/ Surface chemistry | Precipitates | Observations
1DE18 | Ningxia 2005 | 80µm BCP 800°C 100µm EP 130°C | 2h@800°C 48h@120°C w N2 (problems with N2 regulation) |  |  | E3 | LG Heraeus BCP 20 µm |  | Very few
1DE17 | Ningxia 2005 | 80µm BCP 800°C 100µm EP 130°C | 2h@800°C 48h@120°C w/o N2 |  |  | D4 | LG Heraeus BCP 10 µm |  | Whole surface covered Average size:
1DE16 (1st run) | Heraeus Plansee 2005 | 80µm BCP 800°C 100µm EP 130°C | 3h@800°C 48h@120°C w/o N2 |  |  | D5 | LG Heraeus BCP 10 µm |  | Whole surface covered Average size:
1DE16 (2nd run) | “ “ | 80µm BCP 800°C 100µm EP 130°C | 17,5h@300°C 3h@800°C 48h@120°C w/o N2 |  |  | D6 | LG Heraeus BCP 10 µm |  | Very numerous and dense Only at certain crystal orientations
1DE09 | Heraeus 1997 | 119µm BCP 800°C 100µm EP HPR 138°C/48h dry-ice cleaning | 19,5h@300°C 3h@800°C 48h@120°C w/o N2 |  |  | F1 | FG Heraeus EP 200 µm | X | 
Cavity (all FG) | Material | Surface chemistry | Furnace cycle | p_max [mbar] @ 800°C |  | Sample | Material/ Surface chemistry | Precipitates | Observations
1DE18 | Ningxia 2005 | 80µm BCP 800°C 100µm EP 130°C | 2h@800°C 48h@120°C w N2 (problems with N2 regulation) |  |  | E3 | LG Heraeus BCP 20 µm |  | Very few
1DE17 | Ningxia 2005 | 80µm BCP 800°C 100µm EP 130°C | 2h@800°C 48h@120°C w/o N2 |  |  | D4 | LG Heraeus BCP 10 µm |  | 
1DE16 (1st run) | Heraeus Plansee 2005 | 80µm BCP 800°C 100µm EP 130°C | 3h@800°C 48h@120°C w/o N2 |  |  | D5 | LG Heraeus BCP 10 µm |  | 
1DE16 (2nd run) | “ “ | 80µm BCP 800°C 100µm EP 130°C | 17,5h@300°C 3h@800°C 48h@120°C w/o N2 |  |  | D6 | LG Heraeus BCP 10 µm |  | Very numerous and dense Only at certain crystal orientations
1DE09 | Heraeus 1997 | 119µm BCP 800°C 100µm EP HPR 138°C/48h dry-ice cleaning | 19,5h@300°C 3h@800°C 48h@120°C w/o N2 |  |  | F1 | FG Heraeus EP 200 µm | X | 

## Slide 14: DESY cavity - infusion run at Fermilab
Electropolished fine-grain sample
Lots of precipitates!
But: Only on certain grains
Preliminary
1DE03
@ 2K

## Slide 15: Summary
Precipitates
  Grow on certain crystal orientation
  Size may correlate with cavity performance
  Possible phase: Nb2C
Vacuum got better at every run
More important: partial pressures (analyze RGA's)
  Cut 1DE16 for direct surface analysis
  Compare samples
  BCP and EP
  LG & FG
  Material (3 EXFEL suppliers)
Future options
  Installing oil free pumps
  Alternative furnace