## Slide 1
Paolo Prati
INFN & University of Genoa
LNGS, Oct. 2nd 2017
prati@ge.infn.it
Neutron production and monitoring at the new LUNA-MV facility

## Slide 2
LUNA-MV
LUNA MV will be installed in the North part of Hall B of LNGS
See LUNA-MV proposal
a and C induced reactions
≈ 30 m

## Slide 3: Layout of the new LUNA-MV facility
27 m
12.5 m
5.5 m

## Slide 4: The accelerator and the neutron shielding
1H+ (TV: 0.3 – 0.5 MV): 500 μA
He
H
C
1H+ (TV: 0.5 – 3.5 MV): 1000 μA
4He+ (TV: 0.3 – 0.5 MV): 300 μA
4He+ (TV: 0.5 – 3.5 MV): 500 μA
12C+ (TV: 0.3 – 0.5 MV): 100 μA
12C+ (TV: 0.5 – 3.5 MV): 150 μA
12C++ (TV: 0.5 – 3.5 MV): 100 μA
inline Cockcroft Walton accelerator
TERMINAL VOLTAGE: 0.2 – 3.5 MV
Precision of terminal voltage reading: 350 V
Beam energy reproducibility: 0.01% TV
Beam energy stability: 0.001% TV / h
Beam current stability: < 5% / h
A key-point: at low energy LUNA-MV must deliver very intense beams to make the nuclear cross sections measurable…however in the highest part of the energy range, in particular with proton beam, experiments can be performed by beam of much, much lower intensity !
Nevertheless we are committed to guarantee the full operation beyond the peculiar NA needs.

## Slide 5
The “history” of neutron simulation for LUNA-MV
Maximum neutron production calculated in 2014 on the basis of the long-term scientific program:
13C(a,n)16O; Q=2.216 MeV, En ~ 3 MeV @ Ea = 1 MeV
22Ne(a,n)25Mg, Q= - 0.478 MeV, En ~ 0.368 MeV @ Ea = 1 MeV
13C(a,n)16O as 10-5 contaminant of 12C(a,g)16O
Encm ~ 4.89 MeV @ Ea = 3.5 MeV; production rate: 2 103 n/s @ Ia = 200 mA
Simulation input/Design benchmark
Point-like n-source, 2 103 n/s , Enlab = 5.6 MeV

## Slide 6
INFN-FISMEL: MCNP independent simulations (2015-16)
Run with the «source» in the realistic position,
80 cm thick concrete shieding in each direction
R
U
Surface | MCNP (10-7 n/cm2 s) | GEANT4 (10-7 n/cm2 s)
Left | 5.69888 ± 0.00007 | 12.3 ± 0.1
Upper | 2.58450 ± 0.00003 | 5.71 ± 0.05
Lower | 0.67484 ± 0.00002 | 1.53 ± 0.03
Right | 0.29462 ± 0.00002 | 0.64 ± 0.02
Roof | 0.68379 ± 0.00001 | 1.96 ± 0.02
Total (weighted mean) | 1.38535 ± 0.00001 | 3.40 ± 0.02
Xenon (~10 m)
(in Hall B)
LUNA

## Slide 7
INFN-FISMEL: MCNP independent simulations
E* dN/dE
E (MeV)
Fraction of fast neutrons
( En > 1 MeV)
< 1%

## Slide 8
INFN-FISMEL: conclusions
The value obtained with the simulations performed in GEANT-4 by the LUNA collaboration is (3.400.02) 10-7 n/cm2/s to be compared with the value obtained with the MCNP6 code (1.385350.00001) 10-7 n/cm2/s. Both values are lower than the measured value of Arneodo et al., (of a factor 1.2 and 3, respectively). Therefore, the installation of the LUNA-MV accelerator in the Hall C would determine a neutron flux outside the shielding slightly below (in the case of GEANT-4 simulation) or well below (in the case of the MCNP6 simulation) than the natural neutron flux.
MCNP6 has been selected as benchmark Monte Carlo code since it is worldwide considered as one of the most reliable codes for projecting a shielding. It has also been widely validated for the neutron transportation.
The GEANT-4 results for the neutron flux are higher of a factor ≈ 2.2. This can be due to several reasons and would deserve to be deeply investigated. However, the results obtained with MCNP6 are more accurate since their statistical uncertainty is lower than that of GEANT4 results by several orders of magnitude.
February, 23th 2016. Adolfo Esposito, Head of INFN- FISMEL

## Slide 9
Final remarks
Wulandari et al., 2004
Belli
1989
Arneodo
1999
Rindi
1988
Bellotti 1985
Cribier 1995
Aleksan 1989
≈ 3
En < 1 keV
≈ 0.4
1 <En< 10 MeV
10-6 cm-2 s-1

## Slide 10
Other possible sources for neutron emission
Besides those emitted from the target, LUNA-MV could produce neutrons as result of the interaction of the beam with the components of the accelerator and/or of the vacuum line. This could happen in particular with the proton beam at Ebeam > ~1.5 MeV (not negligible at Ebeam > ~2 MeV).
Event | Cause | Frequency | Safeguards to keep Rn < 2 103 s-1
Beam on Faraday cup (beam stopper) | Routine operation | daily | Tantalum coating (passive) Automatic shutdown triggered by the neutron monitors (active)
Beam on slits, collimators, walls | Beam tuning, components failure, operator mistake | daily | Beam Wizard (active) Tantalum coating (passive) Automatic shutdown triggered by the neutron monitors (active)
Vacuum loss | Pumps failure, leaks, breaks | Rare ( y-1) | Automatic shutdown triggered by the accelerator diagnostic (active)
Possible Scenarios

## Slide 11
Passive safeguards: Ta (or Au) coating
At Ep = 3.5 MeV: range of proton in metals ranges form 50 to 100 mm.
The contract between INFN and HVEE foresees the coating by Ta of any part of the accelerator and transport line which could be hit by the beam. Tantalum (or Au…alternative solution for parts of the set-up designed directly by LUNA) is an inert material, however internal contaminants must be controlled.
Ep =3.5 MeV, Ip = 1 mA, 1 ppm impurity in Ta
59Co
55Mn
51V
93Nb

## Slide 12
Clean Ta foils are available in the market
LUNA/HVEE R&D work… in progress
First check: in-house acceptance tests. Spring 2018
Passive safeguards

## Slide 13
Neutron rate due to Ta contaminants: 3.5 vs. 3 MeV (Ep)

## Slide 14
Active safeguards: neutron monitors
Commercial neutron (and X/g) monitors will be installed inside the accelerator room and directly connected/controlled to/by the LNGS safety system. (ref. Authorization procedure)

## Slide 15
Berthold LB6411
Ambient Dose Equivalent ICRP60
Berthold LB6414
Neutron Survey Meter
ELSE Lupin 5401 He3-NP
REM Counter
Active safeguards: neutron monitors

## Slide 16
Model | Berthold LB6411 Ambient Dose ICRP60 | Berthold LB 6414 Survey Meter | ELSE LUPIN 5401 HE3-NP REM Counter
External diameter [mm] | 250 | 310x180x130 | 250
Counter | 3He@2 atm cylindrical | 3He@2 atm cylindrical | 3He@2 (4, 6 e 8) atm spherical
Range [nSv/h] | > 30 |  | > 10
Fluence response [cm2] | 1.09 | 10.7 Am-Be 26.4 fission | 
H*(10) response [counts/nSv] | 2.83 | 68, 27 Am-Be | 4 @2 atm; 6.4 @ 4 atm 8.8 @ 8 atm
Gamma sensitivity [µSv/h] | < 40 @ 10 mSv/h, 662 keV |  | < 0.5 @ 50 mSv/h, 662 keV
h ( 3 < En < 6 MeV) | 3.47 * 10-5 |  | 0.54 ÷ 1.20 * 10-4
Active safeguards: neutron monitors
Berthold LB6411 and ELSE 5401 has been/will be tested underground: background meas./expect. < 1 c.p.m.

## Slide 17
Active safeguards: reacting time shutdown
BERTHOLD
ELSE
@ 2000 n/s  40 s (ELSE, 8 atm)
Safety thres. of 7 counts in ≈ 30 s

## Slide 18
Conclusions and next steps
The LUNA-MV shielding has been designed (and validated) assuming a max. neutron emission of 2 10-3 s-1 , En = 5.6 MeV
Proper coating of the beam line components that will/could be hit by the beam (in particular: Faraday cups and beam stoppers) will prevent uncontrolled n emissions. Work is in progress to select the cleanest coating materials.
Neutron monitors will be placed near (≈ 50 cm) the “hot spots” to guarantee the accelerator shutdown in less than 40 s in case the emission limit is overcome
As final measure, limitations to the maximum beam current and energy (in particular for protons) could be implemented
LUNA suggests to install a LNGS-controlled, long-term, neutron monitor in Hall B, in the area between LUNA-MV and Xenon

## Slide 19
The LUNA collaboration
G.F. Ciani*, L. Di Paolo, A. Formicola, I. Kochanek, M. Junker | INFN LNGS /*GSSI, Italy
D. Bemmerer, M. Takacs, T. Szucs | HZDR Dresden, Germany
C. Broggini, A. Caciolli, R. Depalo, P. Marigo, R. Menegazzo, D. Piatti | Università di Padova and INFN Padova, Italy
C. Gustavino | INFN Roma1, Italy
Z. Elekes, Zs. Fülöp, Gy. Gyurky| MTA-ATOMKI Debrecen, Hungary
M. Lugaro | Monarch University Budapest, Hungary
O. Straniero | INAF Osservatorio Astronomico di Collurania, Teramo, Italy
F. Cavanna, P. Corvisiero, F. Ferraro, P. Prati, S. Zavatarelli | Università di Genova and INFN Genova, Italy
A. Guglielmetti| Università di Milano and INFN Milano, Italy
A. Best, A. Di Leva, G. Imbriani, | Università di Napoli and INFN Napoli, Italy
G. Gervino | Università di Torino and INFN Torino, Italy
M. Aliotta, C. Bruno, T. Chillery, T. Davinson | University of Edinburgh, United Kingdom
G. D’Erasmo, E.M. Fiore, V. Mossa, F. Pantaleo, V. Paticchio, R. Perrino, L. Schiavulli, A. Valentini| Università di Bari and INFN Bari, Italy
.

## Slide 20
Counting rate vs. En

## Slide 21
Bonner Sphere spectrometer
PTB (Physikalisch-Technische Bundesanstalt) property and certification
Neutron flux monitoring in Hall B
(slide shown at the last SC meeting)
A bck. run is going to start today
Data taking actually started Oct 20th 2016…still ongoing !
3He detector for thermal neutrons

## Slide 22
integrated neutron flux = (6.4 ± 1.0) 10-6 n/(cm2 s)
Am-Be ref. spectrum
Neutron flux monitoring in Hall B
(first 142 days)
Net counts in RoI = 157 ± 26
Mean neutron flux = (6.4 ± 2.5) 10-6 n/(cm2 s)
?

## Slide 23
UDO salt mine (490 m depth)
Felsenkeller (50 m depth)
Hall B vs. Reginatto et al., AIP proc. 1553, 77 (2013)
LNGS – Hall B
(preliminary)

## Slide 24
Recent achievements in the neutron bck. Monitoring: the new «LUNA» 3He tubes
Fn(therm)≈ 5 10-7 cm-2 s-1 (LUNA400 site)

## Slide 25
UDO salt mine (490 m depth)
Felsenkeller (50 m depth)
Hall B vs. Reginatto et al., AIP proc. 1553, 77 (2013)
LNGS – Hall B
(preliminary)
New LUNA 3He tubes
Thermal n. only !

## Slide 26
LUNA MV: detail of the shielding pass through
En ~ MeV
Trasm.~ 1/10 each turn
Solid angle fraction at the nearest pass through ≈ 4 10-4
Filled by polyethylene mspheres
Pass through dimension:
H = 20 cm, W = 40 cm
With 4 turns  ≈ 10-7 n/cm2 s through the «critical» pass-through

## Slide 27
LUNA-MV: details of the two sliding doors (SPES-like)