## Slide 1: Mark E. Bier*, Logan Plath*, Stephan Friedrich#, Mohammad Halim*, Francisco Ponce#, Jack Harris+, Jonathan Shulgach*, Robin Cantor++ *Department of Chemistry, Carnegie Mellon University, Pittsburgh, PA, USA #Lawrence Livermore National Laboratory, Livermore, CA, USA +XIA LLC, Hayward, CA, USA ++ STAR Cryoelectronics Inc., Santa Fe, NM, USA
Mass-to-charge and Energy Distributions of Electrosprayed and Matrix-assisted Laser Desorbed Ions Measured by Nb- and Ta- Superconducting Tunnel Junction Cryodetection Mass Spectrometry

## Slide 2: Outline
Introduction
  Historical account of cryodetector MS
  Advantages of using STJs in MS
Applications of STJ MS: Using MALDI TOF MS and ESI LIT MS
  Ultra-high m/z virus particles and gold nanoparticles
  Unique energy deposited measurements
    Energy measurements of metastable fragments, MS/MS
    Energy measurements for various ions and protein conformations.

## Slide 3: History of Cryogenic Detection MS
1996 - Twerenbold et al. (Neuchâtel U.) - First to implement STJ MS.
1996 - Frank (LNLL) / Benner (LBL) et al., – Fast Nb-STJs, analysis up to 66kDa.
1998 - Hilton, et al. (NIST) / Twerenbold et al. (Neuchâtel U.), Impact energy MS measurements using microcalorimeters.
2003 - Comet AG, (Switzerland) Macromizer™ - First commercial STJ MS.
2005-2019
  Zenobi et al, ETHZ - data on a commercial MS, von-WF.
  Ohkubo et al, AIST - proteins, small molecules, metastable ions, higher temperature cryodetectors, large array STJs
  Cotter et al, JHU & Caprioli et al., VU – proteins
  Bier et al., CMU– polymers, capsid, whole virion, ferritin, nanoparticles, first ESI LIT STJ/CEM MS & with Friedrich et al., LLNL – incorporation of Ta-STJs (STAR Cryoelectronics, Cantor), new XIA electronics.

## Slide 4: Cryodetector and MCP efficiency vs. Mass
The Cooper pairs break from thermal excitation or by the phonon energy available after particle collision.
∆ is the band gap parameter of the superconductor. (~meV)
The amount of energy required to break a Cooper pair is 2∆.
KEion >> 2∆ !
Tc (tunnel current) is proportional to the number of quasi-particles formed which is proportional to the collision energy.
+
15 Ǻ oxide barrier
S1
S2
Cooper pairs break resulting in quasi-particles
Tunnel current
+
Principles of ion detection using a super-conducting tunnel junction
0.35°K
KE
Superconducting Tunnel Junction (STJ) Cryodetection
+ B field
+
t (µs)
v
Twerenbold, D.; Gerber, D.; Gritti, D.; Gonin, Y.; Netuschill, A.; Rossel, F.; Schenker, D.; Vuilleumier, J-L. “Single Molecule Detector for Mass Spectrometry with Mass Independent Detection Efficiency”, Proteomics 2001, 1, p69.

## Slide 5: Ta-STJ MS

## Slide 6: What is the noise level Ta-STJs at 350mK?
Would the signals be fast enough?
What would be the maximum resolving power?
X-ray data for StarCryo Ta-STJs acquired at LLNL
400mK
STAR Cryoelectronics 36 Ta-STJs
Friedrich &
Ponce

## Slide 7: MALDI TOF Nb- or Ta-STJ MS
TOF length (linear mode): was 1.6m or 3.8m
Accelerating voltage up to 20,000 V
Delayed extraction linear resolution, R~ 500
16 element Nb-STJ cryodetector, dimension 2.2 x 2.2 mm, active surface 0.64 mm2
~ 4 to 16% of the ion beam!
32 Channel Nb or Ta-STJ detector intercepts ~30% ion beam
20 kV
Signal
32 channel
amplifier
MALDI plate
Nitrogen laser
337 nm, 3 ns pulse
Einzel lens
hn
grid
cryostat
4K
30K
298K
0.3K 3He
3 X 10-8 Torr
4He coldhead
XIA transient recorder
Mass spectrum
COMET AG MacromizerTM
32 element Ta-STJ cryodetector
Carnegie Mellon
+50
IR shields
deflector

## Slide 8: Ta-STJ Cryodetection MS of Large Protein Complexes
IgM
IgG1

## Slide 9: Virus particles Nb-STJ MS Th=Thomson=m/z

## Slide 10: TOP: For the first time, two ~5MDa intact whole virion particles from CPMV have been separated in a mass spectrometer and the individual charge states resolved and detected by STJ MS.
Macromolecular MALDI Nb-STJ Cryodetection MS of Virus Particles
Logan Plath
Intact Separation of Two Whole Virions of Cowpea Mosaic Virus (CPMV) ~ 5MDa
Dr. David Sipe
Jonathan Feldman
Cross-linked
Prof. Nicole Steinmetz
(CWRU)
Anna Czapar (CWRU)
Prof. Roger Hendrix (Pitt)
Prof. Bob Duda (Pitt)
54 nm
kTh
Energy
Capsid of Bacteriophage HK97 at 13MDa
Head II
+1
Notes: Nicole Steinmetz research is focused on modifying plant virus for other uses such as drug delivery. The Bier lab specializes is analyzing these heavy particles and we plan to characterizer her modified plant viruses now that we have shown that we can weigh them. The Bier Lab is seeking out biological discoveries in the study of virus now that we have proven that the analysis can be done. The bottom figure of HK97 show that we can analyze m/z 13 MDa after extending the mass range in software. The HK97 work is in collaboration with Prof Roger Hendrix Lab at the University of Pittsburgh.

## Slide 11: Synthetic gold nanoparticles Nb-STJ MS

## Slide 12
STJ response vs particle type: 5 nm Au particles vs sinapinic acid
Sipe, D.M.; Plath, L.D.; Aksenov, A.A.; Feldman, J.S.; Bier, M.E. Characterization of Mega-Dalton-Sized Nanoparticles by Superconducting Tunnel Junction Cryodetection Mass Spectrometry. ACS Nano. 2018, 12, 2591-2602.

## Slide 13: Energy deposition measurements Nb-STJ

## Slide 14
5% difference in energy deposited by adding one Au atom to angiotensin I
Relative Energy Deposited in Detector
Intensity
[M+H]+
0.607
[M-2H++Au+3]+
0.620
∆E
Energy Scatterplot
M = angiotensin
MW=1296.7 Da

## Slide 15
Increasing STJ Energy Response
Au10(S-C6H4-C4H9)10
Au36(S-C6H4-C4H9)24
Au133(S-C6H4-C4H9)52

## Slide 16: STJ energy response vs number of Au atoms
Angio-Au +, Aun(S-C6H4-C4H9)n+ and 5nm Au+

## Slide 17
STJ
STJ
STJ
STJ
IgM
apoferritin
Au
KE
KE
KE
holoferritin
KE
5 nm
12 nm
12 nm
Which complex will deposit the most energy into the STJ detector?
?

## Slide 18
kTh
A new MS based analytical tool!
Protein
kTh
kTh
5 nm Au
PbS quantum dot
Energy deposited into STJ detector

## Slide 19
STJ Energy Response of multiple samples (+1) vs MW
Curve drawn for clarity
kTh
Protein [M+H]+
Au/Ag
nanoparticle
KE = 20keV
Westmacott, G.; Zhong, F.; Frank, M.; Friedrich, S.; Labov, S.; Benner, W. RCMS 2000, 14(7), 600.

## Slide 20
Relative Energy Deposited in Detector
Intensity
1
Although similar ion size, shape and collisional cross-section, Apo- and HoloFt differ in mass and density. We observe a higher energy response for the Fe-core containing Ferritin.
0.2 0.4 0.6
[HoloFt]+
0.266
[ApoFt]+
0.260
+1
+2

## Slide 21: Metastable fragmentation MS/ MS mproduct= m*precursor (Eproduct/Eprecursor) Nb-STJ MS

## Slide 22
VDEFLECTOR = 0V
All particles
~13 Simultaneous MS/MS spectra based on metastable fragmentation
VDEFLECTOR = 1000V
Only neutral fragments
Au10(S-C6H4-C4H9)9+
Aun(S-C6H4-C4H9)n-1+ Adducts
Aun(S-C6H4-C4H9)n-1+ Fragments

## Slide 23
One MS/MS product scan using metastable fragmentation
Energy resolved metastable fragments of the Au10(TBBT)9+ precursor ion.
Precursor

## Slide 24: ESI Nb-STJ/CEM LIT MS
m/z and energy distributions for multimeric protein ions
energy deposition for different protein conformations

## Slide 25: Electrospray ionization Linear Ion Trap STJ MS
Frequency scanning LIT from 100kHz-5kHz incorporating dual detectors

## Slide 26: Native Protein carbonic anhydrase by ESI-LIT-STJ MS MW=29kDa
16 Nb-STJ
1 Nb-STJ
16 Nb-STJ

## Slide 27: Is there an energy deposition difference for a folded vs unfolded protein?
Native = folded
denatured = unfolded
+18
+18
BSA
MW 67kDa

## Slide 28: Native and Partially Unfolded Bovine Serum Albumin by ESI-LIT-STJ MS

## Slide 29: Relative Energy Deposited into a STJ Detector for Folded vs Unfolded BSA
Single pixel
Native (folded)
The preliminary results for the relative energy deposited into the STJ detector for the +18 charge state of native and partially unfolded BSA ion showed that the relative energy of native BSA was higher than the partially unfolded BSA ion. Ion mobility experiments are need to confirm this result.
Partially unfolded
s

## Slide 30
In memory of
Urs P. Schlunegger
September 27, 1934 - September 28, 2015

## Slide 31: Conclusions
First implementation of Ta-STJs for MS.
First ESI LIT STJ/CEM MS.
ESI-LIT-STJ data showed higher order oligomers including dimer, trimer, and tetramer ions by m/z and energy.
ESI-LIT-STJ data showed low energy ions which suggests that there is more information in this region.
Energy deposited differences for various ion types suggests a new analytical tool for ion content or structure. Preliminary results show a potential application to protein conformational analysis.

## Slide 32
Acknowledgments
Dr. Stephan Friedrich
LLNL
Mohammad Halim
Carnegie Mellon
Prof. Abdil Ozdemir
Sakarya University, Turkey
Prof. Roger Hendrix
Univ. of Pittsburgh
Prof. Bob Duda
Univ. of Pittsburgh
Robert Cantor
STAR Cryoelectronics
NSF REU, HHMI, CMU SURG funding
Department of Chemistry & MCS
Carnegie Mellon University
NSF DBI 0454980 & NSF CHE-CMI-1611146
Logan Plath, GS
Carnegie Mellon
Stephen Tjader, BS
Carnegie Mellon
Jonathan Feldman, UG
Carnegie Mellon
Fan Wang, MS
Carnegie Mellon
Prof. Emeritus Urs Schlunegger
Univ. of Bern
Prof. Rongchao Jin
Carnegie Mellon
Dr. David Sipe
Carnegie Mellon
Dr. Alex Aksenov
UC Davis
Anna Czapar
Case Western Reserve U
Prof. Nicole Steinmetz
Case Western Reserve U.
VIPERdb2
USDA

## Slide 33
In memory of
Urs P. Schlunegger
September 27, 1934 - September 28, 2015