## Slide 1
EAAC2019 15-21 September 2019 Hotel Hermitage, La Biodola Bay, Isola d'Elba, Italy
Initial Studies on a Compact High-Gradient Ka-Band Accelerating Structure for Medical and Industrial Applications
Luigi Faillace
@INFN-Milan

## Slide 2
Outline
Linacs for industrial and medical applications
Compact High-Gradient Ka-Band linac
RF Design
Beam Dynamics
Conclusions and future plans
EAAC2019 15-21 September 2019 Hotel Hermitage, La Biodola Bay, Isola d'Elba, Italy

## Slide 3
Electron linear accelerators (linacs) are used for numerous applications in many different areas;
Both direct electron beams or X-rays from target (e- conversion);
In research, high accelerating gradients are sought for producing ultra-bright and high-energy (>GeV) electron beams for linear colliders as well as FELs, Compton sources, etc.
Thousands of linacs are sold every year, mainly for industrial and medical applications, these structures mostly operate either in S-Band (~3 GHz) or in X-Band (9.3 GHz)  these structures can be very bulky and large robotic systems are needed for operation, representing a crucial issue especially for medical applications.
We propose here a compact linear accelerating structure in Ka-Band (~36 GHz) able to achieve ultra-high accelerating gradients (up to 150 MV/m) and therefore it allows to obtain high beam energies in ultra-compact foot prints.
The proposed Ka-Band accelerator can be either operated, depending on the specific application, either in Standing-wave (SW) or Traveling-wave (TW) mode, is a table-top and inexpensive solution as accelerator for medical and industrial applications. Due to the small geometric dimensions, the maximum average electron current in Ka-Band is lower than lower frequencies and it is therefore suitable for low average power applications;
The Ka-Band accelerator is also designed as a higher harmonics linac for the linearization of the electron beam for the CompactLight project (see Bruno Spataro’s Talk on Thursday)
Compact High-Gradient Ka-Band linac
EAAC2019 15-21 September 2019 Hotel Hermitage, La Biodola Bay, Isola d'Elba, Italy

## Slide 4
 | Application | Beam energy | Average Current
Research | High-Energy Physics | >1GeV | 10’s µA - 10’s mA
Medical Area | Diagnostics |  | > 1 µA
 | Radiotherapy | 4-25 MeV | 10 nA – 1 µA (X-rays of 200-500 cGy/min@1m, flux of 107-8 photons/s/mm2)
 | Radioisotopes Production | Up to 70-90 MeV | Up to 10 mA
 | FLAH therapy | Up to 7 MeV | >10’s µA
 | Blood Irradiation | 1-2 MeV | >10’s µA
Industry | Material Processing | < 10 MeV | 10-150 mA
 | Sterilization | < 10 MeV | Up to 10 mA
 | Non-Destructive Testing (NDT) |  | > 1 mA
Security | Cargo Scanning | < 10 MeV | 0.1 mA – 1 mA
Environmental Area | Water, flue gas treatment, etc. | 0.7-5 MeV | 5-10 mA
Applications
EAAC2019 15-21 September 2019 Hotel Hermitage, La Biodola Bay, Isola d'Elba, Italy

## Slide 5
Modulator
circulator
Directional Coupler
Basic linac accelerator system layout
e-gun
load
RF power source
Vacuum system
Gun driver
Cooling system
linac
RF
windows
target
EAAC2019 15-21 September 2019 Hotel Hermitage, La Biodola Bay, Isola d'Elba, Italy
X-rays
e-

## Slide 6
The choice depends on:
  Requirements of the application
  Availability of the commercial microwave power source
  Knowledge and technology
L-Band
  High beam power ~1MW (e.g. food irradiation)
S-Band is the most common:
  Power source commonly available
  Medium size/weight
  Technology is well established
X-Band
  More compact and lighter  mini-system and portable systems
  More stable and easier to manage (e.g. Radiotherapy)
  Mobile and transportable systems (e.g. Cargo Inspection)
Ka-Band
  Ultra compact  suitable for industrial and medical applications?
Different frequencies for different applications
EAAC2019 15-21 September 2019 Hotel Hermitage, La Biodola Bay, Isola d'Elba, Italy

## Slide 7
2D field plots of the main accelerating cell, electrically coupled, operating in the TM010-like mode.
E-field
H-field
2.9 mm
Main RF Parameters | Value
Frequency | 35.982 GHz
Accelerating Gradient | Up to 150 MV/m 125 MV/m (safe value)
Eff. Shunt Impedance | 158 M/m
Quality Factor Q0 | 4110
Cell length Lc | 2.9 mm
Main RF parameters for Ka-Band accelerating cell
Max Esurface = 250 MV/m
Max Hsurface = 0.33 MA/m
RF Pulsed Heating
ΔT = 16.5 C
(below 50 C safety threshold)
TRF = 50 ns, flat top
Modified Poynting Vector
Sc ~ 8 MW/mm2
From Bruno Spataro’s Talk
 Accelerating gradients are usually lower in industrial linacs due to the particular high shunt impedance cavity geometries.
EAAC2019 15-21 September 2019 Hotel Hermitage, La Biodola Bay, Isola d'Elba, Italy

## Slide 8
Beam break-up (BBU), mainly due to the excitation and amplification of the first dipole deflecting mode (TM110-like), is a major cause of beam instability that limits the maximum current that can be supported inside an accelerating structure
Maximum beam peak current strongly depends on the geometric dimensions (transverse size and length L), that scale linearly with the operating frequency f.
Max current is proportional to (/L)3 and (/L)2 for TW and SW linacs, respectively  measured beam currents in industrial linacs  2 A (TW) and 0.5 A (SW) in S-band (3 GHz).
Here we consider SW operation:
For the Ka-band structure in SW mode, we derive a peak beam current threshold that in the worst-case scenario (trapped deflecting mode) is equal to ~70 mA.
The average output current depends on many factors, e.g. input RF power specifications and the maximum average RF power that can be dissipated by the cooling system of the structure.
Estimation of maximum available beam current in Ka-Band
SW linac
[T. Wangler, “RF Linear Accelerators”]
EAAC2019 15-21 September 2019 Hotel Hermitage, La Biodola Bay, Isola d'Elba, Italy

## Slide 9
RF power supply at ~36 GHz was demonstrated at Yale University more than a decade ago, the Magnicon.
Research activity on the re-design of the old Magnicon in order to improve the output power and the repetition rate at least up to 100 Hz  Klystron version with 3rd harmonic output!
Operation in Ka-band requires relatively low RF power (~5 MW) for industrial and medical applications.
Research activity on the electron gun:
  High energy physics: 500 kV for a beam current of about 200 A and beam power up to 100 MW.
  Industrial/Medical field: 15-30 kV voltage (or higher to increase the beam capture inside the Ka-Band linac).
Collaboration: Pending Memorandum Of Understanding between Yale and INFN-LNF.
Ka-Band RF Power Source
Hirshfield et al., Proceedings of 2005 Particle Accelerator Conference, Knoxville, Tennessee
EAAC2019 15-21 September 2019 Hotel Hermitage, La Biodola Bay, Isola d'Elba, Italy

## Slide 10
The availability of input RF peak power (5MW) with repetition rates between 10 and 100 HZ allows to obtain an average beam current of 1-10 A, which makes the accelerating structure suitable for medium-energy, low current/power applications like medical diagnosis and therapy of tumors as well as intra-operative radiotherapy.
Pulse lengths of the order of 1s, as needed in medical and industrial applications, are compatible with an accelerating gradient of 80 MV/m, since the estimation of the breakdown rate (BDR), that is a major cause of failure of an accelerating structure, is below the safety threshold.
The breakdown limit to the max accelerating gradient is not an issue in this case, since we envision operation below 10 MeV for industrial and medical applications.
Ka-Band RF Power Source for low-energy linacs
Main RF Parameters | Value
Frequency | 35.982 GHz
Accelerating Gradient | ~ 80 MV/m
Linac Length | ~ 10 cm
Peak RF power | ~ 5 MW
Pulse Length | Up to 1.5 s
Repetition Rate | Up to 100 Hz
Duty Cycle | 1.5*10-4
EAAC2019 15-21 September 2019 Hotel Hermitage, La Biodola Bay, Isola d'Elba, Italy

## Slide 11
Electron gun and Linac Design
Lanthanum-based Thermionic cathode 100 A/cm2  1A/mm2
DC diode gun, 15 kV and 250 mA beam current
200 um
1mm
Various types of linac configurations are currently used:
Standing Wave
Pi-mode operation (On-axis electrically coupled linac)
Pi/2-mode (Off-axis/On-axis resonant coupling biperiodic linac)
Traveling Wave
2pi/3 mode (On-axis electrically coupled linac)
5pi/6 mode (Off-axis magnetically coupled linac)
The choice depends on the application and required beam quality (current, spot size, emittance, energy spread…).
Main RF Parameters | Value
Frequency | 35.982 GHz
Accelerating Gradient | ~ 80 MV/m
Linac Length | ~ 10 cm
Peak RF power | ~ 5 MW
Pulse Length | Up to 1.5 s
Repetition Rate | Up to 100 Hz
Duty Cycle | 1.5*10-4
70 mA, 7.5 MeV
Beam Loading Curve
EAAC2019 15-21 September 2019 Hotel Hermitage, La Biodola Bay, Isola d'Elba, Italy
beam current [A]
Energy [eV]

## Slide 12
Initial Beam Dynamics, Input Parameters
PARMELA simulations
15kV DC e-gun (DC electron beam), 250 mA e-beam
DC Input electron beam with gaussian transverse profile (spot size = 0.2 mm diameter, hard edge)
Side-coupled biperiodic accelerating structure (24 cells, L~10cm)
Cell iris diameter = 2mm
No focusing solenoid
Low-beta bunching section
Beam capture ~ 30 % (typically)
Side-coupled biperiodic linac
Beam
path
EM energy
cathode
On-axis Electric Field (MV/m)
EAAC2019 15-21 September 2019 Hotel Hermitage, La Biodola Bay, Isola d'Elba, Italy
(cm)

## Slide 13
Initial Beam Dynamics Output Results
Output current 71 mA
~1 ps
Energy spread <10%
Linac axis
Energy Gain
Beam Transverse Profile
e-beam
Linac exit
EAAC2019 15-21 September 2019 Hotel Hermitage, La Biodola Bay, Isola d'Elba, Italy

## Slide 14
We have presented the preliminary design of a compact linear accelerator in Ka-Band (~36 GHz) which is well suited for industrial and medical applications where a low-medium energy and low average current electron beam is required (from a few MeV’s up to a few tens of MeV’s with average beam currents up to few tens of A’s).
These applications cover the diagnosis and therapy of tumors (both direct electrons or X-rays from e-beam conversion through a target), as well as any other type of diagnostics.
We plan to machine the Ka-Band linac for high gradient applications in two halves with TIG welding of the outer surfaces. As for medium-low energy industrial/medical applications, we plan to apply the same approach but we are also considering an alternative approach like novel clamping techniques (D.A. Dolgashev, L. Faillace, B. Spataro and R. Bonifazi, “Innovative compact braze-free accelerating cavity”, Journal of Instrumentation JINST 13, no. 09 (2018): P09017.)
RF and Beam dynamics Optimization (cavity geometry, focusing solenoid,…)
Conclusions and Future Work
Thanks for your attention!!!
Braze-free cavity
EAAC2019 15-21 September 2019 Hotel Hermitage, La Biodola Bay, Isola d'Elba, Italy