## Slide 1
Un amplificatore SQUID per il segnale del rivelatore KLASH
Paolo Falferi FBK/TIFPA

## Slide 2
The Superconducting Quantum Interference Device combines two phenomena
Josephson Tunneling
F0=h/2e ≈ 210-15 Wb
is the Flux Quantum
F=nF0 (n=0, ±1, ±2...)
I
Flux Quantization
Th: F. London (1950)
Exp: Deaver and Fairbank (1961)
Th: Josephson (1962)
Exp: Anderson and Rowell (1963)
Superconductor
I0 is the Critical Current (maximum supercurrent)
f is the phase of the macroscopic wave-function that describes the Cooper pair condensate
I=I0sind d=f1-f2
dd/dt=2eV/h=2pV/F0
Josephson Equations
The dc SQUID
h ≈ 6.6x10-34 Js is the Planck constant
e ≈ 1.6x10-19 C is the electron charge
30/11/18
Klash kick-off meeting LNF
2

## Slide 3
V
I
R
C
 = (n+1/2) 0
 = n 0
V
I
1
2
3
/0
V
I0 ~ 10-20 µA
Two Josephson junctions connected in parallel on a superconducting loop
Each junction has a self-capacitance C and an added resistive shunt R
Provided that (I condition) the I-V characteristic is nonhysteretic
II condition:
bC= 2pI0 R2C/F0 ≤ 1 .
For bc<<1 V=R(I2-I02)1/2 ≈ RI for I>>I0
The dc SQUID
2I0
I0
I-V Characteristic
V-F Characteristic
2pkBT/I0F0<<1
Ib
Max response to a dFa<<F0
at Fa= (2n+1)F0/4
DV ~ 20-100 µV
30/11/18
Klash kick-off meeting LNF
3

## Slide 4
The white noise of low-Tc SQUIDs is in general in good agreement with these predictions.
From numerical simulation the SQUID response is optimized if
1) bC= 2pI0 R2C/F0 <≈ 1 .
2) bL= 2LI0/F0 = 1 where L=SQUID loop inductance
For these conditions
Transfer Function = VF ≈ R/L ≈ 1/√(pLC)
Flux Noise SF(n) = SV(n)/V2F ≈ 16kBTL2/R
Energy Resolution = e = SF(n)/2L ≈ 9kBTL/R ≈ 16kBT√LC

V
Slope VF ≈ R/L
e ≈ 10-33 J/Hz ≈ 10h for T=0.3K L=50pH R=2W
The dc SQUID
the Noise scales with the Temperature
30/11/18
Klash kick-off meeting LNF
4

## Slide 5
Mi
G
Li
LSQ
Flux-Locked Loop SQUID
For small signals (<0.1F0) the SQUID can be operated in a linear range around the optimum working point W
The dc SQUID
Electronics - Flux-Locked Loop
Flux of the Earth's magnetic field
(50 mT) through 10mm10mm is 2.5F0
For larger signals a linearization of the transfer function is realized with a flux-locked loop (FLL)
The deviation of the SQUID voltage V from that at the optimum working point is amplified, integrated, and fed back via a feedback resistor Rf and a feedback mutual inductance Mf
FLL SQUID transfer function = Rf/Mf
Mf
Rf
Vout
Preamp +
Integrator
30/11/18
Klash kick-off meeting LNF
5

## Slide 6
Feedback Line
Sensor SQUID
Vout
Amplifier
SQUID
Matching
Network
Iin
Two-stage SQUID amplifier
The dc SQUID
Electronics - Methods to Suppress Preamplifier Noise
In principle the noise is thermal it scales with temperature
30/11/18
Klash kick-off meeting LNF
6

## Slide 7
Minimum Energy Resolution e = 35 h
Saturation due to the Hot Electron Effect
Two-stage SQUID amplifier
The dc SQUID
Electronics - Methods to Suppress Preamplifier Noise
30/11/18
Klash kick-off meeting LNF
7

## Slide 8
Fabrication Technology
V
I
Mi
Li
V
I
To improve the magnetic field resolution:
larger pickup loop
+
multi-turn spiral input coil
Lp
Tight coupling (Mi2LiLs) is obtained with multi-turn input coil integrated on top of the SQUID washer
Ls
Flux Transformer
Wire-wound or
Thin-film
Pick-up
The dc SQUID
50-150 mm
Practical realization
Typical dc SQUID layout
Square Washer - Ls=80-300 pH
Uncoupled SQUID
Coupled SQUID
Input Coil
R. Cantor and D. Koelle
30/11/18
Klash kick-off meeting LNF
8

## Slide 9
The conventional square-washer SQUID can be operated as a low-noise amplifier up to about 100 MHz. At higher frequencies the parasitic capacitance between the input coil and the square washer lowers the gain to useless levels
Lo SQUID per KLASH @ 70-250 MHz
Il Problema
dell’alta frequenza
30/11/18
Klash kick-off meeting LNF
9

## Slide 10
M. A. Tarasov et al., IEEE Trans. Appl. Supercond. 2, 79 (1992)
L. Spietz et al., Appl. Phys. Lett. 95, 092505 (2009)
Sono stati sviluppati diversi metodi per estendere la freq di operazione.
D. Drung et al., IEEE Trans. Appl. Supercond. 15, 777 (2005)
30/11/18
Klash kick-off meeting LNF
10
Lo SQUID per KLASH @ 70-250 MHz
Il Problema
dell’alta frequenza

## Slide 11
In contrast to the conventional input scheme the signal is applied between one end of the coil and the washer (the other end of the coil is left open). The washer provides the groundplane for the resulting microstrip
Compared to cold semiconductor amplifiers, MSA offer lower noise (a factor 50 for n < 1GHz) and much lower power dissipation.
Near-quantum-limited performance in the gigahertz range
The resonant frequency can be changed
by connecting a varactor diode
Microstrip SQUID Amplifier (MSA)
Quantum Noise Temperature
at 700 MHz: 33mK
Only one manufacturer of MSA: ez SQUID
A MSA has recently replaced the HEMT amplifier of the Axion Dark Matter eXperiment that uses a microwave cavity at 800 MHz operating at ultracryogenic temperatures
M. Muck and R. McDermott
30/11/18
Klash kick-off meeting LNF
11
Lo SQUID per KLASH @ 70-250 MHz
Il Problema
dell’alta frequenza

## Slide 12
Accoppiamento con cavità
Accoppiamento tipo ADMX cioè
“critically coupled antenna”
“MSA load-matched to 50 Ohm via a terminated Circulator”
Non sente i disturbi a basse frequenze
Su ADMX funziona
Funziona giù fino a 70 MHz?
MSA
SQUID
Accoppiamento magnetico “tradizionale”
Per non sentire rumore vibrazionale  ci deve essere un accoppiamento passa alto o banda ad alto Q
Va simulato e ottimizzato l’intero sistema cavità-pickup-squid
 ΔS = 1 μm2 in 0.6 T => 300 F0
cavità
accoppiamento
risonante
SQUID
30/11/18
Klash kick-off meeting LNF
12
Lo SQUID per KLASH @ 70-250 MHz
Il Problema
dell’alta frequenza

## Slide 13
Li  H
Z
Zin= iLi + Z(, Ibias, , ….)
The SQUID amplifier
Mi
Li
Mf
G
Mi-f
!
Usually negligible.
Important with high Q input load and at high frequency signals (~MHz)
its input impedance
Input Impedance of the dc SQUID Amplifier
Vout
Iin
Instabilità
Accoppiamento magnetico con cavità
Lo SQUID per KLASH @ 70-250 MHz
30/11/18
Klash kick-off meeting LNF
13

## Slide 14
D(f) =D0/(1+if/fD)
Instabilità
Accoppiamento magnetico con cavità
Lo SQUID per KLASH @ 70-250 MHz
Feedback
cooling
30/11/18
Klash kick-off meeting LNF
14

## Slide 15
Power spectral density of the SQUID input current for four different values of the feedback amplifier gain D0
D0=2
D0=10
D0=50
D0=100
Resonance Frequency = 11.5 kHz
Operating Temperature = 135 mK
Minimum Effective Temperature = 14 mK
Instabilità
Accoppiamento magnetico con cavità
Lo SQUID per KLASH @ 70-250 MHz
Feedback
cooling
30/11/18
Klash kick-off meeting LNF
15

## Slide 16
Mi
Li
G
LSQ
FLL SQUID
Back-action Noise
Usually negligible.
Important with high Q input load
and at high frequency (~MHz)
Additive
Noise
Vn
In
Li
Noise Model
!
Back-action noise
Accoppiamento magnetico con cavità
Lo SQUID per KLASH @ 70-250 MHz
30/11/18
Klash kick-off meeting LNF
16

## Slide 17
Things to do list
Verificare la possibilità di utilizzo di un Microstrip SQUID Amplifier fino a 70 MHz e nel caso acquistarne uno (ezSQUID)
Costruzione di una test facility operante in LHe
Acquisire e testare sensori SQUID custom realizzati per alta frequenza (bassa capacità parassita fra input coil e SQUID loop) da Supracon? PTB/Magnicon? IBS/Corea?
Elettronica dedicata, veloce e a basso rumore, su modello di quella del PTB
Cavo criogenico speciale più corto possibile tra elettronica e sensore squid
Progettazione FEM del sistema cavità – pick-up – linea risonante – SQUID con ottimizzazione del rumore
Costruzione di una linea risonante a basse perdite tra pick-up e induttanza di ingresso dello SQUID
30/11/18
Klash kick-off meeting LNF
17

## Slide 18
FINE

## Slide 19
High spatial resolution
small pickup loop (F ≈ 10 mm)
SQUID loop as pickup
(uncoupled SQUID)
High field resolution
large pickup loop
thin-film or wire-wound coil connected to a coupled SQUID
Best performance
SB 1/2 = 0.9 fT/Hz1/2
Magnetometers for High Spatial Resolution
(for scanning magnetic microscopes)
Magnetometers for High Field Resolution
Mi
Li
Lp
Ls
Flux
Transformer
Wire-wound or
Thin-film
Pickup
Magnetometers and Gradiometers

## Slide 20
Magnetometers and Gradiometers
(a) Magnetometer
(b) First-derivative axial gradiometer
(c) Second-derivative axial gradiometer
(d) Thin-film planar first-order gradiometer (parallel and series pick-up loops)
When
the magnetic field source is "distant" or
the environment is magnetically noisy
Magnetometer GRADIOMETER
pick-up balance up to 0.001% (with high-precision fabrication and post-assembly balancing)
Thin-Film Planar Gradiometers
Two-dimensional film
Planar Gradiometers
(activity in brain or heart)
and
Miniature Susceptometers
(magnetic properties of particles down to 10 mm diam and magnetic calorimeter x-ray detectors )
Wire-Wound Axial Gradiometers
deep-lying sources
Long baseline axial gradiometers
wire-wound pickup coils transformer-coupled to the SQUID
A. Braginski and J. Clarke

## Slide 21
Noise Temperature
“True” Energy Resolution
Energy Resolution from additive noise is incomplete
Noise Impedance
Quantum Magnetics two-stage SQUID
8.9 kHz
Noise Sources of the dc SQUID Amplifier

## Slide 22
Vout
A
Lp
Li
L
R
C
Ci
Ci