## Slide 1: A Compact X-ray Source for Advanced Radiological Imaging
THE MESSAGE
BRING IMPROVED IMAGING QUALITY
for BIOMEDICAL APPLICATIONS
to CLINICAL and UNIVERSITY
LABORATORIES

## Slide 2
DRAWBACKS in CONVENTIONAL (X-ray Tubes) ABSORPTION DIAGNOSTICS
on Soft Tissues: Cartilages, Angiography, Breast, Brain, Lung
THE CHALLENGE
ENHANCED IMAGING
BETTER RESOLUTION
SHORTER EXPOSURE / DOSE
LESS CONTRAST AGENTS
ADVANCED IMAGING TECHNIQUES
X-ray SOURCES
Energy-tunable
High flux
Monochromatic
EXIST
NOT (YET) AVAILABLE
IN CLINICS
Quality usually enhanced using contrast agents
not always patient-friendly
Polychromatic, Not Tunable, Low contrast

## Slide 3
Courtesy: G. Tromba et al. / Elettra TS
MEDICAL IMAGING TECHNIQUES
Differential Phase Contrast Imaging (DPCI)
Courtesy: F. Pfeiffer et al. / PSI
ABI - Courtesy: M. Kitchen et al. / Monash University
X-ray Refraction
  Analyzer-Based Imaging (ABI)
Grating Interferometry (GI)
Courtesy: G. Tromba et al. / Elettra TS
Diffraction-Enhanced Computerized Tomography (DEI-CT)
Nuclear Magnetic Resonance (NMR)
CONVENTIONAL MAMMOGRAPHY: LIMITED RESOLUTION FOR DENSE BREASTS
Courtesy: G. Tromba et al. / Elettra TS
X-ray Absorption
Courtesy: G. Tromba et al. / Elettra TS

## Slide 4
NMR / DPCI DIAGNOSTICS COMPARISON
Courtesy: R. Calandrino et al. / SRH Milano
OPTION |  | NMR | DPC
Spatial resolution [mm] |  | >3 | 0.01-0.03
Data registration time [s] |  | 1200 | 30
Image reconstruction time [s] |  | 120 | <1
Primary diagnostics tool? |  | NO | YES
Vessel system | VISIBILITY w/o contr. agent | NO | YES
Nervous system |  | NO | YES
Lymphatic system |  | NO | YES
Soft tissues |  | NO | YES
Skeleton |  | YES | YES
Claustrophobic patients |  | Difficult | Possible
Patients with metallic implants |  | Impossible | Possible
Patients with non-metallic implants |  | Limited | Possible
Real time examination during surgery |  | Impossible | Possible

## Slide 5
ESRF / Grenoble
6 GeV e- / 844 m Circumference
10-40 keV X-rays
270 m
ELETTRA / Trieste
2.4 GeV e- / 260 m Circumference
5-10 keV X-rays
LAB-size LIGHT SOURCES
83 m

## Slide 6
A COMPACT X-ray SOURCE
MuCLS / Munich Compact Light Source
Courtesy: LTI & F. Pfeiffer et al. / TUM
CLS / Lyncean Technologies Inc. / Palo Alto, CA
40 MeV e- / 5x4 m2 footprint
30-40 keV X-rays
5 m
MuCLS @TUM
TOP VIEW
SIDE VIEW
e-
30-40 keV X-rays
at 65 MHz rep. rate
3x10^11 ph/s
e- Desktop Storage Ring
35-40 MeV, 65 MHz, 4.6 m
e-
Laser Optical Cavity
300kW, 65 MHz, 9.2 m
IR Laser
IP
Linac

## Slide 7
MuCLS Dynamic In Vivo Chest X-ray Dark-Field Imaging in Mice R. Gradl et al: IEEE Transactions on Medical Imaging ( Volume: 38 , Issue: 2 , Feb. 2019 )
X-ray grating Interferometry combines
-- Conventional attenuation
-- Phase Contrast
-- Dark Field
Dark-field signal depends on number and size of alveoli
Useful tool in lung tissue imaging
Blur reduction during a full breath cycle
Short exposures + High flux
40 ms, 90 ms and 140 ms exposure times
30, 15 and 10 images per breath.
Absorption
DPC
Dark Field imaging
Dark Field signal

## Slide 8
ADVANTAGES
270 m
e-
lu
UNDULATOR
ESRF: 6 GeV e- / 844 m Circumference
10-40 keV X-rays
Laser pulses
Mini-undulators
ICS
FROM GOLIATH TO DAVID…
X-ray energy has
similar expression
CLS: 40 MeV e- / 5x4 m2 footprint
30-40 keV X-rays
5 m
IMAGING DISTANCE
COLLIMATION

## Slide 9
A COMPACT X-Ray SOURCE
The availability of a compact source
offering high quality X-ray imaging
at moderate cost and limited space requirements
can have an important impact in
Clinical and University-size Laboratories.
BoCXS: A Multidisciplinary Imaging Facility
in the center of a Scientific Triangle

## Slide 10
Double Output S-Band STAR–like System
Phase I
e- Dumping
Dipoles
~9 m
S-Band RF Structure
3m / 25MeV/m
80 MeV
~4 m
S-Band Ph-Inj
Matching
Triplets
e-
200 - 250 Doglegs
Interaction
Regions
Yb-YAG
Laser
hL1
hL2
X-Ray
BPM
H/V Steering
IP Diag.
e-
X-Ray
IP Diag.
USERS
e-

## Slide 11
PHASE I – e- ENERGY: 80 MeV tunable
X-ray ENERGY : (120 – 240) keV
X-ray FLUX : (2 – 0.5)x1010 ph/s
X-ray DIVERGENCE : +/- 6.4 mrad
X-ray REP. RATE: 100 Hz
INITIAL PERFORMANCE

## Slide 12
OUTLOOK
A COMPACT X-ray SOURCE
INSTALLED IN CLINICAL LABORATORIES OR UNIVERSITIES
WOULD PROVIDE AN INVALUABLE TOOL IN SOFT TISSUE IMAGING
IN THE BOLOGNA METROPOLITAN AREA
IMPROVED SPACIAL RESOLUTION
REDUCED BEAM EXPOSURE
RADIOLOGICAL EXAMINATIONS WITHOUT CONTRAST AGENTS
FRIENDLY TO PATIENTS WITH
METALLIC IMPLANTS AND/OR STIMULATORS
HIGH BRILLIANCE, ENERGY-TUNABLE, QUASI-MONOCHROMATIC
X-RAY BEAMS
PROVIDE ADVANCED BIOMEDICAL IMAGING

## Slide 13
A PROJECT IS A DREAM WITH A DEADLINE
THANK YOU FOR YOUR TIME

## Slide 14
SPARE SLIDES

## Slide 15
FREQUENCY BANDS IN PARTICLE ACCELERATORS
BAND | fRF (GHz) | lRF (cm) |  | COMMON LIFE
S | 2.998 – 2.856 | 10.0 –10.5 |  | Satellite communications
C | 5.996 – 5.712 | 5.0 –5.25 | S-Band 2nd harmonic | 5G technology
X | 11.9942 – 11.424 | 2.5 –2.62 | S-Band 4th harmonic | Radar applications

## Slide 16
Double Output S-Band STAR–like System
Phase II
~12 m
155 MeV
~7 m
S-Band Ph-Inj
Matching
Triplets
200 - 250 Doglegs
Interaction
Regions
Yb-YAG
Laser
hL1
hL2
X-Ray
BPM
H/V Steering
IP Diag.
X-Ray
IP Diag.
USERS
e- Dumping
Dipoles
2 x S-Band RF Structures
6m / 25MeV/m
e-
e-
e-

## Slide 17
PHASE I – e- ENERGY: 80 MeV
X-ray ENERGY : (118 – 235) keV
X-ray FLUX : (20 – 5)x109 ph/s
X-ray DIVERGENCE : +/- 6.4 mrad
MODULAR IMPLEMENTATION PHASES
PHASE II – e- ENERGY: 155 MeV
X-ray ENERGY : (0.44 – 0.88) MeV
X-ray FLUX : (20 – 5)x109 ph/s
X-ray DIVERGENCE : +/- 3.3 mrad
PHASE III – e- ENERGY: 155 MeV
X-ray ENERGY : (0.44 – 0.88) MeV
X-ray FLUX : (6 – 1.5)x1011 ph/s
X-ray DIVERGENCE : +/- 3.3 mrad
Rep. rate = 100 Hz
Laser: Fund. & 2nd harm.
Rep. rate = 3 kHz
Laser: Fund. & 2nd harm.