## Slide 1
DCH FEE STATUS
Some ideas for Level 1 Triggered Data Flow
&
A preliminary study on Cluster Counting Data Processing Algorithm
G. Felici

## Slide 2: Outline
DCH TRIGGERED DATA FLOW
  Trigger & OL Specs (reminds)
  ADB main blocks & data frame (remind)
  The trigger burst case – 2 possible readout implementation
  A pushing-mode front-end readout architecture
Cluster Counting Data Processing algorithm
  The preamplifier
  Noise filtering based on DSP correlation technique
  The correlation algorithm
  Example of the technique applied to acquired waveform
Conclusions

## Slide 3: Trigger & OL specs (remind)
System
Trigger rate (average): 150 kHz
Trigger (fixed) latency : ≈ 4 (6 !?) μs
Data OL BW : 16 OL @ 2 Gbits/sec
ECS OL BW : 16 OL @ 2 Gbits/sec
Trigger OL BW : 64 OL @ 1.2 Gbit/sec
Trigger spacing (min) ≈ 60 ns (!?)
Trigger burst : 3 events (?)
Detector
Number of cells (guess): ≈ 9216
Chamber occupancy : 15% (Inner layers)
Chamber gain : 5104 - 1105
Sense wire parasitic (CD) ≈ 25 pF
Specs

## Slide 4: ADB main blocks & data frame (remind)
DISCR
TDC
LPF
FADC
28 MHz sampling rate
1
0
32 samples (≈ 1.14 μs) allow DC full signal development and includes distribution time jitter
2
1
3 Events RO Buffer
Single Channel ROIB Digitizer (ADB)
Trigger Latency Time + n samples
?
3
L1
n
ADC
#
0
ADC
#
1
ADC
#
2
ADC
#
3
ADC
#
4
ADC
#
5
TDC
#
1
ADC
#
7
ADC
#
8
ADC
#
9
TDC
#
2
ADC
#
11
ADC
#
12
ADC
#
13
ADC
#
14
ADC
#
15
ADC
#
16
ADC
#
17
ADC
#
18
ADC
#
19
ADC
#
20
ADC
#
21
ADC
#
22
ADC
#
23
ADC
#
24
ADC
#
25
ADC
#
26
ADC
#
27
ADC
#
28
ADC
#
29
ADC
#
30
ADC
#
31
SAMPLING
7 MHz
Trigger
Position inside the frame  coarse time measurement (5 bits)
Content  fine time measurement (6 bits)
Digitized Data Frame Example

## Slide 5: The trigger burst case
32 samples
Ch n
Trigger 1
32 samples
32 samples
FEE data transfer optimization : download the 3 events as a single “big” event
  Pro : Front-End data transfer optimization
  Cons :
    events are overlapped in the same data frame (further elaboration required to split single events)
    data frames do not have the same lengths (L1 trigger occurrences)
Ch n+i
Trigger 2
Trigger 3
Ch n+i
Data transfer optimization
32 samples
Ch n
Trigger 1
32 samples
32 samples
Single event data transfer : readout each event separately
  Pro :
    Front-End events have the same size
    FEX can be applied while reading the event
    Readout procedure provides event de-randomization
  Cons :
    Partial (previous) L1 event re-reading
Ch n+i
Trigger 2
Trigger 3
Ch n+i
Event management optimization

## Slide 6: A pushing-mode FE readout architecture
COUNTER (0 – n)
L1 FIFO
Sampled data
DATA RO SM
(FEX)
Pushing mode
ADDR
Ev3
Ev2
Ev1
RO buffer
Concentrator board
Trigger Latency Time + n samples (Dual-port memory)
2
1
0
n
Sampling clock
ADDR
L1 (synchronized by sampling clock)
Read ADDR
FiFO Empty
FiFO Read
Data (counter value @ L1)
(counter value @ L1) - Latency
A simulation of the readout architecture will be carried out in the next weeks
NB : minimum trigger spacing > sampling period (≈ 36 ns)
32 word block data readout

## Slide 7
A preliminary study on Cluster Counting Data Processing Algorithm

## Slide 8: Preamplifier
≈ 250 MHz BW
Reduce 8015 CIN sensitivity
Gain ≈ 5 mV/fC
Noise ≈ 1900 erms @ CIN = 3pF
CINJ = 1.8 pF

## Slide 9: Cluster Counting – noise filtering
noise filtering using an algorithm with low area consumption in a FPGA
We are trying to use a DSP technique based on the correlation concept to determine if a target signal (a signal with a shape close to a well separated cluster) is present inside another signal (the digitized signal). This technique is generally used in radar systems where the received signal consist of a shifted and scaled version of the transmitted pulse and overlapped random noise (interference, thermal noise etc)
Sampled signal

## Slide 10: Cluster Counting – The correlation machine
Received signal
Cross-correlated signal
Target signal
The value of the cross-correlation is maximized when the target signal is aligned with the same features in the received signal
Correlation is the optimal technique for detecting a known waveform in random noise (matched filtering technique)
x[n]*t[-n] = y[n]

## Slide 11: Cluster Counting – An example
The Peak Finding algorithm is under study

## Slide 12: Conclusions
A (very preliminary) front-end data readout architecture capable to manage short dead-time (less than 100 ns) has been presented. We plan to validate the architecture by means of (VHDL) simulations in the next weeks
Some work has been done to verify the possibility of using Cluster Counting for dE/dx measurements. A setup made of a short tube (≈ 40 cm), a fast preamplifier/digitizer has been built.
A very preliminary algorithm based on a matched filtering and peak finding procedure to count single ionization events has been implemented. First results look encouraging.