## Slide 1: Lab activities @LNF Status Report
A. Calcaterra, R. de Sangro, G. Felici,
G. Finocchiaro, P. Patteri, M. Piccolo
INFN LNF
XIV SuperB General Meeting
DCH-I parallel session
LNF, 27 September 2010

## Slide 2: Experimental Setup
Prototype 1
6x4 hexagonal cells à la BABAR
  Guard wires to ensure uniformity of electric field among cells
  Aluminized mylar windows on entrance-exit faces
  small gas volume
    low rate on cosmics
    fast gas mixture switching times
LST Telescope:
Two identical assemblies of 26 tubes each
3 cm diameter, 100 μm wires
40%-60% Ar-iC4H10 mixture

## Slide 3: Tracks reconstructed in the telescope
Al-Mylar
windows
10cm lead
pcut~150MeV/c

## Slide 4: Space-time relations
Measured time vs. extrapolated impact parameter – t vs. R
Impact parameter vs. time – R vs. t
Example:
60%He-40%C2H6 @ HV=2010V

## Slide 5: Space-time relations fits
Fit R vs. t relation using Chebyshev polynomials of order (up to) 5
Time range rescaled to [-1,+1] interval
60%He-40%C2H6
HV=2010V

## Slide 6: Drift velocity
60%He-40%C2H6
HV=2010V

## Slide 7: Space-time relations fits (cont.)
LEFT
RIGTH
L+R
Internal cells tend to be left-right symmetric
Column 0 1 2 3
Row
0
1
2
3
4
5
60%He-40%C2H6
HV=2010V

## Slide 8: Space-time relations fits (cont.)
LEFT
RIGTH
L+R
Border cells show some asymmetry
Column 0 1 2 3
Row
0
1
2
3
4
5
60%He-40%C2H6
HV=2010V

## Slide 9: Track fit residuals
Fit X slices with Gaussian + pol0

## Slide 10: Spatial resolution – 60%He-40%C2H6
60%He-40%C2H6
HV=2010V

## Slide 11: Tracks reconstructed in Proto1
Al-Mylar
windows
10cm lead
pcut~150MeV/c
BLUE line is MT track
RED line is PROTO1 track

## Slide 12: Spatial Resolution form Proto1 tracks
60%He-40%C2H6
HV=2010V
Studied layer excluded from track fit in prototype
Preliminary: still need to work out fine details (tails, de- convolution of track resolution… )

## Slide 13: STR & Spatial resolution – 73%He-27%C2H6
73%He-27%C2H6
HV=1910V

## Slide 14: STR & Spatial resolution – 80%He-20%C2H6
80%He-20%C2H6
HV=1860V

## Slide 15: Gas Mixture Comparison
 | r(g/cm3) x1 04 | X0(m) | NP(cm-1) | NT(cm-1) | tmax(ns) | s(mm)
80%He-20%iC4H10 | 6.3 | 810 | 21 | 47 | 420 | 130
90%He-10%iC4H10 | 4.0 | 1410 | 13 | 26 | 550 | 170
95%He- 5%iC4H10 | 4.0 | 1410 | 9 | 15 | 700 | 200
35%He-65%CH4 | 4.9 | 1010 | 26 | 37 | 210 | 200
51%He-49%CH4 | 4.1 | 1260 | 20 | 29 | 300 | 220
67%He-33%CH4 | 3.3 | 1690 | 15 | 21 | 350 | 200
79%He-21%CH4 | 2.7 | 2270 | 11 | 15 | 470 | 200
60%He-40%C2H6 | 6.4 | 780 | 19 | 49 | 550 | 130
73%He-27%C2H6 | 4.9 | 1080 | 15 | 36 | 600 | 130
79%He-21%C2H6 | 4.2 | 1310 | 12 | 30 | 650 | 130
tmax and space resolution depend on other parameters. Figures shown are preliminary and indicative
Results for iC4H10 and CH4 –based gas mixtures to be updated with analysis software used for C2H6 gases

## Slide 16: Waveform Digitization & Cluster Counting
Read 3 channels of proto1 through DRS4(*) “evaluation board”
Channel arrangement chosen to see cosmic ray tracks
Same channels (after active splitters) as in the standard DAQ chain
(*) A switched capacitor array (SCA) with 1024 cells, capable of digitizing eight analog signals with high speed (6 GSPS) and high accuracy (11.5 bit SNR) on a single chip (http://drs.web.psi.ch/)
[only 4 channels read-out with current version of firmware]
Same acquisition trigger (scintillator counters coincidence) as telescope+proto1 setup shown so far
  Proto1 preamp bandwidth <100 MHz
  Channel #1 of DRS4 board connected to drift tube with faster preamplifier (see next slides)

## Slide 17: Drift Tubes for Cluster Counting
Two square-section tubes (24mm side) built to study cluster counting issues in simplified environment. The anode - as in Proto1 - is a 20mm Tungsten wire.
One tube (400mm long) already equipped with 300MHz bandwidth preamplifier
  Gain ~5mV/fC; noise ~1900 e- RMS

## Slide 18: Drift Tubes for Cluster Counting
400mm long tube currently on top of our setup
Dump the four waveforms when at least one of them exceeds 25mV
In spite of x3.5 higher preamp bandwidth, a noise level similar to proto1 (~1.4 mV RMS) is obtained with a careful shielding (…)

## Slide 19: Drift Tubes for Cluster Counting (cont.)
Adequate preamp bandwidth is (clearly) crucial to count clusters
85%He-15%C2H6 nP=10.2/cm

## Slide 20: Counting Clusters
INPUT SIGNAL
1024 channels,
2Gsample/sec
TARGET SIGNAL
(cluster)
Position of peaks
from finding algorithm

## Slide 21: More Detectors for Counting Clusters
1700mm long tube to study signal propagation along the wire
  built, strung and tested for gas tightness
  in the gas flow circuit and with HV on for some days now
  to be equipped with fast preamp soon
2000 mm long square cell prototype (~20 cells) for realistic study (X-talk, undetected clusters, signal propagation…) being designed
  hope to have it in hand early 2011

## Slide 22: Summary
More He-based mixtures, using ethane as quencher, have been analyzed.
  preliminary vdrift and spatial resolution results show good performance
  Methane still seems best candidate for light, fast mixtures (short readout time window to reduce background)
Very first attempts at identifying clusters in (almost) realistic conditions
  need fast preamps, low noise, low nP mixtures, good control of track length