## Slide 1: Photodiodes
Large, small and new PD

## Slide 2: Old large and small PD with LYSO
Standard assembling configuration
Measurements with Amptek
Two scintillators in coincidence
MIP is measurable with small PD

## Slide 3: Old large and small PD with LYSO Small PD

## Slide 4: Old large and small PD with LYSO Large PD

## Slide 5: Old large and small PD with LYSO Offset PDS Gain 92,55

## Slide 6: Old large and small PD with LYSO Offset PDS Gain 92,55

## Slide 7: Old large and small PD with LYSO Offset PDSL Gain 37,686

## Slide 8: Old large and small PD with LYSO MIP ratio
 | PD large | PD small
MIP | 572,363 | 24,757
Gain | 37,686 | 92,55
Offset | 6,189 | 2,964
MIP norm | 15,023 | 0,2355
MIP norm = (MIP-offset)/Gain
MIP ratio = 15,023/0,2355=63,93
Area ratio = 84,64/1,6=52,9

## Slide 9: Old large and small PD with LYSO LED

## Slide 10: Old large and small PD with LYSO LED

## Slide 11: NEW PD 2 pieces No protection!

## Slide 12: NEW PD Electrical characterization Vdep1=103,8 Vdep2=97,9 Ileak < 1nA

## Slide 13: NEW PD Spectral measurements with AMPTEK First attempt no visible MIP signal with LYSO
Very noisy
High amplitude LED signal is only measurable
Strange dependence on bias voltage
Probably bad PD contact with preamp

## Slide 14: NEW PD Spectral measurements with AMPTEK First attempt no visible MIP signal with LYSO New PD was mounted on CsI crystal

## Slide 15: NEW PD Spectral measurements with AMPTEK First attempt no visible MIP signal with LYSO New PD was mounted on CsI crystal MIP measurement
Normalized signal ratio 9,07

## Slide 16: NEW PD Spectral measurements with AMPTEK First attempt no visible MIP signal with LYSO New PD was mounted on CsI crystal LED measurement @ 415nm
Normalized signal ratio 6,7

## Slide 17: NEW PD Spectral measurements with AMPTEK New PD QE is unknown LED with collimator D =1,2 mm
Signal ratio 0,9

## Slide 18: NEW PD Spectral measurements with AMPTEK First attempt: no visible MIP signal with LYSO Second measurement
Normalized signal ratio 5.78

## Slide 19: NEW PD Spectral measurements with AMPTEK First attempt: no visible MIP signal with LYSO Second measurement
~ 20% difference between two LPD

## Slide 20: New PDL
  Coupled to LYSO with Scionix EJ-550 Optical Grade Silicone Grease

## Slide 21: New PDL
  New PDL were coupled to LYSO with Scionix EJ-550 Optical Grade Silicone Grease
  PDS was coupled with silicon.
  Previous results:
  PDL/PDS (MIP)= 63
  PDL/PDN = 3,2
  Expected PDN/PDS = 63/3,2=20
  Measured PDN/PDS = 14,6

## Slide 22: New PDL/PDS
  Signal ratio for new PDL/PDS was studied with MIP signal using two Amptek A250 CSA and LeCroy HDO6054 oscilloscope for simultaneous acquisition.
  New PDL has a maximal signal in order of 510 mV. PDS has a signal ~ 20 times lover that is too small for precise measurements.
  Few hundred waveforms were acquired and 5 with maximal signal were analyzed.
  PDL/PDS signal ratio varies from 19 to 35.