## Slide 1: Methodologies and diagnostics of advanced time-of flight ion spectrometry for high-repetition-rate high-intensity laser-matter experiments
M. Salvadori1,2,3, F. Consoli1, C. Verona4, M. Cipriani1, M. P. Anania5, P. L. Andreoli1, P. Antici2, F. Bisesto5, G. Costa5,
G. Cristofari1, R. De Angelis6, G. Di Giorgio1, M. Ferrario5, M. Galletti5, D. Giulietti6,7, M. Migliorati3,8, R. Pompili5,
O. Rosmej9,10 M. Scisciò1, S. Zähter9,10 , A. Zigler11
1 ENEA Fusion and technologies for nuclear safety department, C. R. Frascati, Rome, Italy 2INRS-EMT, Varennes, Quebec, Canada 3 Università di Roma La Sapienza, Italy 4Industrial Engineering department, University of Rome Tor Vergata, Italy 5INFN-LNF Frascati, Rome Italy 6Department of Physics, University of Pisa, Italy 7INFN of Pisa, Italy 8 INFN Sezione di Roma, Italy, . 9GSI Helmotzzentrum für Schwerionenforschung GmbH, Darmstadt, Germany 10 Goethe University Frankfurt, Frankfurt am Main, Germany 11Racah Institute of Physics, Hebrew University, Jerusalem, Israel
martina.salvadori@uniroma1.it
EAAC 2021, 20th-24h September, Frascati
This work has been carried out within the framework of the EUROfusion Consortium and has received funding from the Euratom research and training program 2014-2018 and 2019-2020 under agreement No 633053. The views and opinions expressed herein do not necessarily reflect those of the European Commission

## Slide 2: OVERVIEW
01
LASER PLASMA ACCELERATION
a simple introduction
02
TIME OF FLIGHT (TOF)
the basics of the technique
03
ADVANCED TOF
EMP robustness and enhanced dynamic range
04
THE SPECTRUM RECONSTRUCTION
from the raw signal to a calibrated spectrum
05
THE TECHNIQUE APPLIED
some results from the PHELIX experimental campaign

## Slide 3: Laser-plasma acceleration: a simple introduction
Ions+
1. Plasma generation
2. Laser energy absorption by plasma electrons
  Collisional processes (I < 1015 W cm-2)
  Collisionless processes (resonance absorption, vacuum heating, JB heating)
, ps
Ions+
Suprathermal electrons
2.1. electrons transport through the dense target
A. Macchi et al., Rev. Mod. Phys. 85 (2013) 751, H. Daido et al., Reports Prog. Phys. 75 (2012) 056401 and references there in

## Slide 4: Laser-plasma acceleration: a simple introduction
4. The oncoming electrons are thus going to be confined under Coulombic forces to a Sheath field, directed as the Target Normal, that extends for a distance and reaches an intensity of the order of TV/m.
3. The escaping electrons will give rise to an electrostatic potential derived by the charge-separation between themselves and the positively charged target.
5. The intensity of this electric field is high enough to ionize the atoms present at the rear surface of the target.
+
6. These ions, which are subjected to an intense electric field, will be then accelerated to the order of MeV energies. Protons, having the highest charge to mass ratio, are in the most favorable condition for Acceleration
Proton energy up to a few tens of MeV
Short proton bunch duration at the source (fs-ps)
Charge of a few nC
Conical beam of tens of degrees ( 20°)
The spectrum can be approximated with a Maxwellian
Highest energy → lower flux
We need a suitable
diagnostic system!
A. Macchi et al., Rev. Mod. Phys. 85 (2013) 751, H. Daido et al., Reports Prog. Phys. 75 (2012) 056401 and references there in
P. Puyuelo-Valdes et al.
Sci. Rep. 11:9998 (2021)

## Slide 5: Diagnostic system requirements
The ideal diagnostic system should have:
High sensitivity
High energy resolution
and should allow to retrieve:
Spectrum of accelerated ions
Angular distribution of accelerated ions
but it also has to provide:
Electro Magnetic Pulses (EMPs) robustness
Real-time detection and characterization of ions
Layered Imaging systems (RCF, Imaging plates…)
Solid state nuclear track detectors
Thomson spectrometers
Electrostatic-magnetostatic spectrometers
Time Of Flight technique

## Slide 6: Diagnostic system requirements
The ideal diagnostic system should have:
High sensitivity
High energy resolution
and should allow to retrieve:
Spectrum of accelerated ions
Angular distribution of accelerated ions
but it also has to provide:
Electro Magnetic Pulses (EMPs) robustness
Real-time detection and characterization of ions
Layered Imaging systems (RCF, Imaging plates…)
Solid state nuclear track detectors
Thomson spectrometers
Electrostatic-magnetostatic spectrometers
Time Of Flight technique

## Slide 7: The Basics of Time-Of-Flight (TOF)
time
Ion, v1
Ion, v2 < v1
Photons, c
TOF technique relies on the measurement of the time needed by a particle to travel for a known distance
The temporal position T0 of the photopeak provides a reliable signature of the laser-matter interaction instant
Particle energies are computed evaluating the delay between the signal detection time and
T0
Photopeak
Ions
Accelerating potential: temporal window with pure proton contribution
Or in terms of proton energy:

## Slide 8: Challenges in TOF: EMP and Dynamic Range
TOF signals have an intrinsic high dynamic range.
It is necessary to develop a technique able to appreciate the full dynamic range
EMP pollution hinders the employment of TOF in high energies facilities
It is necessary to optimize the system to work in these environments
Typical signal from Time Of Flight in Laser-Plasma environment
Low precision in defining T0
  Poor energy resolution
X-rays peak
Ion signal
Poor Signal to Noise Ratio

## Slide 9: Challenges in TOF: EMP and Dynamic Range
τ = 25 fs
Target : 10 µm Al
FLAME
E = 2.5 J
W
τ = 750 fs
Target: 10 µm Titanium
PHELIX
E = 88 J
W

## Slide 10: Challenges in TOF: EMP and Dynamic Range
Broadband EMP spectra
Associated electric field at 30 cm: 25 kV/m
Coupling with the detection system
Coupling with the acquisition system
τ = 25 fs
Target : 10 µm Al
FLAME
E = 2.5 J
W
M. Salvadori et al., JINST 15, C10002 (2020)
F. Consoli et al., Scientific Reports 9, 8551 (2019)
F.Consoli et al. High-Power Laser Science and Engineering 8, e22 2020)

## Slide 11: Detection system optimization
Developing a proper shielding
Compact metallic enclosure: exploting skin effect
resistivity; permeability
The 2 mm Al enclosure attenuate EM waves with frequencies down to 2 KHz
Skin depth for aluminum
M. Salvadori et al. Scientific Reports 11, 3071 (2021)

## Slide 12: Detection system optimization
40 mm Aluminum pipe
Exploiting the attenuation provided by waveguides
1. Cutoff frequency
2. Attenuation of the propagating EM waves:
M. Salvadori et al. Scientific Reports 11, 3071 (2021)
Geometrical parameters
Skin depth
Material conductivity

## Slide 13: Acquisition system optimization
Long double shielded coaxial cables
  Provide high frequency attenuation,
  Introduce a temporal delay between the TOF signal and the EMP contribution,
  Allow to place the scope far from the experimental chamber.
Scope with minimal bandwidth to reduce coupling with high frequency electromagnetic pulses.
Scattering parameter S21 Transmission coefficient
M. Salvadori et al. Scientific Reports 11, 3071 (2021)

## Slide 14: Dynamic range enhancement
Where is the photopeak?
Need a higher dynamic range of the acquisition system
Raw signal: no numerical filtering technique is used!

## Slide 15: Dynamic range enhancement
M. Salvadori et al. Scientific Reports 11, 3071 (2021)

## Slide 16: The spectrum reconstruction
PARTICLE
SPECTRUM
01
SIGNAL
ACQUISITION
02
ENERGY
ESTIMATION
03
NUMBER
COMPUTATION
Acquisition and diagnostic system optimization for EMP robustness
Enhanced dynamic range
Enabling high-repetition rate measurements
Measure the Time-Of-Flight of particles
Compute their velocity
Retrieve their energy
De-embedding procedure to retrieve the amplitude at the source
Characterization of the Detector used for TOF measurements
Particle number estimation from the signal amplitude
M. Salvadori et al. Scientific Reports 11, 3071 (2021)

## Slide 17: 1. De-embedding: signal at source
What we have
What we need
M. Salvadori et al. Scientific Reports 11, 3071 (2021)
Obtained characterizing the transmission line by means of a Network Analyzer

## Slide 18: Detector characterization
The detectors that can be used in the time-of-flight scheme can be grouped in three categories:
Electrostatic detectors (i.e. Faraday Cups)
Scintillators coupled to vacuum phototubes
Semiconductor detectors
Diamond Detectors
Wide bandgap (5.5 eV)
High charge carrier mobility
High breakdown field
Not sensible to visible and IR light High radiation hardness
Fast time response
Low leakage current

## Slide 19: 2. Detector characterization: Diamond detectors
1. Production of free electron-hole pairs ( = 13.1 eV).
E
(> 1 MV/m)
2. Charge collection at the electrodes
3. Signal Generation according to the collected charge
Charge Collection Efficiency
Carriers drift velocity
mobility, Electric field
C. Verona et al. , J. Appl. Phys. 118 (2015) 184503,
R. Sussmann CVD diamond for electronic devices and sensors, Wiley and Sons (2009)

## Slide 20: 2. Detector characterization: Diamond detectors
1. Production of free electron-hole pairs ( = 13.1 eV).
E
(> 1 MV/m)
2. Charge collection at the electrodes
3. Signal Generation according to the collected charge
Charge Collection Efficiency
Carriers drift velocity
mobility, Electric field
C. Verona et al. , J. Appl. Phys. 118 (2015) 184503,
R. Sussmann CVD diamond for electronic devices and sensors, Wiley and Sons (2009)

## Slide 21: 2. Detector characterization: Diamond detectors
C. Verona et al., JINST 15, C09066 (2020)
50 µm thickness: Protons up to 3 MeV are completely stopped inside the detector, higher energies require a correction factor.
Different electrodes layout for different needs!
Single crystal structure: 1. Allow to have high CCE 2. Allow to have fast temporal response 3.Small area detectors

## Slide 22: 2. Detector characterization: Diamond detectors
15 mm
15 mm
Wide active surface:
  Enhanced sensitivity
  More susceptible to direct EMP coupling
150 µm thicnkess:
  Suitable for protons up to 6 MeV
  For higher energies a correction factor must be introduced
Polycristyalline structure:
  Allow to have large area detectors
  CCE lower than single crystal structures
  Lower temporal response than single crystal structures
Type: IIa-Electronic grade

## Slide 23
2. Detector characterization: Diamond detectors
The detector assembly mounted on the PHELIX chamber
ions
Diamond
detector

## Slide 24: 2. Detector characterization: Diamond detectors
FWHM = 3.4 ns
CCE ≈ 30 %

## Slide 25: 2. Detector characterization: Diamond detectors
FWHM = 520 ps
FWHM = 460 ps
CCE = 70 %
CCE = 98 %
C. Verona et al., JINST 15, C09066 (2020)

## Slide 26: 3. Particle number: the final spectrum
FWHM
.

## Slide 27: 3. Particle number: the final spectrum
Proton spectrum
Application of suitable filters helps to down extend the region of pure proton contribution
Pure proton contribution
Energy range with pure proton contribution:
Ions accelerated by a potential
τ = 25 fs
Target : 10 µm Al
FLAME
E = 2.5 J
W

## Slide 28: Experiment at PHELIX
10 µm Al filter
10 µm Al filter
20 µm Al filter
FWHM = 3.4 ns
CCE ≈ 30 %
15 × 15 mm2 area
τ = 750 fs
Target: 300 µm TAC Foam 2 mg/cm3
PHELIX
E = 88 J
W
2 × 2 mm2 area
FWHM = 520 ps
CCE ≈ 70 %
FWHM = 520 ps
CCE ≈ 70 %
2 × 2 mm2 area
Provide information on
the EMP level

## Slide 29: Experiment at PHELIX
Broadband spectra
Associated estimated electric field at ≈30 cm : 250 kV/m

## Slide 30: Experiment at PHELIX

## Slide 31: Experiment at PHELIX

## Slide 32: Experiment at PHELIX
Detector inside the experimental chamber 60 cm from target

## Slide 33: Talk summary
An advanced TOF diagnostic methodology was developed to cope with the high level of EMPs typically produced during laser-plasma interaction enabling the possibility to work at high repetition rate
The enhanced dynamic range of the methodology allow to appreciate the whole dynamic range of the signal resulting in an optimal identification of the minimum TOF delay
The technique was coupled with diamond detectors characterized by a wide band gap, a high charge carrier mobility and a fast temporal response resulting in measurement with a high energy resolution
Knowing the detector characteristics and the type of accelerated particles, the methodology allows for the reconstruction of a calibrated spectrum.
It has been shown how the advanced technique has been employed during an experiment performed at the PHELIX laser facility (GSI). Successful tests were also carried out in other environments strongly affected by EMPs :
  PALS laser facility (up to 600 J, hundreds of ps pulse) High fluxes moderate energy ions
  CLPU – Vega 3 (up to 25 J, tens of fs pulse) High energy ions with low fluxes

## Slide 34
Thank you for your attention
Any questions?
Martina.salvadori@uniroma1.it
This work has been carried out within the framework of the EUROfusion Consortium and has received funding from the Euratom research and training programme 2014-2018 and 2019-2020 under agreement No 633053. The views and opinions expressed herein do not necessarily reflect those of the European Commission

## Slide 35: Back up – filter application
Al filter cut
SHOT #5
EL on Target = 2.65 J
Al filter cut
According to the performed SRIM simulations, (10±1.5) µm aluminium filter stops protons up to 750 keV and heavier ions.

## Slide 36: Back up – CCE calibration
Effect of a thickness-dependent electric field on CCE

## Slide 37: Back up – IEC calibration
Effect of a thickness-dependent electric field on CCE

## Slide 38: Back up – Correction factor (1)
Without thickness limit:
The whole energy is deposited in the detector
This amount of charge is generated
150 thick diamond layer:
Only a portion of the energy is deposited
This amount of charge is generated
2 MeV
5 MeV
6 MeV

## Slide 39: Back up – Correction factor (2)
The whole energy is deposited in the detector
Only a portion of the energy is deposited in the detector