## Slide 1: Quality of Amplification Foils (Cond’ive & Res’ive)
Top (Cu/DLC)
Bottom (PI)
Conductive Foil
Resistive Foil (1st prod)
Resistive Foil (2nd prod)
75-80um
um
um
um
78um
um
2019
2020
2020

## Slide 2: Observations / Concerns
Cu electrode (July 2020)
  Smooth surface, hole quality is great, grain-size < 1um
First production DLC foil (July 2019)
  Smooth surface, hole quality is great, grain-size < 10um
  Very circular bottom holes
2nd production DLC foil (July 2020)
  Canyon-like surface, no grains visible, distored holes
  Distored bottom holes
Check USTC details of 2019 and 2020 FTM foil production
DLC Foil Production Procedure: Breaking of DLC in last step is clean / high quality ?

## Slide 3: Top (DLC)
2
1
3
4
5
6
7
8
Bottom (PI)
1
2
3
4
5
6
7
8
also inside detector
also inside detector
78um
100um
108um
87um
69um
93um
99um
100um
75um
69um
105um
88um
80um
60um
81um
inside detector
92um
inside detector
2nd Foil prod (2020)

## Slide 4: Top (DLC)
2
1
3
4
5
6
7
8
Bottom (PI)
1
2
3
4
5
6
7
8
also inside detector
also inside detector
Second Foil inside detector
2
Second Foil inside detector
no meas
90um
66um
90um
70um
75um
93um
100um
124um
83um
88um
57um
Second Measurement
86um
97um
inside detector
inside detector
2nd Foil prod (2020)

## Slide 5: Top (DLC)
2
1
3
4
3
3
4
4
Bottom (PI)
1
4
3
3
4
8
also inside detector
no meas
90um
66um
60um
58um
105um
104um
105um
106um
59um
61um
86um
97um
inside detector
inside detector
1st Foil prod (2019)
um
um
1
1
2
Second Foil inside detector
2
also inside detector
Second Foil inside detector
inside detector
inside detector

## Slide 6: Selection of Best Foil
 | Foil 1 | Foil 2 | Foil 3 | Foil 4 | Foil 5 | Foil 6 | Foil 7 | Foil 8
Top (1) | 81 | 78 | 87 | 69 | 69 | 75 | 80 | 60
Top (2) | 92 | - | 90 | 66 | 70 | 75 | 88 | 57
Bot (1) | 86 | 100 | 108 | 93 | 99 | 100 | 105 | 88
Bot (2) | 97 | - | - | 90 | 93 | 100 | 124 | 83
Rank | ? | 5 | 7 | 2 | 3 | 4 | 6 | 1
 | Foil 1 | Foil 2 | Foil 3 | Foil 4 | Foil 5 | Foil 6 | Foil 7 | Foil 8
B/T (1) | 1,14 | 1,28 | 1,24 | 1,35 | 1,43 | 1,33 | 1,31 | 1,47
B/T (2) | 1,13 | - | - | 1,36 | 1,33 | 1,33 | 1,41 | 1,45
Rank |  |  |  |  |  |  |  | 

## Slide 7: Gain Depends on Top & Bot diam
Simulation
Measurement
Cu Foil: top: 75-80um, bottom: 45um
DLC Foil: top: 78um, bottom: 100um Gain factor 2 higher for DLC foil
Cu Foil: top: 70um, bottom: 50um
DLC Foil: top: 50um, bottom: 70um
Gain factor 2 higher for small top diameter

## Slide 8: Gain: Old vs New DLC Foil
um
78um
Foil prod 2019 – Test 2020
Foil prod 2020 – Test 2021

## Slide 11: Paschen Curve
Ar:CO2 70:30
Air
Ne:C2H6 90:10
Ne:CF4 90:10

## Slide 12: FTM Foil production procedure 1 – The actual procedure
Starting FCCL: single DLC layer
Photo-Lithography & Cu Etching
Kapton Etching
Cu Etching & Cr removal (and DLC partially goes away)
Clean DLC with water jet & pummice  suggestion: try with plasma etching
1
2
3
4
5
Layer | Thickness | Color code
Cu | 7um | 
Cr | 10nm | 
DLC | 100nm | 
Polyimide | 50um | 
Cu | 5um | 

## Slide 13: FTM Foil production procedure 2 – With Delaminating DLC
Photo-Lithography & Cu Etching
(a) Kapton Etching: etching (b) Delamination => more etching (c) Kapton Etching: remove PI
Cu Etching & Cr removal
Remove DLC with water jet  try with plasma etching
Delamination allows the etching liquid to attack more Polyimide than foreseen by the mask, in an irregular pattern, resulting in irregular hole shape both in top (DLC) as bottom (PI).
1
2b
3
4
Layer | Thickness | Color code
Cu | 7um | 
Cr | 10nm | 
DLC | 100nm | 
Polyimide | 50um | 
Cu | 5um | 
2c
2a

## Slide 14: FTM Foil production procedure 3 – Suggestion & Future: Create Rim
Starting FCCL: single DLC layer
Photo-Lithography & Cu Etching
Kapton Etching
Cu Patterning on top (larger), leave Cu as protection on the bottom
Etch / Remove DLC (scrubbing: water jet with small pummice) or with Plasma Etching
Cu & Cr removal
Procedure can be scaled up to 125um polyimide, inserting a 2nd kapton etching step, starting from the top
1
2
3
4
5
Layer | Thickness | Color code
Cu | 7um | 
Cr | 10nm | 
DLC | 100nm | 
Polyimide | 50um | 
Cu | 5um | 

## Slide 15: Res GEM Foil production procedure
1
2
3
4
5
6
Layer | Thickness | Color code
Cu | 5um | 
Cr | 10nm | 
DLC | 100nm | 
Polyimide | 50um | 
DLC | 100nm | 
Cr | 10nm | 
Cu | 5um | 
Starting FCCL: double DLC layer
Photo-Lithography & Cu-Cr Etching
DLC “breaking” (sand blasting - pummice)
Kapton Etching
DLC “breaking” (sand blasting)
Cu Etching & Cr removal

## Slide 16: FTM Foil production procedure 4 - Future
Starting FCCL: single DLC layer
Photo-Lithography & Cu-Cr Etching
DLC “breaking” (sand blasting?)
Kapton Etching
DLC “breaking” (sand blasting?)
Cu Etching & Cr removal
1
2
3
4
5
6
Layer | Thickness | Color code
Cu | 7um | 
Cr | 10nm | 
DLC | 100nm | 
Polyimide | 50um | 
Cu | 5um | 

## Slide 17: Alternative Detection Structures to be considered?
Picosec micromegas (with pre-amp) arrive at few 105 – 106 with pre-amplification. Assume pre-amp of factor 10, then MM in Compass gas (Ne:CF4:C2H6 80:10:10) still has G = few 104 – 105
RETGEM shows gain 105 (Peskov) 104 (Confirmed Chinese Colleagues)
REGEM (Japan) however not successful
Possibility to make Micromegas with Resistive Mesh? E.g. Glass Fiber mesh coated with DLC?
  Polyester mesh to be coated with DLC is a possibility …
Possibility to make RETGEM without Cu layer on top of PCB. PCB of 0.2 – 0.4 – 1mm ?
Increased thickness of amplification stage will result in reduced Capacitance (C=eA/d), and hence reduced transparency / signal induction

## Slide 19: For the Graphene Collaboration