URANIA-V
10-14 Maggio 2021
CHI
COSA
W1 – planare – cremat 2ch
W6 – grooved – cremat 1ch
W8 – grooved – caen 1ch
W6 – grooved 6mm – cremat 2ch
THR SCAN @ different gains
COME
Misurando allo scaler i conteggi in 10s:
- 1 canale discriminato e formato a 20us
- 1 canale discriminato e formato a 20us (attivo solonel caso del cremat a 2 ch)
- 1 OR dei due canali, ri-formato a 35us

---
PLANARE 2 CREMAT
HV SIDE
mV
# (log)
NO HV SIDE
# (log)

---
PLANARE 2 CREMAT
OR planare 2cremat
# (log)
Subtraction
# (log)

---
GROOVED 1 CREMAT
HV SIDE
# (log)

---
GROOVED 1 CREMAT
OR grooved 1 cremat
# (log)
sub grooved 1 cremat
# (log)

---
GROOVED 1 CAEN cameretta non calibrata in HV
HV SIDE
# (log)
OR grooved caen
# (log)

---
GROOVED 2 CREMAT 6 + 3 mm
HV SIDE
# (log)
NO HV SIDE
# (log)

---
GROOVED 2 CREMAT 6 + 3 mm
OR grooved 2 cremat
# (log)
sub grooved 2 cremat
# (log)

---
CONFRONTO GAIN 500
PLANARE 2CREMAT / GROOVED 1CREMAT / GROOVED 2CREMAT6+3
#

---
CONFRONTO GAIN 1000
PLANARE 2CREMAT / GROOVED 1CREMAT / GROOVED 2CREMAT6+3
#

---
RATE - CONFRONTO
PLANARE 2CREMAT / GROOVED 1CREMAT / GROOVED 2CREMAT6+3
Hz
500
Hz
1000

---
CONFRONTO GAIN 500
GROOVED 1CAEN / GROOVED 1CREMAT / GROOVED 2CREMAT6+3
#

---
CONFRONTO GAIN 1000
GROOVED 1CAEN / GROOVED 1CREMAT / GROOVED 2CREMAT6+3
#

---
RATE - CONFRONTO
GROOVED 1CAEN / GROOVED 1CREMAT / GROOVED 2CREMAT6+3
Hz
500
Hz
1000

---
GROOVED 1CAEN / GROOVED 1CREMAT / GROOVED 2CREMAT6+3
Hz
G500
PLANARE 2CREMAT / GROOVED 1CREMAT / GROOVED 2CREMAT6+3
Hz
G500
Planare conteggia circa ¼ in meno; Grooved 1CAEN non era ben calibrata in HV;
1 o 2 cremat non sembrano cambiare significativamente; Conteggi al plateaux per grooved (no caen) ~2000

---
RATE ATTESO
Planare conteggia circa ¼ in meno;
Grooved 1CAEN non era ben calibrata in HV;
1 o 2 cremat non sembrano cambiare significativamente;
Conteggi al plateaux per grooved (no caen) ~2000
Sorgente 750 Hz/cm2
1% di efficienza = 750Hz
75000 Hz
Area Attiva 100 cm2
Efficienza ~3%
Nostri Conteggi GROOVED ~ 2000Hz
Efficienza ~2%
Nostri Conteggi PLANARE ~ 1500Hz

---
CONFRONTO GAIN 500
PLANARE 2CREMAT / GROOVED 1CREMAT / GROOVED 2CREMAT6+3
#

---
CONFRONTO GAIN 500
GROOVED 1CAEN / GROOVED 1CREMAT / GROOVED 2CREMAT6+3
#

---
CONFRONTO GAIN 1000
PLANARE 2CREMAT / GROOVED 1CREMAT / GROOVED 2CREMAT6+3
#

---
CONFRONTO GAIN 1000
GROOVED 1CAEN / GROOVED 1CREMAT / GROOVED 2CREMAT6+3
#