## Slide 1
MRI-guided PET Reconstruction with Adaptive Prior Strength
Jorge Cabello1, Udunna Anazodo2
1Siemens Medical Solutions USA, Inc.
2Lawson Health Research Institute, London, Ontario, Canada and Research Centre for Studies in Aging, McGill University, Montreal, Quebec, Canada

## Slide 2: Siemens Medical Solutions, Knoxville
Jorge Cabello, PhD

## Slide 3: Introduction
PET suffers from limited spatial resolution and noise, partly due (but not limited) to crystal size, injected dose and scan time, while anatomical MRI has high spatial resolution and soft tissue contrast.
The combination of MRI in the PET reconstruction is a method to regularize the PET reconstruction inverse problem, under the general assumption that similar voxels in MRI are similar in the PET image.
PET and MRI data are combined in the objective function, where the influence of the MRI is controlled by a hyper-parameter (β), which is usually fixed for every voxel and iteration, and tweaked for each dataset (patient).
This work presents a method that adaptively calculates β, and out-performs regular OSEM for a wide range of statistics and different radiotracers, applied to brain PET.

## Slide 4: Methods
The prior model, written of the form
was based on the Bowsher model, where only voxels similar in the MRI are considered in the calculation of the prior.
The similarity weight was based on the Burg joint entropy1:
The hyper-parameter β2 is calculated for each voxel j, at each iterations, as
where α is a scanner-dependent constant, and uj is the intensity in voxel j, making β dependent on the statistical noise in that voxel.
Distance weight
Similarity weight
Potential function
PET
MRI
1A. Mehranian et al., PMB (2017) 62 (15) 5975-6007.
2A. Salomon et al., arxiv Physics (2019).
System model, norm and attenuation

## Slide 5: Materials
The method was evaluated with simulated and experimental data.
Simulated data:
Experimental data:
Regional analysis was performed using Statistical Parametric Mapping2.
1https://brainweb.bic.mni.mcgill.ca/
2https://www.fil.ion.ucl.ac.uk/spm/software/spm12/
Data were generated using a normal case from the BrainWeb brain database1 and GATE (v8.1).
The phantom for emission data, attenuation map and T1-weighted were calculated.
22.6 MBq activity of a 5- and 1-minute scans were simulated, resulting in 152x106 and 30x106 trues, respectively.
PET and T1-weighted MRI were acquired on a Biograph mMR.
  30 min [18F]-FDG brain scan of 1 patient with suspected epilepsy (200 MBq, 40min p.i)
  20 min [18F]-florbetapir brain scan of 2 patients with suspected Alzheimer's disease (~370MBq, 50 min p.i)
All scans were re-framed to 5- and 1-minute scans.

## Slide 6: Results – simulated data
MRI-guided β=102
MRI-guided adaptive β
OSEM
5-min scan
152x106 trues
1-min scan
30x106 trues
57.0
0.0
[kBq/mL]

## Slide 7: Results – simulated data
Lesions only present in the PET data but absent in the MRI data
5-min scan
1-min scan
MRI-guided β=102
MRI-guided adaptive β
MRI-guided adaptive β
Lesions in T1w
Lesions NOT in T1w
50.0
0.0
[kBq/mL]

## Slide 8: Results – [18F]-FDG Epilepsy case
Data from Lawson Health Research Institute, London, Ontario, Canada
30.0
0.0
[kBq/mL]
[kBq/mL]
[kBq/mL]
[kBq/mL]
30 min
5.4 108 trues
10 min
9.6 107 trues
1 min
1.9 107 trues
OSEM
MRI-guided
Adaptive β
MRI-guided
Fixed β (102)
PSF
3 iterations
21 subsets

## Slide 9: Results – [18F]-FDG Epilepsy case
30.0
0.0
[kBq/mL]
[kBq/mL]
[kBq/mL]
[kBq/mL]
30 min
5.4 108 trues
10 min
9.6 107 trues
1 min
1.9 107 trues
OSEM
MRI-guided
Adaptive β
PSF
10 iterations
21 subsets
Data from Lawson Health Research Institute, London, Ontario, Canada

## Slide 10: ROI analysis
Cerebellum
Hippocampus
Ventricles
30-min scan
5-min scan
1-min scan
Relative difference between OSEM and MRI-guided for the 30-, 5- and 1-min scans.
Mean intensity vs noise
Data from Lawson Health Research Institute, London, Ontario, Canada

## Slide 11: β-amyloid imaging - [18F]-florbetapir
Data courtesy of Hongyu An from the Mallinckrodt Institute of Radiology, Washington University School of Medicine (St Louis, MO,USA)
Positive amyloid scan
Negative amyloid scan
Patients injected ~375 MBq, scanned for 20 min, 50 min p.i.

## Slide 12: β-amyloid imaging - [18F]-florbetapir VOI analysis
Data courtesy of Hongyu An from the Mallinckrodt Institute of Radiology, Washington University School of Medicine (St Louis, MO,USA)
Regions subject to show β-amyloid deposition.
Taking the 20 min scan as reference, relative difference for the 5- and 1-min scans.
Amyloid+ scan
Amyloid- scan

## Slide 13: PET Quantification - Epilepsy
Poirier et al EJHI 2020; Poirier et al Epilepsy Research 2021

## Slide 14: PET Quantification - Epilepsy
Poirier et al EJHI 2020; Poirier et al Epilepsy Research 2021; Raymond et al Brain & Brain PET 2022
20 Epilepsy Patients

## Slide 15: PET Quantification - Epilepsy
Poirier et al EJHI 2020; Poirier et al Epilepsy Research 2021; Raymond et al Brain & Brain PET 2022
T1+ZAI
20 Epilepsy Patients

## Slide 16: Conclusions
Fixed β showed how the MRI was overly enforced on the PET data for the different levels of statistics, producing in some cases artefacts, noise and loss of PET-only features. The adaptive β did not show this problem at the cost of the MRI not being so enforced.
The clinical cases showed how the usual acquisition time of 20/30 min could be reduced to 5 min without any visible loss of image quality (which could be translated into dose reduction).
The method not only showed to work for more than FDG, but also for activity distributions where the PET does not just follow one type of tissue according to the MRI.
Clinical data showed that differences between OSEM and MRI-guided are <10%. Same comparison with simulated data, comparing to ground truth, is on the way.
Notes: 3rd point. PET shows uniform distribution across white and gray matter in the amyloid case, even though MRI shows a clear division.