## Slide 1: real-time 3D dose imaging and source localization in brachytherapy
A. Giaz1, N. Amplilogov1, S. Cometti1, W. Kam3, M. Martyn2, R. Santoro1, M. Caccia1 and S. O’Keeffe3
1Università degli Studi dell’Insubria, DiSAT, via Valleggio, 11, Como, Italy
2Department of Radiotherapy Physics, Galway Clinic, Galway, Ireland
3Optical Fibre Sensors Research Centre, University of Limerick, Limerick V94 T9PX, Ireland

## Slide 2: Oncological brachytherapy
Radiotherapy is the use of ionizing radiation for treatment. It is delivered into external beam radiotherapy using accelerators, or internal, using sources, known as brachytherapy (BT).
Brachytherapy is a form of radiation therapy where a sealed radiation source is placed inside or next to the cancer area. The source placement is vital to ensure radiation to the tumor, while ensuring minimum exposure to nearby critical organs.
Brachytherapy is divided into Low Dose Rate (LDR) and High Dose Rate (HDR).
 | Source | Activity | Implantation | <Eγ> | Abs. Length
LDR | 125I | 15 MBq | Permanent | 35 keV | 3 cm
HDR | 192Ir | ∼100 GBq | Temporary | 380 keV | 10 cm
prostate cancer
gynecological cancer

## Slide 3: THE ORIGIN PROJECT
The ORIGIN project is an initiative of the Photonics Public Private Partnership (www.photonics21.org), and has received funding from the European Union's Horizon 2020 Research and Innovation Programme under Grant Agreement n° 871324
ORIGIN CONSORTIUM
PI from Limerick University
Optical Fibre Dose Imaging for Adaptive Brachytherapy
ORIGIN aims to deliver more effective, photonics-enabled, brachytherapy for cancer treatment through advanced real-time radiation dose imaging and source localisation. This will be achieved by the development of a new optical fibre-based sensor system to support diagnostics-driven therapy through enhanced adaptive brachytherapy.

## Slide 4: THE ORIGIN PROJECT
Optical Fibre Dose Imaging for Adaptive Brachytherapy
ORIGIN aims to deliver more effective, photonics-enabled, brachytherapy for cancer treatment through advanced real-time radiation dose imaging and source localisation. This will be achieved by the development of a new optical fibre-based sensor system to support diagnostics-driven therapy through enhanced adaptive brachytherapy.
 | LDR - BT | HDR - BT
Sensitivity to Dose | up to 3 cm | up to 10 cm
Spatial Resolution | 3 mm @ 3 cm | 1 mm @ 5 cm
Dose Rate Range | 1 mGy/s up to 15 Gy/s | 
Statistical Precision | 5% in 0.5 s | 5% in 0.1 s
At least 16 dosimeters are need to localize the source.

## Slide 5: THE ORIGIN PROJECT: Dosimeter System
The project’s goal will be achieved by developing a 16 optical fiber-based system with scintillating light detected by SiPM to reconstruct the dose map.
Sensor: 1x1 mm2 Ketek
The signal associated with an X-ray or 𝛾-ray interaction consists of a trail of single photons, and the dosimetry is based on photon counting.
Scintillator | Gadox (LDR) | YVO (HDR)
𝜏 [𝜇s] * | 500 | 500
𝜆max [nm] | 545 | 600 - 650
LY [ph/MeV]* | 7.1 ∙ 104 | 4.8 ∙ 104
Transmittance* | 10.2 % | 4.2 %
*S. Cometti et al. “Characterization of scintillating materials in use for brachytherapy fiber based dosimeters”, accepted for publication to NIMA
1 p.e.
2 p.e.
3 p.e.
PDE @ 600 nm: 14 %
DCR @ 0.5 p.e.: 100 kHZ
OPT: 26%
Op. Voltage: Vbd + 7 V

## Slide 6: Single Channel Measurements
From single channel measurement was possible to extract all the specification of the project.
The results prove that the single fibre dosimeter system complies with the specifications for both the HDR-BT and LDR-BT.
The data follow the Treatment Planning System (TPS)
HDR
LDR
Sensitivity to Dose: 3 cm
Spatial Resolution: 1.4 mm
Statistical Precision: 2.6 % in 0.5 s
Sensitivity to Dose: 10 cm
Spatial Resolution: 0.8 mm
Statistical Precision: 3 % in 0.1 s

## Slide 7: THE 16 Channel readout system
64-channel SiPM readout board (CAEN* FERS) equipped with 2 WEEROC CITIROC1A ASICs**
150 mm
60 mm
Simplified Block-schema
Single p.e. counting capability
Maximum counting rate: 20 MHz
1 HV power supply (20 – 100V) with temperature compensation
Ethernet, usb2 and optical link interface for readout (up to 6.25 Gbit/s)
*https://www.caen.it/products/a5202/
**https://www.weeroc.com/products/sipm-read-out/citiroc-1a

## Slide 8: Laboratory-qualification
An automatic procedure is used to estimate the VBreakDown (VBD) for all SiPMs.
V BIAS adjustment allows to equalize the 𝛥pp and to operate all SiPMs at the same PDE (V BIAS = VBD + 7 V).
Staircases acquired at V BIAS to select 0.5 p.e. and 1.5 p.e. threshold.
Typical Staircase
1.5 p.e.
0.5 p.e.
Baseline crossing
Same Threshold on all channels

## Slide 9: Laboratory-qualification
Measurements with X-ray cabinet to evaluated the uncertainties:
Fiber positioning: 0.8%
Fiber connection: 0.1%
Geometrical acceptance: 10 %
Raw Data
Fiber Non-Uniformity: 16 % (pre-production) 1.2 % (new production)
non-uniformity: 16%
non-uniformity: < 1%
1 fiber
+
16 channels
16 fibers
1 channel
Normalization with a scaling factor including all the contributions

## Slide 10: EQUALIZATION Measurement
Water-phantom with circular geometry
All fibres at the same distance from the source
Measured counting rate normalised at mean value
Run to check equalization procedure; systematics variation less than 1%
Non – Uniformity: < 1 %
Non – Uniformity: 17 %

## Slide 11: EQUALIZATION Measurement
Water-phantom with circular geometry
All fibres at the same distance from the source
Measured counting rate normalised at mean value
Run to check equalization procedure; systematics variation less than 1%
Non – Uniformity: < 1 %
The normalization scaling factor was check at different distances: 5 cm

## Slide 12: EQUALIZATION Measurement
Water-phantom with circular geometry
All fibres at the same distance from the source
Measured counting rate normalised at mean value
Run to check equalization procedure; systematics variation less than 1%
Non – Uniformity: < 1 %
The normalization scaling factor was check at different distances: 10 cm

## Slide 13: CLINICAL MEASUREMENTS
16-channel system, with dosimeters placed around a 192Ir radioactive at the Queen’s University Hospital (Belfast, Ireland).
z = 0 cm
z = 3 cm
z = 10 cm
Step: 1 cm
Step: 0.2 cm
x
y
z
11 Fibres:
3 fibres at r = 4 cm,
3 fibres at r = 2.5 cm,
2 fibres at r =1.5 cm,
3 fibres at r = 1 cm.

## Slide 14: CLINICAL MEASUREMENT Results
The source is moved from a 10 cm distance on the z axis down to z = 0 cm and the corresponding counting rate as a function of distance was measured.

## Slide 15: CLINICAL MEASUREMENT Results
The sensors are capable of monitoring the radiation dose at distances compatible with brachytherapy.
Data has to be compared with TPS.
One Fiber (r = 2.5 cm): pile-up corrected
Trend ∝1/r2
Last 3 points excluded
𝜸 Att. Coeff. at 300 keV in water ∼ 6 cm
Tree Fibres: (r = 4 cm, r = 2.5 cm, r =1.5 cm): pile-up corrected - DCR subtracted
Tree Fibres: (r = 4 cm, r = 2.5 cm, r =1.5 cm): pile-up corrected - DCR subtracted
After Equalization
Preliminary
Preliminary

## Slide 16: Source localization
Source localization has been implemented using triangulation algorithm. Initial results looks promising.
A quantitative assessment was not completed yet.
Preliminary
Preliminary
z = 0 cm
z = 3 cm

## Slide 17: OUTLOOK
The laboratory characterization shows that the uncertainties of the system are less than 1%.
The clinical HDR characterization of the ORIGIN dosimeter shows the expected trends. The data has to be compared with TPS.
3D-printed semi-anatomical phantom prototypes were developed to provide be used in the evaluation of the dosimetry system.
will be integration of the ORIGIN dosimeter prototype in a clinical systems started.

## Slide 18
Contact Us ORIGIN@ul.ie
@ORIGIN_2020
ORIGIN2020.eu