## Slide 1: IHEP SCD simulation
Rui Qiao; WenXi Peng; Ke Gong

## Slide 2: background
SCD should measure Z=1~26， large dynamic range.
Two requirements:
  Large dynamic range of ASIC (VA140)
  Optimize charge sharing in detector
charge sharing
charge sharing ratios of AGILE
From G. Barbiellini, NIMA 490 (2002) 146

## Slide 3: The charge sharing of AMS & DAMPE
AMS
DAMPE
 | Ratio A1->A2
AMS | ~40%
DAMPE @beamtest | <4%
Notes: 来源：G:\keyan\科研任务\211212_硅微条模拟\211212_硅微条模拟-软件比较.pptx

## Slide 4: The in-flight calibration of DAMPE charge sharing ratio: A1->A2
Parameter | Selection
Angle (perpendicular strip) | <5 deg
Angle (along strip) | <15 deg
Impact Position | ETA<0.05 ¦¦ ETA>0.95
Heavy ions | A1> 2200 ADC
No Saturation | A1< 2800 ADC
A1 (ADC)
A2/A1
Z=6
Z=5
Z=8
Z=7
A2/A1
Ratio of A1->A2 is ~0.5%
Notes: 来源：G:\keyan\科研任务\180115_轻粒子电荷\180301_用重核刻度更远的电荷分配系数\4 AnaLonger_调整策略后的2年的信号谱\TH2I_R2R0_vsR0_Lad000.png
G:\keyan\科研任务\180115_轻粒子电荷\180301_用重核刻度更远的电荷分配系数\5 FitLonger_调整策略后的2年的信号谱\TH1I_R2_R0.pdf

## Slide 5: Problems to solve…
The charge sharing (0.5%) of DAMPE is too weak. When strip A1 encounter saturation (~3000 ADC)， the amplitude of strip A2 is too small (~15 ADC).
We need to design a SCD with larger charge sharing ratios.

## Slide 6: Simulation tools
Name | Source | 运行环境
Sentaurus TCAD | Commercial | Windows
Silvaco TCAD |  | 
Allpix | Open-source；CERN | Linux；C++
KDetSim |  | 
TRACS |  | Linux；GUI
WeightField | Open-source；INFN | 
The four open-source simulation tools cannot simulate the capacitive charge sharing.
The commercial TCAD can simulate the capacitance.

## Slide 7: The TCAD simulation
Parameters | Values |  | 
 | DAMPE | IHEP HERD(20um) | IHEP HERD(60um)
Thickness (um) | 320 |  | 
implant Pitch (um) | 121 | 80 | 
Readout pitch (um) | 242 | 160 | 
implant width (um) | 48 | 20 | 60
Al width (um) | 56 | 28 | 68
concentration | 50V full depletion |  | 
Sio2 thickness | 520 pF coupling |  | 
Implant thickness (um) | 4.4 pF interstrip cap. |  | 

## Slide 8: The TCAD simulation
DAMPE design with 5 strips
Thickness 320um
Implant pitch 121um
Implant width 48um
Al width 56um
electrode
doping
Notes: 来源：G:\keyan\科研任务\211212_硅微条模拟\MyTCAD\220107——DAMPE 5个注入条\

## Slide 9: Simulation Results
Result: CV curve
Red: interstrip cap.
Green: backplane cap.
Full depletion: 50V
Det_Name | Interstrip cap. (pF) | Coupling cap. (pF) | Backplane cap. (pF)
DAMPE_Exp | 4.4 | 520 | 4.3
DAMPE_Sim | 4.4 | 517 | 3.9
HERD_20um | 3.3 | 255 | 2.5
HERD_60um | 8.8 | 651 | 2.6
Notes: 来源：G:\keyan\科研任务\211212_硅微条模拟\MyTCAD\220110——HERD\电容值统计.xlsx

## Slide 10: Spice simulation
The TCAD capacitances are inserted into SPICE.
Pulse current source generate charges
Probes measure amplitutde of A1 & A2
Spice simulation of DAMPE charge sharing.
Det_Name | Charge sharing ratio: A1->A2
DAMPE_Exp | 0.28%
DAMPE_Inflight | ~0.5%
HERD_20um | 0.47%
HERD_60um | 0.58%

## Slide 11: Conclusion
Capacitive charge sharing is reported by AGLIE, AMS-02, DAMPE
When strip A1 is saturated, we should use strip A2/A3 for reconstruction.
The DAMPE charge sharing ration A2/A1 is small (0.4%), compared to AMS-02(~40%) and AGILE(10%). We should optimize SCD with larger charge sharing ratio.
TCAD+SPICE simulation was implemented for DAMPE & IHEP SCD. The ratio is still <1%.

## Slide 12: Backup: AGILE
Notes: G. Barbiellini, NIMA 490 (2002) 146

## Slide 13: Backup: AMS-02
Not hitting readout strip, but d~0.3
Not normal incidence, but 10deg.