## Slide 1
Characterization of Superpix0 chips
Noise scan results
A. Manazza, G. Traversi, L. Ratti
UNIVERSITA’ DEGLI STUDI DI PAVIA
Laboratorio di Strumentazione Elettronica

## Slide 2
SuperPix0 - Fine Pitch Hybrid Pixel Sensor
The prototype chip contains 4096 readout cells of 50 μm × 50 μm arranged in a 32 × 128 matrix and organized in MacroPixels (MP)
We have tested 3 chips with no sensor connected:
  CHIP4
  CHIP6
  CHIP7
SuperPix0 layout
Photograph of the prototype
Cross sectional view of the metal layers used in the layout of the SuperPix0 chip
ALUCAP
CHIP1 could not be tested because some bonding wires were damaged

## Slide 3
SuperPix0 - Analog Cell
PA
IFB
Cinject
Cdetector
CFB
DISC
+
-
VTH
FFD
InjectMaskClock
In-pixel
digital
logic
HIT
Qinject
Q
D
InjectMaskIn
InjectMaskOut
Pixel layout integrated in the SuperPix0 chip
ENC = 90 e- @Cd=0
Q = 66000 e-
ENC = 120 e- @Cd=0
ENC = 120 e- + 0.51 e-/fC * Cd
ENC = 90 e- + 0.55 e-/fC * Cd

## Slide 4
Noise Scan Measurements
We measured the occupancy as a function of the comparator threshold and we fitted the histogram and extracted noise and dispersion parameters with the function:
O(Vth - Vbl) = 1 – exp [- Toss νo exp (- (Vth - Vbl)/2 σ2)]
There are 8 pixels that have larger occupancy than the others. This happens only for pixels in a particular position in the MP
The fastOR and the latch enable lines run below the bump bonding pad and when they change state it is like the pixel is injected
Frequency
at zero threshold
Observation
window (120 μs)
Baseline
voltage
Noise

## Slide 5
Noise scan results
 | Simulations | CHIP 4 | CHIP 6 | CHIP 7
THRESHOLD [mV] | 180 | 208.3 ± 0.2 | 207.3 ± 0.2 | 207.8 ± 0.2
THRESHOLD DISPERSION [mV] | 2.52 (350 e-) | 2.43 ± 0.17 | 2.19 ± 0.15 | 1.98 ± 0.13
NOISE [mV] | 0.63 (90 e-) | 0.51 ± 0.01 | 0.503 ± 0.01 | 0.504 ± 0.01
The threshold dispersion and the ENC parameters are calculated with post-layout charge sensitivity (Gq = 44 mV/fC)
The threshold and noise histograms are referred to chip 4