## Slide 1
Instrumentation and diagnostics for Test Facility Dipole at Fermilab
G. Velev
2nd Instrumentation and Diagnostics for Superconducting Magnets Workshop

## Slide 2: Outline
Introduction - HTS cable test facility at Fermilab
Test Facility Dipole (TFD) from LBNL - 15 T background field
Quench protection and characterization
Instrumentation and diagnostics for the TFD and HTS cable sample
  New multichannel Digital Integrator
  Nanovolt Multiplexer Meter for splice and VT measurements
  New signal conditioner system
  EMMA: Software Framework for Measurement Systems
  Fibers - covered by Maria’s and Steve’s talks
  Quench antennas - Joe’s talk
Summary
Examples:
EMMA – Extensible Magnetic Measurement Application

## Slide 3: Introduction
US DOE Offices of High Energy Physics and Fusion Energy Sciences joined their effort to build HTS cable-testing facility.
This facility is currently in construction at Fermilab.
  For the US FES and HEP, it will serve as a test stand for HTS cables in high dipole field and at increased temperatures.
  It is designed to provide similar or better capabilities than EDIPO at PSI, and FRESCA2 at CERN
  For the US Magnet Development Program, it will be the main facility for testing of magnets with 16-20 plus T fields, including hybrid magnets
  15 T dipole (TFD) - provided by LBNL
For more information - see the talk in the TF workshop

## Slide 4: Test Facility Dipole
Field: 16 T design target with >15% margin on the load line at 1.9 K
  15 T is used as reference for operation
Clear aperture: 144 mm (H) x 94 mm (V) superimposed with 106 mm round for compatibility with FRESCA2 samples
Magnet design is based on past development and studies of block-coil dipoles (HD, FRESCA2 and HEPdipo)
Mechanical design: Rutherford cable, accelerator coil with shell-based structure with axial pre-load
Quench Protection: energy extraction (dump resistor) is sufficient for safe magnet discharge: stored energy ~ 12 MJ
Coupling Loss induced Quench (CLIQ) system is included in the baseline plan for redundancy, no quench heaters
Layer 1: 40 turns
Coil 1
Coil 2
Layer 2: 40 turns
Layer 3: 44 turns
Layer 4: 44 turns
Rmin:
54 mm
D. Arbelaez, P. Ferracin, R. Hafalia, P. Mallon, D. Martins Araujo, J. L. Rudeiros Fernandez, G. Vallone et al.

## Slide 5: Quench Protection Simulation
Studies were performed with STEAM-LEDET code
Conductor: Cu/NonCu: 0.8-1.17 (selected 0.9)
At operation at current corresponding to 15 T
Dump resistor: two options 130 mΩ (2 kV) – 75 mΩ (1 kV)
For redundancy CLIQ is used: 40 mF, 600 V (AUP specs)
Expected Detection + validation time: 15 ms
If quench detection is increased by factor of 3 (45 ms), temperature increases by ~ 30K
Failure to activate dump on time may damage the coils - need reliable QP system
V. Marinozzi

## Slide 6: QP system - Requirements and Specifications
QP developed for the muon-to-electron (mu2e) conversion experiment to protect 3 solenoids with ~105 MJ stored energy
It is designed to minimize false quench triggers and unscheduled downtime that causes unnecessary interruptions to the operation of experiment
The quench detection system monitors the superconducting magnet coils and bus, the HTS power leads, and the vapor-cooled copper power leads
QP system: 3 tiers
  Tier-I – Digital Quench Detection (DQD);
  Tier-II (redundant) – Analog Quench Detection (AQD);
  Tier-III – Quench System Monitoring and Data Manager + Operational GUIs & Quench Analysis
DQD and AQD provide independent quench triggers and designed to minimize the false trips
DQD and AQD features:
  Two operation modes – production and expert modes
  User settable validation times
  Built in signal conditioning: filtering, gains, and attenuation
First application: AUP cryo-assembly testing as we speak

## Slide 7: DQD System
Modular Hardware System:
Minimum number of connections
Built-in high voltage isolation;
Simple FPGA algorithms implemented on lower level (Verilog)
Data is stored in the DQD module SRAM independently from PC
DQD Module:
Digital Isolation
Programable:
  Gains;
  Trigger validation and delay times;
  Multiple current dependent thresholds;
  Channels for bucking;
  Fast quench loggers
  Pre and post trigger data windows
DQD Controller:
Communicates with TIER-III
Provides current ramp rate (dI/dt) for DQD & AQD
Note: The DQD quench detectors are commercially produced
using Class 3 assembly standards (high reliability electronics).
Digital Quench Detector 4 ch board

## Slide 8: Analog Quench Protection
AQD Hardware System Features:
Channel isolation
Ability to perform analog bucking of two near by channels
Ability to buck a channel against magnet current I or dI/dt
Settable QD threshold
Balance adjustment
Selectable gains for each channel (0.25 to 256)
Selectable validation time for each channel
Selectable input filter
Isolated quench output signals
Isolated AQD status outputs
Full chassis of Analog Quench Detector Modules – 28 Channels
AQD Channel/Bucking Circuit
Courtesy of T. Cummings (FNAL)
AQD Module

## Slide 9: TIER-III Software System
Expert Mode: QD Balance
Quench Software Layout
Tier-III is based on National Instruments C-RIO.
Configurable GUI includes a Change of State Monitor, an Event Logger, and launch buttons for Quench Monitoring, Quench Analysis, and Expert Mode.
Tier-III collects quench data pre- and post- event analysis.

## Slide 10: QP Architecture for HTS Cable Test Facility
Quench Detection & Management
Quench Logic
Module & Protection
PS Control and Monitoring
Branch 1
Branch 2
Two branches:
  Branch 1: Magnet protection
  Branch 2: Test Sample protection

## Slide 11: HFVMTF QP Details – Quench Detection & Management
Primary Custom Digital Quench Detection System (DQD) – 96 Channels
Quench management system, based on National Instruments C-RIO (Tier III). Provides quench configuration, real time control and monitoring, and quench data management for offline analysis.
Redundant Custom Analog Quench Detection System (AQD) – 28 Channels
The detectors can monitor:
Temperature sensors
Strain gauges
Hall sensors

## Slide 12: HFVMTF QP Details – Quench Logic & Protection
A set of two redundant custom Quench Logic Modules (QLMs) carry out the critical hardware-based quench logic, such as:
Energy extraction system enable and discharge control via DUMP switches
Heater control logic via Heater Firing Units (HFU)
CLIQ control logic (CLIQ)
Slow Ramp Down request for the Current control
Phase Back control for the power system

## Slide 13
Powering Hi-Lumi test stand to Nominal I (16.3kA), Quench at 17.289kA

## Slide 14: Instrumentation and data acquisition for test sample
On November 21-22, we hosted a workshop with the fusion user community. The users' feedback is summarized as follows:
Provide a standard “base” instrumentation layout for all tests
  User samples: Voltage taps (high-field, joints), Cernoxes, etc.
  Facility: Voltages (e.g. sample-facility joints), Cryogenics, B-field, etc.
Flexibility to accommodate user-requested/supplied instrumentation
  Additional voltage taps and Cernoxes specific to user sample
  Surface heaters (with larger gauge wires and in-house power supplies)
  Strain gauges (w/ 3 or 5 wire cable runs), acoustic sensors, hall probes, fiber optics
  Cryogenic measurements (pressure, mass flow, etc.)
  Ensure sufficient spare low- (~0.1Hz) and high-rate channels (~100 Hz) in DAQ

## Slide 15: 10 channels Digital Integrator
Items | Value | Units
Channels | 10 | 
Input Range | ±10 | V
Drift | <5x10-9 | Vs/s
Resolution | 1.49x10-14 | Vs
Trigger modes | Pulse Encoder Auto | 
Max Sample rate | 20 | MHz
Buffer Size | 1M | Samples/ch
Channel-Channel Isolation | 1 | kV
Channel-Chassis Isolation | 1 | kV
Max Chassis Power | 60 | W
Output Gain Range | 1/64 - 16384 | V/V
Trigger Interface | Differential | 
Trigger input impedance | 120 | Ohm
Fermilab is developing a Digital Integrator needed for magnet characterization and measurements
It is based on the 24-bit ADC
  Individual channels have own ADCs - not multiplexing
  Cost efficient
The goal is to substitute the expensive Metrolab FDIs
Can be used as precision digitizer needed for HTS samples
Current Status: one 10-Channel chassis has been assembled and compared to commercial alternatives (FDI-2056)
A. Makulski and T. Cummings

## Slide 16: Nanovolt Multiplexer Meter (NMM)
It very effective and useful tool to measure splice resistance
Effective segment voltages
  Provide information for potential conductor damage and type of the damage
Recently build - 31 channels
Using Keithley 2182A Nanovolt Meter for measurement
Standalone Python UI - no license/drivers needed
Easy readout: single USB Interface
Splices bellow 1 nOhm is measured
Courtesy of T. Cummings and S. Stoynev (FNAL)

## Slide 17: Example: 15 T Dipole Measurements
Used after magnet degradation
A characteristic V-I curve is observed in the two segments quenching most with the limiting one showing faster growth after 2 kA.
Other segments are showing similar signs but at much lower level.
NMM is going to be used during the commissioning of the TFD and for better diagnostics of the test samples
Courtesy of S. Stoynev (FNAL)

## Slide 18: Signal Conditioner (SC) Module
Needed for testing HTS samples
  convert signals to desired levels and types, and to allow longer cable runs
Main Requirements:
  Number of input channels: 10
  One output for each signal conditioning channel, and one bucking output
  Input and Output Type: Differential
  Input voltage range: ± 10mV to ± 200V
  Input overvoltage protection: ± 300V
  Input voltage offset:
    50μV non calibrated
    <5μV after calibration
  Crosstalk channel to channel: < -130dB
  Output voltage range: ± 10V
  Attenuation and amplification of the signals in 0.01 - 1000 range
  Development is in progress

## Slide 19: EMMA Framework for Measurement Systems
EMMA is a modern component-based framework and an example of flexible system, where a family of applications is built based on a common design and implemented on a common framework.
Architecture based on a publish-subscribe software bus via socket-based TCP/IP.
Applications configurable from standardized local and remote components.
Measurements automated via Python scripts.
On-line analysis developed in MATLAB and integrated as DLLs.
Easy debugging capabilities (data flow monitoring, component debug info)
Configurable user interfaces
Software and Comp. system group

## Slide 20: Measurement System Family
The software product line method is used to develop a family of measurement systems. These systems are based on a common architecture and a common framework.
All EMMA components have properties that are externally settable (configuration file) to tailor their behavior. The components in the system are coordinated by a Python script, which allows for additional flexibility in designing measurement systems.
SPLICE SYSTEM
SSW ROTATING COIL FIELD MAP MISCALLENOUS
SSW
(QUADRUPOLE)
SSW (DIPOLE)
VIBRATING WIRE
FLAT COIL (PPU)
SUPERCONDUCTING HORIZONTAL (HL-LHC)
CONVENTIONAL HORIZONTAL
SUPERCONDUCTING VERTICAL
Mu2e FMS
POINT SCAN (VERTICAL)
POINT SCAN (HORIZONTAL)
3D Field Mapping (PPU)
Commissioned
Under development
Planned development

## Slide 21: Summary
Fermilab has developed or is in the process of developing a full set of tools needed to protect TFD, as well as to protect and characterize the HTS cable samples.
  QD/QP system for TFD is commissioned
  QD/QP/QC for test samples
  The system is tested in real operation as we speak
We always increasing our “toolbox”:
  Nanovolt meter module – done
  Digital integrator module – in progress
  Signal conditioning module - in progress
  Development of the fiber systems
  Quench antennas and new MM probes
New software monitoring and measurement system is developed.

## Slide 22
Extras

## Slide 24: Anti-cryostat and sample holder
We have a conceptual design of the anti-cryostat
Expected operational mode: field is perpendicular to the test sample
Clear aperture: (EDIPO) 144 mm (H) x 94 mm (V)
B=15 T
F=1.5MN/m
Imax=100 kA
EDIPO/Sultan configuration
HTS CICC cable (Viper)
Instrumentation tree for HTS cable

## Slide 25: EMMA – User Interface