## Slide 1
Instrumentation and Diagnostics for Superconducting Magnets Workshop
Apr 24 – 28, 2023, Paestum
Maria Baldini, S. Krave
Fiber optic sensing for strain and temperature

## Slide 2: Fiber optics sensor
FBG sensor: discrete sensor
Rayleigh sensor: distributed sensor

## Slide 3: Outline
Fiber optics implementation at FNAL
FBG fibers:
    Calibration test
    Feedthrough line for the Vertical magnet test facility
Distributed fiber optics for strain and temperature measurements
    Calibration test
    Strain maps (Steve K.)
    Quench detection
Future plans

## Slide 4: Tensile stress test at 300 K (2020)
AISI 1080 steel and Ti
Instron press in IB3 up to 200 MPa
Fibers vs strain gauges
Tensile vs Poisson strain
Young modulus measured with FBG sensors is in very good agreement with what is expected for steel and Ti.
Purchase FBG fiber for HBM
8ch interrogator
Polyimide or Ormocer fibers with 4 sensors
Sensor length 6 mm

## Slide 5: Tensile stress test in Liquid Nitrogen
Tensile stress was applied up to 200 MPa in Liquid Nitrogen
Strain variation is around 800 me for both strain gauge and fibers
Poisson ratio is 0.3
Issues during cooldown
Signal was lost on some sensors
Absolute value after cooldown was not correct
Strain gauges
FBG sensors

## Slide 6: Fiber feedthrough in vertical magnet test facility
4x4 Fibers con FC/APC
Fiber Cables to Top Plate
G-10 Lambda Plug
Fisher connectors
Connector potted with Stycast 2850

## Slide 7: Mirror magnet cooldown (March 2022)
FBG2= 1508.013 nm
FBG3= 1514.049 nm
FBG4=1520.063 nm
@4.5 K only two sensors are still alive
Strain is consistent for each FBG sensor
At cold the strain variation without T compensation is around 4000 um/m
FBG2 Azimuthal
FBG4 Azimuthal
FBG3 Longitudinal

## Slide 8: Mirror magnet powering
Training and quench 1
At cold only two sensors are still alive
FBG3 is longitudinal
FBG4 is azimuthal
The closest strain gauge is labeled ULE60
Training and quench 2
Training and quench 3

## Slide 9: New feedthrough line for the Vertical magnet test facility
T < 70 K
TOP PLATE
LAMBDA PLATE
Implemented solution: use of pigtail cables spliced below the lambda plate
Very low signal < 70 K
Implemented solution: use of pigtail cables spliced below the lambda plate
Both FBG and Rayleigh
No signal loss at the splice
Line will be tested in the next couple of months

## Slide 10: Outline
Fiber optics implementation at FNAL
FBG fibers:
    Calibration test
    Feedthrough line for the Vertical magnet test facility
Distributed fiber optics for strain and temperature measurements
    Calibration test
    Strain maps (Steve K.)
    Quench detection
Future plans

## Slide 11: Distributed fibers: calibration on steel bar
3 segments of fiber were glued on each side of the steel bar
Observe the spectrum of entire length of the fiber
Pinch the fiber and identify specific position along the sensor
Laboratory Directed R&D
Purchase the interrogator
  strain sensors
  fibers encapsulated in Teflon tube: temperature sensors

## Slide 12
Distributed fiber sensor: quench detection
Fabricate several small solenoids with NiCr wire and wind a standard strain and temperature fibers sensors

## Slide 13
TEST # | Voltage | Capacitance | t | Energy stored
 | V | mF | ms | J
1 | 10 | 27 | 13.5 | 1.3
2 | 35 | 27 | 13.5 | 16.5
3 | 40 | 27 | 13.5 | 21.6
4 | 28 | 54 | 27 | 21.1
5 | 40 | 54 | 27 | 43.2
6 | 28 | 108 | 54 | 42.3
7 | 20 | 162 | 81 | 32.4
8 | 25 | 162 | 81 | 50.6
9 | 30 | 162 | 81 | 72.9
10 | 21 | 324 | 162 | 71.4
TEST PARAMETERS
5 m long temperature sensors
Sample rate 160 Hz and spatial pitch 2.6 mm.
Spot heater size (covering 3 solenoid turns): 5 mm wide, 40 mm long and 2.5 um thick
Capacitor bank (up to 12 capacitors): 27 to 324 mF
Test in Liquid N
Distributed fiber sensor: quench detection

## Slide 14: Distributed fiber sensor: quench detection
38 mm
25 cm
40 mm

## Slide 15: Distributed fiber sensor: quench detection
40 V 54 mF

## Slide 16: Distributed fiber sensors: quench detection
Temperature variation vs time constant
200 K variation in 100ms which is consistent with Nb3Sn coils during quench

## Slide 17: Distributed fiber sensors: quench detection
Outside the coil

## Slide 18: Temperature sensor in Li He
Test sensitivity of
Strain and temperature sensors in Liquid He
Nb3Sn witness sample
Ic measurements
Repeat experiment with a spot heater

## Slide 19: Future plan: test REBCO cable in Liquid N and He
Star REBCO cable
Corc REBCO cable
Wrap a temperature and a strain distributed optical sensor around it
Perform a quench propagation study in Li Nitrogen
Repeat the test in Liquid He

## Slide 20: Conclusions
Fiber feedthrough line have been implemented in the vertical magnet test facility at FNAL
Stable diagnostic probe for testing of R&D Magnets at FNAL
DISTRIBUTED SENSORS
QUENCH DETECTION:
  temperature fiber sensors were successfully used to measure T variations in quench like conditions in small solenoid
  Data collected with temperature fiber sensors are easier to interpreter than strain sensors
FUTURE PLAN
  Investigate T sensitivity at 1.9 K
  Quench propagation study on a REBCO cable

## Slide 21: Distributed fiber optics