## Slide 1
VHEE-FLASH radiotherapy facility
Beam Delivery - WP2
FRIDA meeting
Pisa 27-28 Settembre 2022
Lucia Giuliano

## Slide 2
VHEE-FLASH radiotherapy facility
Dependence of FLASH effect on treatment parameters
7 MeV S-band linac for FLASH therapy: experience in design and realization
Compact C-Band System at 5.712 GHz
C-band Very High Energy Electrons (VHEEs) LINAC for FLASH regime
  Injector
  TW structure
Goal of the talk:
Where are we ?
Where are we going?

## Slide 3
The mean dose-rate was initially used as a simplistic definition for FLASH irradiation (>40 Gy/s).
However, the instantaneous dose rate and the total treatment time appear to be the most critical beam parameters.
Friedl AA, Prise KM, Butterworth KT, Montay-Gruel P, Favaudon V. Radiobiology of the FLASH effect. Med Phys. 2022 Mar;49(3):1993-2013. Epub 2021 Sep 20.
Total dose
Dose per pulse
Mean dose rate
Instantaneous dose rate
Pulse length
Pulse Repetition Frequency
Number of the pulses
FLASH effect depends on different inter-dependent parameters
Beam parameters to induce the FLASH effect with electrons

## Slide 4: Conditions to obtain or miss the FLASH effect
Montay-Gruel P, Acharya MM, Jorge PG, et al. Hypofractionated FLASH-RT as an effective treatment against glioblastoma that reduces neurocognitive side effects in mice. Clin Cancer Res. 2021;27(3):775-784.
Available data on FLASH radiotherapy from in vivo and in vitro experiments
Highest dose rate in the pulse,
Shortest irradiation time,
X rays
Protons
highest is the FLASH effect.

## Slide 5: Experience with 7 MeV FLASH LINAC
Cathode
Chart: 5MeV pulse width 1ms





Es: 2,75V/25W = 110 mA

## Slide 6
E field accelerates the particles inside the LINAC
The resonant modes have a fixed spatial configuration
They oscillate in time
The electromagnetic energy is reflected back and forth at both ends of the cavity
S-band linac (f=2,998 GHz)
Bi-periodic geometry
Standing wave structure (SW)
p/2 mode: Coupling cells are empty (no field inside)
Magnetic coupling with hole off axis
Input power
Cathode
Standing wave structure
E field
Beam axis
Accelerating cell
Coupling cell

## Slide 7
S-band linac (f=2,998 GHz)
Standing wave structure (SW)
Bi-periodic geometry
p/2 mode: Coupling cells are empty (no field inside)
Magnetic coupling with hole off axis
Input power
Cathode
p/2 mode
E field

## Slide 8
Magnetic coupling with off axis holes
S-band linac (f=2,998 GHz)
Standing wave structure (SW)
Bi-periodic geometry
p/2 mode: Coupling cells are empty (no field inside)
Magnetic coupling with hole off axis
Off axis magnetic coupling
H field

## Slide 9
Nose-cone type geometry
Nose-cone type geometry very high electric field in the middle of the accelerating cells very efficient acceleration of the beam
E field

## Slide 10: Possible beam Parameters of the new VHEE-Linac
 | Description | Proposed Value for New Linac #
E | Beam Energy | 60 - 130 MeV
f | RF frequency | 5.712 GHz
PRF | Pulse repetition frequency | > 100 Hz
 | Pulse width | 1 - 3 µs
 | Pulse Charge | 200 - 600 nC
 | Pulse Current | 200 mA
 | Dose in a single pulse | 200 Gy* in ∅ 5 cm
 | In-Pulse Dose-Rate | >> Gy/s
 | Description | Proposed Value for New Linac #
E | Beam Energy | 60 - 130 MeV
f | RF frequency | 5.712 GHz
PRF | Pulse repetition frequency | > 100 Hz
 | Pulse width | 1 - 3 µs
 | Pulse Charge | 200 - 600 nC
 | Pulse Current | 200 mA
 | Dose in a single pulse | 200 Gy* in ∅ 5 cm
 | In-Pulse Dose-Rate | 
 | Description | Measured Value
E | Beam Energy | 7 MeV
f | RF frequency | 2.998 GHz
PRF | Pulse repetition frequency | > 100 Hz
 | Pulse width | 1 - 4 µs
 | Pulse Charge | 500 nC
 | Pulse Current | 125 mA
 | Dose in a single pulse | 19 Gy in ∅ 3 cm applicator*
 | In-Pulse Dose-Rate | > Gy/s
 | Description | Measured Value
E | Beam Energy | 7 MeV
f | RF frequency | 2.998 GHz
PRF | Pulse repetition frequency | > 100 Hz
 | Pulse width | 1 - 4 µs
 | Pulse Charge | 500 nC
 | Pulse Current | 125 mA
 | Dose in a single pulse | 19 Gy in ∅ 3 cm applicator*
 | In-Pulse Dose-Rate | 
* homogeneous (95%) field size at 55 cm of the exit window
Scaling to high energy the beam parameters for FLASH irradiation at 7MeV energy range
*Courtesy of WP4, “Beam characterization of the ultra-high dose rate research linac ElectronFLASH for preclinical study”
accepted in Physica Medica
Possibility of beam intensity modulation
Explore the FLASH effect both in the large field and pencil beam case

## Slide 11: Possible beam Parameters of the new VHEE-Linac
 | Description | Proposed Value for New Linac #
E | Beam Energy | 60 - 130 MeV
f | RF frequency | 5.712 GHz
PRF | Pulse repetition frequency | > 100 Hz
 | Pulse width | 1 - 3 µs
 | Pulse Charge | 200 - 600 nC
 | Pulse Current | 200 mA
 | Dose in a single pulse | 200 Gy* in ∅ 5 cm
 | In-Pulse Dose-Rate | >> Gy/s
 | Description | Proposed Value for New Linac #
E | Beam Energy | 60 - 130 MeV
f | RF frequency | 5.712 GHz
PRF | Pulse repetition frequency | > 100 Hz
 | Pulse width | 1 - 3 µs
 | Pulse Charge | 200 - 600 nC
 | Pulse Current | 200 mA
 | Dose in a single pulse | 200 Gy* in ∅ 5 cm
 | In-Pulse Dose-Rate | 
*Courtesy of WP4, “Beam characterization of the ultra-high dose rate research linac ElectronFLASH for preclinical study « submit to physica medica
We design accelerator by knowing the CHARGES
We use accelerator to deliver DOSE
FRIDA

## Slide 12
Compact electron linac for FLASH-RT
Reduce the size
Starting from the S-band (2.998 GHz) accelerator’s structure I scale down the geometric dimensions to get C-band design (5,712 GHz).
Increase Energy gradient
Increase the energy gain per meter (energy gradient)
Example
3GHz  6GHz
Ds-band = 7cm  DC-band =3,5 cm
GradientS-band 15 MV/m  GradientC-band 30MV/m
D=72.45 mm
D=36,54 mm
S-band C-band

## Slide 13
VHEE Linac Layout

## Slide 14: VHEE Linac modular Layout
The modular section is resulted the best choise because:
In a constant impedance structure the gradient it decays exponentially with z (lorger is worst).
The brazing is easier and allows the use of smaller ovens
The vacuum system is more efficient
VS

## Slide 15: VHEE Linac Layout
Injector

## Slide 16: Choice of main Accelerator structures: Source and Injector
Electron source: pulsed DC thermionic electron gun at 15 kV.
Bunching section is designed (L= lb/2 ) to match the low energy charges continuously emitted by the gun with the linac
the particles lost at the entrance of the linac are dramatically reduced beam capture (> 40%) .
Injector Linac with from e-gun up to ≈ 10 MeV and 200 mA beam peak current.
C-band linac (f=5,712 GHz)
Standing wave structure (SW)
p/2 mode
Bi-periodic geometry

## Slide 17: VHEE Linac Layout
TW section

## Slide 18: Choice of main Accelerator structures: TW linac
The frequency is 5.712 GHz;
TW Constant Impedance: Iris aperture radius a = 5 mm;
d= 2lb/3 allows to synchronize the beam passage in each cell with the accelerating field;
Initial single cell RF parameters simulations
2D side sketch of one single cell
Parameter | Size [mm]
a | 5
b | 21,03
d | 17,45
t | 3
R0 | 6
r1/r2 | 1,25
Unloaded
gradient
loaded gradient
@I=200 mA
a = 5mm

## Slide 19: RF parameters studies
Structure length = 90 cm;
Average gradient for is equal to = 40 MV/m, with an RF input power of Prf = 22.5 MW;
Modified Poynting vector below threshold of 2 MW/mm2
Eliptical shape to increase the shunt impedance (LINAC efficiency).
Courtesy of D. De Arcangelis

## Slide 20: VHEE C-band Linac
Dual-feed input and output RF power couplers;
Reflection Coefficient S11 = -30 dB at 5.712 GHz : very low reflective power ;
S11 parameter
Power in
Power out
Electric field animation
Power in
Power out
L. Giuliano, D. De Arcangelis, D . Francescone
dB
5.57 5.65 5.711 5.76 5.8
Frequency [GHz]
Electric field on the axis
Efield
Arbitrary Unit
0
-10
-20
-30

## Slide 21
Electron beam energy (MeV)
Iinj = 44 % Igun
Iout = 90% of Iinj
Igun = 600 mA
Energy = 15 keV
Beam Dynamics for the VHEE-Linac
energy gain and beam trasport
Output e-beam >130 MeV and 200 mA output e-beam;
Input e-gun current = 600 mA;
No focusing solenoids needed;
Total beam capture 40%;
Lost e-current phase-space evaluated for required radiosafety protocols;
Very low energy spread 200 keV @ 130 MeV
200 keV
Energy spectrum
courtesy of Luigi Faillace

## Slide 22
Quadrupole duplet
Grad. 47 T/m
3 cm bore
3 mm
3 cm
X (cm)
Y (cm)
X (cm)
Y (cm)
X (cm)
Y (cm)
Beam Spot Evolution
Xrms, Yrms (mm)
1 m
Beam Dynamics – Variable Field size
130 MeV
courtesy of Luigi Faillace

## Slide 23
12 m
7 m
LINAC
Y dipole
Magnetic field 1.6 T
Bore 3 cm
35deg
1 m
1.5 m
6 m
1 m
Quadrupole duplet
Gradient 47 T/m
Bore 3 cm
Infrastructure Layout @ Sapienza University
TECHNICAL ROOM
courtesy of Luigi Faillace

## Slide 24
Cell mechanical structure
42.049 mm
tuner
Water holes in rectangular configuration
10 mm
bolts to join the cells
courtesy of Marco Magi

## Slide 25: Mechanical structure
water for cooling
tuner
power in
power in
courtesy of Marco Magi

## Slide 26: Nearest future work
Optimization of mechanical structure with Marco Magi, D. Alesini, L. Faillace
Production of the prototype -> brazing tests
RF measurement on the prototype

## Slide 27
VHEE Linac Layout
Pulse compressor possibility
Giuseppe Torrisi
Talk

## Slide 28
Conclusions
We designed and developed (SIT-Sapienza University cooperation) LOW energy S-band linac for FLASH preclinical irradiation;
We set design parameters for novel accelerator fo FLASH RT according with measurements performed at Institut Curie (Orsay - France)
We did the upgrade to C-band technology (5.712 GHz) to reduce the size, the weight and increase the energy gradient;
Design and study of C-band Very High Electron Energy linac is under study at Sapienza University, in collaboration with INFN
  RF and beam dynamics design are been performed
  Mechanical design is been realized
  A prototype is under construction
Thank you
L. Faillace, L. Ficcadenti, D. Francescone, M. Migliorati, A. Mostacci, L. Palumbo, V. Favaudon, S. Heinrich, A. Patriarca,

## Slide 29
Tuesday, October 25, 2022
12.00 pm
Aula Francesco Calasso -Law Building
(Edificio di Giurisprudenza)
piazzale Aldo Moro 5, Roma
LECTIO MAGISTRALIS
Vincent Favaudon
Institut Curie (Orsay- France)
"From conventional dose-rate to FLASH radiotherapy: unveiled aspects of the dynamic of cell and tissue response to ionizing radiation".