## Slide 1: 15th workshop on breakdown science and high gradient technology 16-20 October 2023, Frascati, ITALY
Dielectric Assist Accelerating (DAA) structures for compact linear accelerators of low energy particles in hadrontherapy treatments
P. Martinez-Reviriego1, A. Grudiev3, D. Esperante1,2, B. Gimeno1, C. Blanch1,
N. Fuster-Martínez1, D. Gonzalez-Iglesias1, P. Martín-Luna1, E. Martinez1,
A. Menendez1 and J. Fuster1
1 Instituto de Física Corpuscular (IFIC), CSIC-University of Valencia, Parque Científico, C/ Catedrático José Beltrán, 2 46980 Paterna (Valencia)
2 Electronics Engineering Department, University of València, 46100 Burjassot, Spain
3 CERN, 01631 Meyrin, Switzerland

## Slide 2: Outline
Linear accelerators for hadrontherapy treatments
Dielectric Assist Accelerating (DAA) structure design procedure
Comparison for different materials and particle velocity
Multipactor analysis
Electromagnetic performance

## Slide 3: Hadrontherapy
Accelerator | Beam always present during treatment? | Energy variation by electronic means? | Time needed for varying the energy
Cyclotron | Yes | No | 80-100 ms (*)
Synchrotron | No | Yes | 1-2 s
Linac | Yes | Yes | 1-2 ms
Relative Biological Effectiveness (RBE)
Linear Energy Transfer (LET)
Marx, V. (2014, April 4). Sharp shooters. 508. Nature, p. 137
Compact and efficient accelerators
Physics challenges for linacs
Notes: LET: density of energy deposition along the radiation track.
RBE: measures the effects of the irradiation. Protons = 1.1
Hacer hincapié en órganos móviles

## Slide 4: Linear Accelerators for Hadrontherapy
~ 28 m
proton pulses 2.5us @ 200 Hz
18-20 MV/m
LINAC Conference 2014,
S. Benedetti et al.
RF DESIGN OF A NOVEL BACKWARD TRAVELLING WAVE LINAC FOR PROTON THERAPY
Backward Travelling Wave (BTW) High Gradient cavity testing at IFIC
ADAM, spin-off of CERN and TERA foundation is developing a proton linear accelerator to be installed in a hospital in England
Normal Cavities
High-Gradient Cavities
50 MV/m

## Slide 5: Dielectric Assist Accelerating (DAA) cavity
Copper
Dielectric:
High
Low
D. Satoh, M. Yoshida, and N. Hayashizak, "Dielectric assist accelerating structure." Physical Review Accelerators and Beams, vol. 19, 1, pp. 1011302, 2016
Investigations Into X-Band Dielectric Assist Accelerating Structures for Future Linear Accelerators. Yelong Wei, Alexej Grudiev.
Parameter | Calculation
 | mm
Parameter | Calculation
Resonant frequency for the mode depends on the combination of
Working under mode:
  High .
  Dielectric helps to decrease cavity size.
  Low electric field in metal.
  Axial symmetry
Low particles
S-band
Standing wave
Notes: .

## Slide 6: DAA cavity single cell solution
Electric field
Magnetic field
Electric field focus on accelerating area.
  Low field on metallic surface:
  High breakdown limit
Magnetic field concentrated on vacuum.
  Low losses on metal:
    High
Notes: .

## Slide 7: DAA cavity single cell solution
Electric field
Electric Energy
1st: Parallel boundary
2nd : is conserved along the dielectric
3rd : Perpendicular boundary
is constant
High inside dielectric
is constant
High in vacuum
Notes: .

## Slide 8: DAA cavity single cell design
Resonant frequency for the mode depends on the combination of :
  Scan for and we look for the value of that makes MHz.
  Look for the values of that maximizes
Example for ideal material: and
Notes: .

## Slide 9: DAA cavity single cell iris optimization
Scan in iris thickness:
Iris thickness
High Power Test Results of X-Band Dielectric Disk Accelerating Structures. Ben Freemire
Notes: .

## Slide 10: Energy range for hadrontherapy
Protons: MeV
: MeV/u
Bencini, V. (2020). Design of a novel linear accelerator for carbon ion therapy (Doctoral dissertation, Rome U.).
Real material
Ideal material
Notes: .

## Slide 11: Scan in losses
Conductor:
Dielectric:
Notes: .

## Slide 12: SEY Measurements
DLC Coating
Ceramic
Metals
Dielectric
Notes: a-C -> Amorphous Carbon
DLC -> Diamond-Like Carbon -> above 1 MOhm per square ~500 nm

## Slide 13: Multipactor in the DAA structure
  Some of the main results of the multipactor simulations in the DAA cell are summarized on this slide
Multipactor growth factor σ as a function of the RF electric field amplitude at the cell axis in the down zone (left) and up zone (right)
Ne , number of electrons
N0 , initial electron number
growth factor
a) Number of electrons in the structure as a function of time normalised to the period of the RF signal.
Results for E0 = 1 MV/m (down zone)
Multipactor simulation in the down zone for E0 = 1 MV/m
b) Colour map with the number of electrons impacting at each wall position being able to generate two or more secondary electrons.
Courtesy of Daniel Gonzalez-Iglesias
Notes: a-C -> Amorphous Carbon
DLC -> Diamond-Like Carbon -> above 1 MOhm per square ~500 nm

## Slide 14: Coating losses
Amorphous and diamond-like carbon coatings for sey reduction of dielectric material for accelerating structure application. A. Grudiev et al. Tech. Rep. 2022.
Coating losses:
internal
external
a-C coating can reduce multipactor but it has an impact on electromagnetic performance
Notes: .

## Slide 15: Coupling and dispersion curve
High electrical coupling between cells.
  No need for coupling cells
Low leads to dispersion curves crossing.
Thicker irises lead to mode overlapping.
D16,
Notes: .

## Slide 16: Thermal simulation
Boundary condition: 22 ºC
Thermal load:
Heat flux: Surface losses in copper.
Heat generation: Volumetric losses in dielectric.
Electromagnetic normalization:
DUT =
https://www.americanelements.com/magnesium-titanate-12032-30-3
https://www.makeitfrom.com/material-properties/Magnesium-Titanate-IEC-60672-Type-C-320
https://accuratus.com/alumox.html
Notes: .

## Slide 17: Critical points
Triple junction points can lead to instabilities and singularities in electric field in the neighborhood of the point
In the surroundings of the junction, so we can use a 2D constant field approach.
Notes: .

## Slide 18: Analytical approach
Techaumnat, B., Hamada, S., & Takuma, T. (2002). Effect of conductivity in triple-junction problems. Journal of electrostatics, 56(1), 67-76
The contact-point electric field will be infinitely large if and will be zero if . A non-zero and non-singular value will exist only if .
Notes: .

## Slide 19: Conclusions
  DAA have extremely high and :
    Results are better for high and
    Lower input power needed.
  worse than copper structures
    Higher total energy needed.
  Iris length is an important parameter to add in the optimization process.
  Multipactor must be suppressed by coatings with high sheet resistivity.
  Very high coupling between cells
    Modes overlapping can be a problem for low and .
  Very low thermal conductivity can be a source of mechanical problems.
  Triple junction point and sharp corners must be carefully studied to avoid singularities and instabilities.
Notes: .

## Slide 20: Back up

## Slide 21: Standing Wave Acceleration Cavities
P RF in
z
Ez
z
 mode
0 mode
Ez
Synchronization condition
: particle velocity
d: distance between cells
fRF: RF frequency
c: speed of light in vacuum
Cylindrical single (or multiple cavities) working on the TM010-like mode are used
Figures of merit:
Shunt impedance: efficiency of the acceleration mode.
Quality factor: efficiency to store RF energy .
R/Q: pure geometric qualification factor.
  NC cavity R1M
  SC cavity R1T
  NC cavity Q104
  SC cavity Q1010
100
 mode

## Slide 22: HG-CCL Copper structure single cell
, mode
Electric field
Magnetic field
Electric field
Magnetic field
Pillbox cavity
Optimization
High losses in metallic walls: low RF efficiency.
High peak electric field in metal: field emission and RF breakdown.
Notes: .

## Slide 23: DAA cavity single cell solution
Electric field
Magnetic field
Electric field
Magnetic field
Pillbox cavity
Dielectric and Optimization
DAA mode
Low losses in metallic walls
Low peak electric field in metal.
Notes: .

## Slide 24: DAA cavity single cell CST
Resonant frequency for the mode depends on the combination of :
  Scan for and we look for the value of that makes MHz.
  Look for the values of that maximize
In CST the minimum volume to simulate is 1/8 of the total volume using symmetry planes in XY, XZ, YZ.
Then the mode we are interested in is mixed with many modes with no revolution symmetry
Notes: .

## Slide 25: DAA cavity single cell Superfish
Resonant frequency for the mode depends on the combination of :
  Scan for and we look for the value of that makes MHz.
  Look for the values of that maximizes
Superfish: 2D electromagnetic solver for axial symmetric RF cavities.
  Fast.
  Low computational cost.
  Free.
  Reduce the number of solutions due to the symmetry.
Magnetic wall
Electric wall
cm
cm
Notes: .

## Slide 26: DAA cavity single cell Superfish
Notes: .

## Slide 27: HFSS Results
Geometry | SUPERFISH |  |  | HFSS |  | 
 | (MHz) |  | () | (MHz) |  | ()
edge | 2998.7 | 46949 | 944 | 2998.2 | 47444 | 949
mm | 2999.3 | 47316 | 959 | 2999.1 | 47755 | 960
Geometry | SUPERFISH |  |  | HFSS |  | 
edge | 2998.7 | 46949 | 944 | 2998.2 | 47444 | 949
 | 2999.3 | 47316 | 959 | 2999.1 | 47755 | 960
D16,
Conductor:
Dielectric:
Coating:
Notes: https://www.sciencedirect.com/topics/engineering/dielectric-loss-tangent

## Slide 28: Rounded points
A
B
C
D
AB
CD
Notes: .

## Slide 29: SEY Measurements
DISK 2: Sample 5: MgTiO with one side DLC coated (US vendor).
D16-2: Sample 9: D16 with one side DLC coated (Japan vendor).
MgTiO: Sample 4: MgTi Oxide based Conductive Ceramic uncoated.
DISK 7: Sample 6: MgTiO with one side DLC coated (Japan vendor).
D16-1: Sample 8: D16 with one side DLC coated (US vendor).
Coating
Ceramic
Metals
Dielectric
Notes: a-C -> Amorphous Carbon
DLC -> Diamond-Like Carbon -> above 1 MOhm per square ~500 nm

## Slide 30: Multipactor in the DAA structure
  Multipactor breakdown is an electron avalanche-like discharge occurring in components operating under vacuum conditions and high-power RF electromagnetic fields
  The onset of multipactor discharge leads to various detrimental effects that degrade the device performance. Thus, this phenomenon poses a significant limitation on the maximum RF power handling capability of devices
  The risk of suffering a multipactor discharge in the DAA structures was analyzed by means of numerical simulations with an in-house developed code based on the Monte-Carlo method
Primary electron (free)
Secondary electron
Diagram of the DAA cell
vacuum
e-
e-
d
Waveguide wall
Waveguide wall
e-
e-
e-
e-
e-
e-
e-
e-
e-
e-
e-
e-
e-
e-
e-
e-
e-
e-
e-
e-
e-
e-
e-
e-
e-
e-
e-
e-
e-
e-
e-
e-
e-
e-
e-
e-
e-
e-
e-
e-
e-
e-
e-
e-
e-
e-
e-
e-
t
5 TRF/2
0 1 2 3 4 5 6 7 8 9 10
Wall collisions
  The simulations explore RF electric field amplitudes in the range E0 = [0.01, 200] MV/m
  The cell is divided into two zones: up and down, with separate simulations launched for each zone
  In the DAA cell multipactor appears in the range 1-200 MV/m
Courtesy of Daniel Gonzlez-Iglesias
Notes: a-C -> Amorphous Carbon
DLC -> Diamond-Like Carbon -> above 1 MOhm per square ~500 nm

## Slide 31: ANSYS Workbench
Ceramic
copper
Material | MgTiO3 | 99.5% | TiO2-doped Al2O3 | CVD Diamond | Copper
Density (g/cm^3) | 3,6 | 3.89 |  | 3.52 | 8,933
Thermal conductivity (W/m·K) | 3,8 | 35 |  | 2000 | 400
Thermal expansion (µm/m·K) | 8,0 | 8.4 |  | 1.0 | 17,7
 | 16.66 | 9.8 | 9.64 | 5.7 | 
Material | MgTiO3 |  | TiO2-doped Al2O3 | CVD Diamond | Copper
Density (g/cm^3) | 3,6 | 3.89 |  | 3.52 | 8,933
Thermal conductivity (W/m·K) | 3,8 | 35 |  | 2000 | 400
Thermal expansion (µm/m·K) | 8,0 | 8.4 |  | 1.0 | 17,7
 | 16.66 | 9.8 | 9.64 | 5.7 | 
https://www.americanelements.com/magnesium-titanate-12032-30-3
https://www.makeitfrom.com/material-properties/Magnesium-Titanate-IEC-60672-Type-C-320
https://accuratus.com/alumox.html
Notes: .

## Slide 32: Tolerance studies
Parameter | (kHz/m) |  |  |  |  | 
material | D9 |  | D16 |  | D50 | 
 | 0.4 | 1 | 0.4 | 1 | 0.4 | 1
 | 1.7 | 2.3 | 4.0 | 2.1 | 2.7 | 1.1
 | -11.0 | -42.8 | -28.0 | -51.8 | -53.1 | -63.9
 | -2.0 | -5.5 | -2.2 | -3.6 | -1.3 | -1.4
 | -29.5 | -12.1 | -39.1 | -12.8 | -34.5 | -12.6
 | 1.1 | 1.4 | 1.6 | 1.1 | 2.0 | 0.8
 | -71.3 | -4.8 | -38.8 | -3.1 | -10.3 | -1.2
Parameter |  |  |  |  |  | 
material | D9 |  | D16 |  | D50 | 
 | 0.4 | 1 | 0.4 | 1 | 0.4 | 1
 | 1.7 | 2.3 | 4.0 | 2.1 | 2.7 | 1.1
 | -11.0 | -42.8 | -28.0 | -51.8 | -53.1 | -63.9
 | -2.0 | -5.5 | -2.2 | -3.6 | -1.3 | -1.4
 | -29.5 | -12.1 | -39.1 | -12.8 | -34.5 | -12.6
 | 1.1 | 1.4 | 1.6 | 1.1 | 2.0 | 0.8
 | -71.3 | -4.8 | -38.8 | -3.1 | -10.3 | -1.2
Notes: .

## Slide 33: Dielectric Disk Loaded Accelerating (DDA) Cavity
Gap/2
Parameter | With nose | Without nose
 | 21180 | 20299
() | 2154 | 1234
() | 45.6 | 25.0
T | 0.9076 | 0.8055
 | 6.27 | 2.38
Parameter | With nose | Without nose
 | 21180 | 20299
 | 2154 | 1234
 | 45.6 | 25.0
T | 0.9076 | 0.8055
 | 6.27 | 2.38
Same
Notes: .