## Slide 1
High Precision X-Ray Measurements 2023
19 – 23 June 2023, Laboratori Nazionali di Frascati, Italy
A. Buttacavoli1*, F. Principato1, G. Gerardi1, D. Cascio1, G. Raso1, V. Taormina1, M. Bettelli2, A. Zappettini2, S. Zanettini3, G. Bertuccio4, F. Mele4, J. Quercia4 and L. Abbene1
1Department of Physics and Chemistry-E. Segrè, University of Palermo, Italy
2IMEM/CNR, Parma, Italy
3due2lab s.r.l, Reggio Emilia, Italy
4Politecnico di Milano, Como, Italy
Development of energy-resolved X-ray scanners for contaminant detection: results from the AVATAR X project
*corresponding author: antonino.buttacavoli@unipa.it

## Slide 2: Motivations
Goal: Development of room temperature spectroscopic imagers for X-ray imaging up to 150 keV.
AVATAR X Project (MUR)
Energy-resolved X-ray scanners
for food inspections
Department of Physics and Chemistry (DiFC),
University of Palermo (Italy) (L. Abbene)
IMEM-CNR, Parma (Italy) (A. Zappettini)
Politecnico di Milano, (G. Bertuccio)
Due2lab company, Italy (N. Zambelli)
Collaboration

## Slide 3
Proposed Detection Systems
CZT linear array detectors
Front-end electronics with
the preamplifier (CSP) stage
Digital Readout Electronics for Pulse Shape and
Height Analysis
(PSHA)
Room temperature operation
Energy resolution < 3 keV at FWHM at 60 keV
Sub-millimetre spatial resolution (< 0.5 mm)
Requirements

## Slide 4: DETECTORS AND ELECTRONICS

## Slide 5
The Detectors
Detectors are fabricated at IMEM-CNR (Parma, Italy) and by due2lab s.r.l (Reggio Emilia, Italy).
CZT crystals (10.4 x 3 x 1 mm3) grown by using the Boron Oxide Encapsulated Vertical Bridgman (B-VB) growth technique and gold electroless contacts.
Linear arrays with pixel pitch of 225 µm.
Inter-pixel gap of 25 µm.
Intermediate guard-ring to minimize charge sharing.
A. Zappettini et al., J. Cryst. Growth 307 (2007) 283

## Slide 6
B-VB CZT Detectors
Physical properties of B-VB CZT detectors
Low leakage currents
High bias voltage operation (> 7000 V/cm)
Good electron charge transport properties
L. Abbene, et al. J. Synchrotron Rad. (2020) 27 (2), 319.

## Slide 7
Front-end Electronics
Hybrid charge sensitive preamplifiers developed at University of Palermo
Front-end electronics with preamplifier stage (no analog pulse shaping).
ENC < 100 electrons.
Resistive-feedback circuit with a decay time of 20 us.

## Slide 8
Digital Electronics
Digital Pulse Processing (DPP) electronics.
8 digitizers (100 MHz) with 32 channels
(modified version of CAEN DT5724, CAEN S.p.A., Italy).
Custom Firmware based on a DPP strategy developed at DiFC (University of Palermo).
L. Abbene, et al. J. Synchrotron Rad. (2018) 25 (1), 257.

## Slide 9
Outputs
the arrival time
the pulse height (energy)
the pulse time width and the peaking time (shape)
Digital Real-Time Pulse Shaping by using the Single Delay Line (SDL) and Trapezoidal filtering.
The system also allows the acquisition of the CSP waveforms for quick look or off-line analysis.
Digital Pulse Processing

## Slide 10: SPECTROSCOPIC RESULTS

## Slide 11
Energy Spectra from Anode Pixels
- 700 V

## Slide 12
Energy Spectra from internal guard-ring

## Slide 13
Energy Spectra from Cathode

## Slide 14
Spectroscopic Correction with Cathode/Anode Pulses

## Slide 15: PRELIMINARY RESULTS WITH A NEW ASIC

## Slide 16
Spectroscopic Performance with a new ASIC
SIRIO ASIC
Ultra-high resolution at room temperature… very close to the Fano Noise (0.4 keV at 60 keV)!
Preamplifier Stage
Technology: CMOS 0.35 µm
Conversion gain: 40 mV/fC
ENC (without detector) 3.2 el. r.m.s.
Rise-time (without detector): <10 ns
F. Mele, et al. IEEE TNS (2021) 68, 379-383.

## Slide 17: CHARGE SHARING MEASUREMENTS

## Slide 18
Charge Sharing Results
Shared events of 49 % with inter-pixel gap of 25 µm.
For comparison, shared events of 60 % with inter-pixel of 50 µm.
Time coincidence events between a pixel with all adjacent pixels and guard-ring (simulating a pixel layout)

## Slide 19
Time coincidence events between a pixel with the two adjacent pixels (linear array approach)
Shared events of 36 % with collimator and guard-ring (linear array approach).
Charge Sharing Results

## Slide 20
Time coincidence events between a pixel with the two adjacent pixels (linear array approach)
Charge Sharing Results

## Slide 21: RESULTS FROM ENERGY-RESOLVED IMAGES

## Slide 22
The Phantom
X-ray scanning with Ag-target X-ray tube (50 kV)
Scanned area

## Slide 23
Photon Counting Images and Charge Sharing
X-ray scanning with Ag-target X-ray tube (50 kV)
Raw Image (with charge sharing)
CNRsteel = 6 CNRplastics = 1
G-CSD Image (after CSD with the internal guard-ring)
CNRsteel = 11 CNRplastics = 5
CSD Image (after CSD)
CNRsteel = 11 CNRplastics = 6

## Slide 24
Energy-Resolved X-ray Images
New Technique: Window-Based Energy Selecting X-ray Imaging
A. Buttacavoli et al., Sensors Vol 23, Issue 6 (2023) 3196
Taormina et al.
21 June 2023 5:10 pm
Session: X-ray detectors 3
This conference

## Slide 25
Energy-Resolved X-ray Images
New Technique: Window-Based Energy Selecting X-ray Imaging
A. Buttacavoli et al., Sensors Vol 23, Issue 6 (2023) 3196
Taormina et al.
21 June 2023 5:10 pm
Session: X-ray detectors 3
This conference

## Slide 26
Conclusions and Ongoing Activities
B-VB CZT linear array detectors with good room temperature performance (FWHM < 3 keV at 60 keV) are developed.
Inter-pixel gap of 25 µm allows the reduction of charge sharing and charge losses.
A new energy-resolved approach, termed window-based energy selecting, was developed and excellent CNR enhancements in X-ray images were obtained, if compared with photon counting and other energy-resolved techniques;
20
Ultra-high energy resolution (0.6 keV FWHM at 60 keV) was obtained with a new ASIC, opening exciting activities for spectroscopic and imaging enhancements.

## Slide 27
THANK YOU FOR YOUR KIND ATTENTION