## Slide 1: Innovative X-ray detection systems based on monolithic arrays of Silicon Drift Detectors
Griseld Deda1,2, Giacomo Ticchi1,2, Beatrice Pedretti1,2, Davide Di Vita1,2, Giacomo Borghi1,2, Marco Carminati1,2, Nicola Zorzi3 and Carlo Fiorini1,2
Dipartimento di Elettronica, Informazione e Bioingegneria, Politecnico di Milano, Milan, Italy
INFN, Sezione di Milano, Milan, Italy
Fondazione Bruno Kessler, Trento, Italy

## Slide 2: Silicon Drift Detector and CUBE Preamplifier
Silicon Drift Detectors provides the best energy resolution (small anode capacitance, independent of detector area)
CUBE Charge Sensitive Amplifier (CSA) with low series noise (high ) is bonded close to SDD anode (minimize parasitics)
SDD+CUBE = versatile solution for high count-rate application (Bombelli et.al.,2012)
Example of SDD array mounted on a ceramic
SDD array and CUBE
Radiation Entrance Window
Charge Drift
Charge Generation
SDD
CUBE
Goal: adopt SDDs in different configurations to enhance properties like detection efficiency, throughput capability, and compactness, in addition to high-energy resolution.

## Slide 3: Motivation
Nuclear physics experiments
Industrial bench-top and portable X-ray instruments
Example: high speed mapping in XFM (X-ray Fluorescence Microscopy)
High-speed X-ray mapping for cultural heritage applications
High speed mapping in X-ray Fluorescence Microscopy
Measurement taken at ID16 (ESRF) with ARDESIA detector (courtesy P.Cloetens)
SIDDHARTA-2 experiment

## Slide 4: Stacked-SDDs Detection Module
Detection area
X-Ray source
SDDs
Goal: Development of a new detection module based on 1 mm thick monolithic arrays of Silicon Drift Detector (SDD) coupled with CMOS-based Charge Sensitive Amplifiers (CSA) arranged in a stacked structure to reach 2 mm of effective thickness.
Quantum Efficiency
Energy (keV)

## Slide 5: Stacked Structure: SDD and PCB Carrier
20 mm
36 mm
8 mm
8 mm
3 mm
TOP side
BOTTOM side
1
2
3
4
5
6
7
8
CSA
41 mm
Board cutout
26 mm
16 mm
32 mm
Connector
Top SDD
Aluminum
spacer
Bottom PCB
2x4 monolithic array of Silicon Drift Detector
1 mm thick
64 mm2 pixel area
36 mm x 20 mm chip dimensions
Produced by Fondazione Bruno Kessler (FBK)

## Slide 6: X-ray Spectroscopy Measurements
10
20
30
40
50
60
70
kCounts
0
0
1
2
3
4
17,8keV
13,9keV
26,4keV
20,8keV
59,5keV
Top SDD
Bottom SDD
10
20
30
40
50
60
70
Energy [keV]
kCounts
0
0
1
2
3
4
26,4keV
59,5keV
17,8keV
13,9keV
20,8keV
Energy [keV]
Am241 source -> characteristic lines in both soft and hard x-ray energy range
Read-out with SFERA ASIC, a 16-channel APP for X and -ray detectors
A first demonstrator adopting 450 m thick SDDs, similar to the ones to be used in the final module has been realized.

## Slide 7: X-ray Spectroscopy Measurements
10
20
30
40
50
60
70
kCounts
0
0
1
2
3
4
17,8keV
13,9keV
26,4keV
20,8keV
59,5keV
Top SDD
Bottom SDD
10
20
30
40
50
60
70
Energy [keV]
kCounts
0
0
1
2
3
4
26,4keV
59,5keV
13,9keV
20,8keV
17,8keV
10
20
30
40
50
60
70
Energy [keV]
0
0
1
2
3
4
26.4keV
59.5keV
17.8keV
kCounts
13.9keV
20.8keV
Ch3 Top
Ch3 Top + Ch3 Bottom
Total spectrum obtained from the sum
Two-fold increase of the total count rate of the system at the 59.5 keV characteristic line of Am241.

## Slide 8: SCARLET Project
Future SCARLET 8x8 Illustration
Excitation beam (photons or electrons)
X-rays
Bump bonding
Collimator
SDD monolithic array 8x8
Substrate with high thermal conductivity
Two separated connections, for DAQ and power supply, respectively
Mechanical support and thermal sink
Sample
4x4 readout ASIC
for each quadrant
ASIC features:
4 readout channels
1Mcps/pixel output count rate
full analog processing in the ASIC pixel
Res. <200eV @200ns pulse width
ADC integrated in the ASIC
Goal of the project: Development of a readout ASIC to be coupled to pixel detectors (focus: monolithic SDDs arrays) for ultra-high rate X-ray spectroscopy.
Gold-stud bump-bonding[*]
Chip
Heat and pressure
Substrate
Au stud bump
(after leveling)
[*] Caselle et al., J.Inst. 2016

## Slide 9: SCARLET Hybrid Detection Module
2 mm
12 mm
4x4 monolithic array
4 mm2 pixel area
12 mm x 12 mm chip dimensions
450 μm thick; available also 1 mm thick
Produced by Fondazione Bruno Kessler (FBK)
Entrance window
SCARLET
Flip-chip assembly
Bump-bonding: SDD bias; anode current signal readout
Wedge-bonding: ASIC power supply; ASIC programming

## Slide 10: Spectroscopic Experimental Results
233eV
161eV
148eV
168eV
220eV
155eV
149eV
174eV
229eV
185eV
169eV
191eV
twidth [ns]
400
1000
2000
200
FWHM [eV]
120
140
160
180
200
220
240
Mn-K
6,49keV
Mn-K
5,89keV
148eV FWHM
(10,3 e- rms)
1us pulse width
ADC bin
Counts
Iron-55 X-ray source
Spectra from internal ADC
T ~ 0°C
twidth

## Slide 11: ARDESIA-16 Spectrometer
Detection Module
57 ø
170
255
[All dimensions in mm]
140
Synchrotron Radiation Source
Focusing
Optic
Incident X-rays
Emitted X-ray
Fluorescence
Energy Dispersive
Detector
Sample
Monochromator
XRF
XAFS
XFM

## Slide 12: ARDESIA-16 Spectrometer
4 × 4 5-mm-side
24 mm
5 mm side for each pixel
Thicknesses from 450 µm to 1 mm
Produced by FBK
16 elements monolithic array
Detection Module
57 ø
170
255
[All dimensions in mm]
140
[*] Utica, G., et al. "ARDESIA-16: a 16-channel SDD-based spectrometer for energy dispersive X-ray fluorescence spectroscopy." Journal of Instrumentation 16.07 (2021): P07057.

## Slide 13: ESRF ID16A – DESY PETRA III P06
High-Definition nano-XRF (Courtesy: Peter Cloetens)
Step size: 10 nm
Dwell time: 15 ms
Fe-Kα
Cu-Kα
Atherosclerotic plaques on an arthritic bone
Dwell time: 10 ms
Courtesy: Peter Modregger (Univ. Siegen)

## Slide 14: Towards ARDESIA-64
57 ø
170
255
[All dimensions in mm]
140
X-ray beam
Higher-intensity X-ray beam
Sample
Sample
Fluorescence emission
Increased fluorescence emission
Advantages:
Lower experiment time
Higher rate
Increased mapping resolution
100 Mcps
17 Mcps
ARDESIA-16
ARDESIA-64

## Slide 15: ASCANIO Detection Module
PCB
Mo needle
Mo collimator
SDD
X-ray beam
Sample
Sample fluorescence light
Aluminum enclosure
Top Beryllium window
Bottom
Beryllium window
Copper block
Peltier cell
SDDs modules
14
4 channels monolithic array
1 mm thickness
5 mm x 5 mm pixel area
2 mm wide guard rings
Energy Range: 0.2 ÷ 20
Designed for X-ray fluorescence microscopy (XFM) imaging in synchrotron beamlines

## Slide 16: ASCANIO Output Count Rate
Assumption: largest solid angle channels operate at 2 Mcps*
[*] Scoullar P A, McLean C C and Evans R J 2011 Real time pulse pile-up recovery in a high throughput digital pulse processor AIP Conference Proceedings vol 1412 (American Institute of Physics) pp 270–277.

## Slide 17: 55Fe source Spectroscopy Measurements
FWHM = 129 eV
PT = 2 us
FWHM = 175 eV
PT = 32 ns
High energy resolution for high count rate measurements
Measures performed with Dante DPP (XGlab)

## Slide 18: Towards a 4π Detection System
57,5 mm
22,8 mm
8 mm
Goal: Development of a 4π detection system based on arrays of SDD detectors able to measure directionality of the signal (X-ray emitted by central target)

## Slide 19: Conclusions and Future Work
We presented the adoption of SDDs in different configurations to enhance properties like, detection efficiency, throughput capability and compactness, in addition to high-energy resolution.
Absorption efficiency
Stacked SDDs
Bump-bonding technology
SCARLET project
Solid angle
ASCANIO
4π
ARDESIA-64
Miniaturization of SDD element
Ongoing activities:

## Slide 20
Thank you for your attention!
griseld.deda@polimi.it