## Slide 1
Aristotle University of Thessaloniki
Comparison of different methods for evaluating quantitative XRF data in copper-based artefacts
Eleni Konstantakopoulou, Annalaura Casanova Municchia, Loredana Luvidi, Marco Ferretti
High Precision X-Ray Measurements 19-23 June 2023, Frascati, Italy

## Slide 2
X-Ray Fluorescence Analysis of Heritage Materials
Simultaneous multi elemental analysis
Non-destructive
Portable devices
Application of XRF in Cultural Heritage
Fabrication technology
Authenticity
Provenance of the artefacts

## Slide 3
Hand-Held XRF devices
Highly miniaturised hardware
Powerful software which provides quantitative data in real time
Built-in manufacturer calibration
Customized manufacturer calibration
Independent external software
Aim of this study is the comparison of the quantitative results of copper-based alloys from a commercial HH-XRF, using 3 different approaches

## Slide 4
Materials and Methods (1)
Bruker Tracer 5g
Instrument
Reference Materials
26 certified standards compositionally significant of heritage copper-based artefacts
Rh anode X-Ray tube
50 kV max. voltage
195 uΑ max. current
Silicon Drift Detector
Collimators: 3mm, 8mm
5 different filters
ISPC Bruker Tracer 5g
26 Copper - based coupons

## Slide 5
Materials and Methods (2)
Experiments
10 series of measurements on 26 standards for each calibration approach by 3mm
 | Built-in Bruker For Cu Alloys | Customized Bruker | Off-Line PyMca
Voltage | 15kV | 49kV | 49kV
Current | 23.2 uA | 39.3uA | 39.3uA
Filter | - | Al(300um) Ti(25um) | Al(300um) Ti(25um)
Meas.Time | 30s | 30s | 30s
Quantification Method | Empirical Coefficients | Empirical Coefficients | Fundamental Parameters

## Slide 6
Results – Tin
Calibration | R square | RMSE
Built-in | 0.997 | 0.251
Customized | 0.999 | 0.122
PyMca | 0.999 | 0.145

## Slide 7
Results – Antimony
Calibration | R square | RMSE
Built-in | 0.982 | 0.149
Customized | 0.999 | 0.025
PyMca | 0.999 | 0.024

## Slide 8
Results – Tin and Antimony
PyMca
Sb-K Lines
K Lines
L Lines
Sn
Sb

## Slide 9
Results - Iron (1)
Calibration | R square | RMSE
Built-in | 0.860 | 0.092
Customized | 0.982 | 0.034
PyMca | 0.996 | 0.015

## Slide 10
Customized Calibration
Results - Iron (2)
EasyCal™
Overlap correction

## Slide 11
PyMca
Results - Iron (3)
PyMca
PyMca Calibration
Peaks Spectrum
Fe-Kα
Fe-Kβ
Cu-EP
Co-Kβ
Ζn-EP
Cu-EP

## Slide 12
Results – Root Mean Squared Error

## Slide 13
Conclusions
Hand-Held XRF devices cannot be considered as black-boxes. The operator must have knowledge of XRF principles
Quantitative data obtained by Built-in Bruker Calibration show the largest deviation from the nominal values
Customized Bruker and PyMca Calibrations provide more accurate results and can be combined for the quantification of copper-based artefacts
Customized Bruker Calibration immediately provides quantitative results and is useful during the field acquisition, while PyMca Calibration can be used for subsequent more detailed spectrum processing

## Slide 14
Aristotle University of Thessaloniki
Thank you for your attention