## Slide 1
Use of hydrogenated amorphous silicon to precisely measure X-ray fluxes
Keida Kanxheri on behalf of HASPIDE collaboration
HIGH PRECISION X-RAY MEASUREMENTS 2023

## Slide 2
HASPIDE PROJECT
INFN CSN5 CALL
STARTED IN 2022
3 YEARS DURATION
8 INSTITUTIONS:
6 INFN ( FI, LE, LNS, MI, PG, TO)
1 UNIVERSITY OF WOLLONGONG (AUS)
1 EPFL NEUCHATEL (SWITZERLAND)
~ 45 RESEARCHERS
2

## Slide 3
HYDROGENATED AMORPHOUS SILICON
1969
1976
SYNTHESIZED FOR THE FIRST TIME BY CHITTIK ET AL. FROM GROWTH BY PLASMA-ENHANCED VAPOR DEPOSITION (PE-CVD) OF SiH4 (SILANE)
SPEAR AND LECOMBER DEMONSTRATED THAT THIS MATERIAL COULD BE SUBSTITUTIONALLY DOPED BOTH IN N AND P-TYPES. DEVELOPMENT OF TRANSISTORS, SOLAR CELLS AND MEMORIES.
FROM MID ‘80 ON
DIRECT DETECTION OF
IONIZING RADIATION BASED
ON P-I-N OR SCHOTTKY DIODE STRUCTURES (THICKNESS 1 TO 50 um)
3

## Slide 4
WHY a-Si:H AS MATERIAL?
IT IS INTRINSICALLY RADIATION RESISTANT;
IT CAN BE DEPOSITED IN THIN LAYERS
(~ 1-100 um);
IT CAN BE DEPOSITED ON DIFFERENT
SUBSTRATES, EVEN FLEXIBLE ONES LIKE
MYLAR AND KAPTON;
IT HAS A BAND GAP VALUE JUST HIGHER THAN c-Si: 1.7-1.9 eV;
MATERIAL DEPOSITION PROCESS CAN BE DONE ON WIDE AREAS.
4

## Slide 5
SEVERAL STANDARD PRODUCTION TECHNIQUES:
PECVD (PLASMA ENHANCED CHEMICAL VAPOR DEPOSITION); [EPFL]
IN THE PROCESS IS USED A MIXTURE OF SILANE (SiH4) AND
HYDROGEN, WITH WORKING TEMPERATURE BETWEEN 160° AND 300°C
RADIOFREQUENCY PECVD (RF OR VHF PECVD), AT VARIOUS
FREQUENCIES.
TWO NEW PROCEDURES ARE UNDER DEVELOPMENT IN THE HASPIDE PROJECT:
→ PULSED LASER DEPOSITION (PLD) [LECCE]
→ REACTIVE SPUTTERING [LECCE]
DEVICE PRODUCTION
5

## Slide 6
HASPIDE GOALS
FABRICATION OF THIN a-Si:H (1- 10 um) IONIZING RADIATION DETECTORS DEPOSITED OVER THIN PLASTIC SUPPORTS TO BE USED FOR:
BEAM MONITORING OF LINACS AND OTHER TYPES OF ACCELERATORS
DETECTION OF RADIATION BURSTS IN SPACE, FOR EXAMPLE SOLAR
ENERGETIC PARTICLES EVENTS;
NEUTRON DETECTION VIA 10B DEPOSITION OVER AN a-Si:H LAYER TO DETECT α PRODUCED BY NEUTRON CONVERSION.
6

## Slide 7
PIN CONTACTS
7
a-Si:H FILM BETWEEN TWO P OR N DOPED LAYERS
DOPED LAYER THICKNESS < 30 nm; INTRINSIC LAYER (1-10 um)
SAME STRUCTURE ON CRYSTALLINE SILICON AND ON KAPTON SUBSTRATE FOR FLEXIBLE DETECTORS

## Slide 8
8
CHARGE SELECTIVE CONTACTS (CSC):TWO THIN LAYERS OF METAL OXIDES SANDWICHING THE SENSITIVE A-SI:H FILM.
FIRST PROTOTYPES:
ELECTRON SELECTIVE CONTACTS:
ZnO, AL OR TiO2
HOLE SELECTIVE CONTACTS:
MoOx
SELECTIVE CONTACT LAYER THICKNESS < 30 nm;
INTRINSIC LAYER (1-10 um)
CSC CONTACTS

## Slide 9
HOMOGENEITY OF SAME SUBSTRATE DEVICES
a-Si:H DEPOSITIONS ON POLYAMIDE (PECVD)
2X2 mm2 AND 5X5 mm2 DEVICES (P-I-N)
THICKNESS: 2.5 um.
POLYAMIDE THICKNESS: 25 um
GOAL: PRODUCE DEVICES TO MONITOR IN REAL TIME IONIZING BEAMS, BOTH IN SHAPE AND FLUX
9

## Slide 10
SENSITIVITY
MEASUREMENTS
LINEAR CORRELATION BETWEEN DOSE RATE
AND SIGNAL
HERE WE CAN SEE THE SENSITIVITIES
NORMALIZED TO THE SENSITIVE VOLUME
MEASURED FOR THE VARIOUS
TYPES OF DETECTORS.
10

## Slide 11
1.TYPE OF CONTACT
NORMALIZING THE SENSITIVITY TO THE VOLUME OF EACH SENSOR, THE BEST SENSITIVITY IS THAT
ASSOCIATED WITH PIN-TYPE DEVICES DEPOSITED ON c-Si.
2. THICKESS
WITH THE SAME CONTACT AND AREA, THE SENSITIVITY SCALES WITH THE THICKNESS
3. AREA
WITH THE SAME THICKNESS AND TYPE OF CONTACT, THE SENSORS WITH A GREATER AREA HAVE A
GREATER SENSITIVITY.
4. SUBSTRATE
WITH THE SAME CONTACT AND VOLUME, DIODES DEPOSITED ON CRYSTALLINE SILICON HAVE GREATER
SENSITIVITY.
5. PERIMETER
WITH THE SAME CONTACT, AREA, AND THICKNESS STRIPS HAVE A LOWER SENSITIVITY THAN SQUARE PADS
SENSITIVITY MEASUREMENTS
11

## Slide 12
MEDICAL LINAC
FIELD SIZE 10 X 10 cm2
6 MV PHOTON BEAM
DR=500 cGY/MIN
FOUR s-Si:H PIN DIODE STRUCTURE OF 2.5 um
THICKNESS ARE FABRICATED ON KAPTON SUBSTRATE.
THE PAD DIODE AREA IS 2.5X2.5 mm2.
SAMPLES ARE MECHANICALLY FIXED TO A KAPTON PIGTAIL 35 CM LONG VIA DOUBLE SIDED STICKY TAPE.
ELECTRICAL CONTACTS ARE MADE WITH CARBON CONDUCTIVE PAINT AND INSULATED. COPPER WIRE WITH 50 um DIAMETER.
UOW ELECTRONICS
12

## Slide 13
SENSITIVITY
THE SENSITIVITY OF THE FOUR PIXELS IS IN THE RANGE OF 140 AND 160 fC/cGY WITH A LINEARITY 0,999 AND THE SIGNAL TO NOISE RATIO IS ABOUT 70.
Florence Hospital
13

## Slide 14
BEAM PROFILE
FIELD SIZE 10 X 10 cm2
6 MV PHOTON BEAM
DR = 500 CGY/MIN
Florence Hospital
14

## Slide 15
SYNCHROTRON
THE DEVICE CONTAINS 4 PIXELS
WITH AN AREA OF 4 × 4 mm2.
THE ELECTRON SELECTIVE CONTACT IS FORMED THROUGH THE DEPOSITION OF 10 nm TiO2 ON GLASS. THE ACTIVE A-SI:H
DETECTOR LAYER IS INTRINSIC (UNDOPED) 0.8 um. THE
HOLE SELECTIVE CONTACTS ON THE TOP OF THE DEVICE IS A 20 nm THICK MOOX LAYER AND
PROTECTED BY 60 nm OF INDIUM TITANIUM OXIDE (ITO)
UOW ELECTRONICS
ANSTO AS AT THE IMAGING AND MEDICAL BEAMLINE (IMBL).
15

## Slide 16
MICROBEAM PROFILE
Australian Synchrotron
16

## Slide 17
SENSITIVE VOLUME
17
Australian Synchrotron

## Slide 18
TRANSMISSION DEVICES
DEVICE WITH 2X2 mm2 P-I-N ON KAPTON SENSORS,
2.5 um THICK
TOTAL THICKNESS ~ 150 um
INSTRUMENTED ACCELERATOR’S EXIT WINDOW
18

## Slide 19
SMALL LINEAR ARRAYS (3-5 PIXELS):
3D CONFORMAL BREAST TANGENTIAL (SKIN DOSE MONITORING IN STEEP DOSE GRADIENT AREAS AT THE SURFACE OF THE BREAST DURING 3D-CRT)
SKIN DOSIMETRY IN THE SOFT TISSUE SARCOMA
LARGE LINEAR ARRAYS (12+ PIXELS):
LINEAR ARRAY COVERING A DISTANCE OF 5 TO 12 cm WITH A RESOLUTION OF 2 mm WOULD BE BENEFICIAL FOR STANDARD LINACS AND TOMOTHERAPY MACHINE QA
FUTURE APPLICATIONS
possible clinical use cases:
Single pixel devices
Skin dosimetry
Total body irradiation in-vivo dosimetry
Dosimetry of extreme high dose rate radiotherapy modalities such as
Synchrotron Microbeam Radiotherapy and FLASH THERAPy
19