## Slide 1: Preliminary analysis at SPS beam test: PSD report
C. Altomare, L. Di Venere, F. Gargano, M.N. Mazziotta, R. Pillera, D. Serini

## Slide 2: Set-up on the beamline
Tile 1: Trigger Tile
Tile 2: PSD Tile
Main trigger (Plastic scintillator on beam axis):
TRIGGER: S0 10x10x0.5 cm tile.
Tile S0
Beam
Beam

## Slide 3: Readout configuration
Tile 1 (bc-408): Trigger Tile
Tile2 (bc-404): PSD Tile
Small SiPM PCB with 3 1x1 Hamamatsu SiPM
Large SiPM PCB with 3 3x3 Hamamatsu SiPM
BetaCHIP configuration
HV = 42.5 V
trigger = 3EXT (TTL) «Syncronicity» with PSD in I2C
mux_path= 1 (low gain)
gain= 15 (lowest gain)
Data Sample
Trigger Threshold 600 mV -> He Cut
Trigger Threshold 100 mV -> proton Visible
PSD
Tile 2
Tile
PSD
Bar L2 (BC-404) 40 cm (BetaChip ASIC)
  EndCap 1: 1 SiPM 3x3 mm
  Top 1: PCB with 3 SiPM 3x3 mm
  Top 2: PCB with 3 SiPM 1x1 mm
  EndCap 2: 1 SiPM 3x3 mm
PCB 3 SiPMs 3x3 mm
PCB 3 SiPMs 1x1 mm
SiPM Model | S14160-3015 | S14160-1315
Size | 3x3 mm2 | 1.3x1.3 mm2
Pixel pitch | 15 um | 15 um
Number of pixels | 39984 | 7284
Refractive index | 1.57 | 1.57
Peak sensitivity wavelength | 460 nm | 460 nm
PDE | 32% | 32%
Vbr | 38 V | 38 V
Beam
Particle fragments 150 GeV/n
A/Z 2.2
Beam size: ~4 cm

## Slide 4: Beam composition expectation from Test Beam 2018
H
Be
Li
He
B
C
N
O
F
Ne
Na
Mg
Al
Si
P
S
Cl
Ar
K
Be
Li
H-He
B
C
N
O
F
Ne
Na
Mg
Al
Si
P
S
Cl
Auto Trigger
Tile 1
Low gain – 8
Test Beam SPS 2022
Trigger scintillator
PMT
Test Beam SPS 2018
Sh
S4
S0
Beam
PMT
Fragment simulation
Preliminary
SiPM

## Slide 5: Tile 1 analysis (Front)
4
Run 201 -202
no veto no finger
Vthr: -600mV
LG15
T 1
3
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
4
3
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
SiPM 3x3 mm
SiPM 1x1 mm
22
T1
TRG
PSD
T2
Data Selection:
From T1 3x3 SiPM size
Preliminary

## Slide 6: Tile 2 analysis (Back)
4
T 2
3
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
4
3
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
SiPM 3x3 mm
SiPM 1x1 mm
22
T1
TRG
PSD
T2
Data Selection:
From T1 3x3 SiPM size
22
Z:22 fragments
Z:22 fragments
Run 201 -202
no veto no finger
Vthr: -600mV
LG15
Preliminary

## Slide 7: PSD L2 analysis (EndCup)
4
3
5
6
7
8
4
3
5
6
7
8
9
10
11
12
SiPM 3x3 mm
SiPM 3x3 mm
T1
TRG
PSD
T2
Data Selection:
From T1 3x3 SiPM size
Z:8 fragments
Z:12 fragments
L2
L2
8
12
Run 201 -202
no veto no finger
Vthr: -600mV
LG0
Preliminary

## Slide 8: PSD L2 analysis (Top)
4
3
5
6
7
8
4
3
5
6
7
8
9
10
11
12
SiPM 3x3 mm
SiPM 1x1 mm
T1
TRG
PSD
T2
Data Selection:
From T1 3x3 SiPM size
Z:11 fragments
Z:19 fragments
L2
L2
9
10
11
11
13
15
17
19
14
16
18
19
Run 201 -202
no veto no finger
Vthr: -600mV
LG0
Preliminary

## Slide 9: PSD S2 analysis
T1
TRG
PSD
T2
S2
S2
Run 201 -202
no veto no finger
Vthr: -600mV
LG0
Preliminary

## Slide 10: PSD S3 analysis
S3
T1
TRG
PSD
T2
S3
Run 201 -202
no veto no finger
Vthr: -600mV
LG0
Preliminary

## Slide 11: Peak resolution comparison
T1 3x3 mm
T1 1x1 mm
T2 3x3 mm
T2 1x1 mm
L2 END0 3x3 mm
L2 END1 3x3 mm
L2 TOP0 3x3 mm
L2 TOP0 1x1 mm
Preliminary
High Z are better resolved
1x1 mm SiPMs have less resolution than 3x3 mm
Bars and back Tile have more or less the same resolution
3 Top SiPM have more or less the same resolution of 2 side SiPM
Back detectors have less resolution than front tile (probably due to the fragmentation)

## Slide 12: Quenching effect in plastic scintillator
T1 3x3 mm
T1 1x1 mm
T2 3x3 mm
T2 1x1 mm
L2 END0 3x3 mm
L2 END1 3x3 mm
L2 TOP0 3x3 mm
L2 TOP0 1x1 mm
BAR LG0
BAR LG0
Tile LG15
Preliminary

## Slide 13: Quenching effect in plastic scintillator
Core scintillation
Halo scintillation
Beam test 2018
*Fixed for Tile fits
[1]Robert DWYER and Dazhuang “ZHOU PLASTIC SCINTILLATOR RESPONSE TO RELATIVISTIC NUCLEI, Z< 28”
[2] P.S. Marrocchesi at all “Beam test performance of a scintillator-based detector for the charge identification of relativistic ions”
[3]De Benedictis et al “Characterization of plastic scintillator tiles equipped with SiPMs for the High Energy cosmic-Radiation Detection (HERD) experiment
[4]Altomare , Serini et al “A full and customizable simulation of a scintillation tile equipped with SiPMs for Plastic Scintillator Detectors in the next generation of satellite experiments.”
Preliminary
Chart: 
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
Fh | 0.392 | 0.36 | 0.38 | 0.38 | 0.2993 | 0.3395 | 0.3344 | 0.3825
Ref | num | Fh | Err | Kb*
Dwyer [1] | 1 | 0,392 | 0,004 | 
Marrocchesi [2] | 2 | 0,36 | 0,01 | 
Beam Test 2018 [3] | 3 | 0,38 | 0,062 | 
Simualtion [4] | 4 | 0,38 | 0,019 | 
T1 3x3 | 5 | 0,2993 | 0,01297 | 0,0039
T1 1x1 | 6 | 0,3395 | 0,02363 | 0,0039
T2 3x3 | 7 | 0,3344 | 0,03051 | 0,0039
T2 1x1 | 8 | 0,3825 | 0,04701 | 0,0039
Preliminary

## Slide 14
To Do List
Sync analysis with other subsystems (e.g. IT SCD …): Ongoing
Gaussian/Multigaussian fit for Peak resolution: Done
Quenching effect studies with PSD scintillators and literature results : Done
Channel equalization with simulation : Ongoing
Comparison of runs with and without fingers: Ongoing
Comparison of runs with other beam position: Ongoing

## Slide 15: BackUp Slides

## Slide 16
Ref | num | Fh | Err | kb
Dwyer | 1 | 0,392 | 0,004 | 
Marrocchesi | 2 | 0,36 | 0,01 | 
Beam Test 2018 | 3 | 0,38 | 0,062 | 
simualtion | 4 | 0,38 | 0,019 | 
T1 3x3 | 5 | 0,2993 | 0,01297 | 0,0039
T1 1x1 | 6 | 0,3395 | 0,02363 | 0,0039
T2 3x3 | 7 | 0,3344 | 0,03051 | 0,0039
T2 1x1 | 8 | 0,3825 | 0,04701 | 0,0039
L2E03x3 | 9 | 0,6 | 0,3008 | 0,0039
L2E13x3 | 10 | 0,6 | 0,3612 | 0,0039
L2TOP03x3 | 11 | 0,6 | 0,3681 | 0,0039
L2TOP01x1 | 12 | 0,6 | 0,289 | 0,0039
Chart: 
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8
Fh | 0.392 | 0.36 | 0.38 | 0.38 | 0.2993 | 0.3395 | 0.3344 | 0.3825

## Slide 17: PSD @ SPS Tile T1 1x1 mm
4
Run 201 -202
no veto
no finger
Vthr: -600mV
3
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22

## Slide 18: PSD @ SPS Tile T1
Run 201 -202
no veto
no finger
Vthr: -600mV

## Slide 19: PSD @ SPS T1 Saturation 3x3 mm

## Slide 20: PSD @ SPS T1 Saturation 1x1 mm

## Slide 21: PSD @ SPS T2 3x3

## Slide 22: PSD @ SPS T2 1x1

## Slide 23: PSD @ SPS Scatter plots

## Slide 24: PSD @ SPS T2 3x3 Saturation

## Slide 25: PSD @ SPS T2 1x1 Saturation

## Slide 26: PSD @ L2 Bar EndCup0

## Slide 27: PSD @ L2 Bar EndCup1

## Slide 28: PSD @ SPS L2 Bar EndCup
Tile

## Slide 29: PSD @ SPS L2 Bar EndCup0 Saturation

## Slide 30: PSD @ SPS L2 Bar EndCup1 Saturation

## Slide 31: PSD @ SPS L2 Bar TOP0 3x3

## Slide 32: PSD @ SPS L2 Bar TOP0 1x1

## Slide 33: PSD @ SPS L2 Bar TOP0 3x3 Saturation

## Slide 34: PSD @ SPS L2 Bar TOP0 1x1 Saturation

## Slide 35: PSD @ SPS L2 Bar TOP0