## Slide 1: Readout Electronic Update
Andrea Fabbri on behalf of Electronic Group

## Slide 2
Electronic Readout Scheme
We had two electronic readout proposal until the FDR (end of April 2021).
The review committee suggestd to merge the two proposal:
  FEB (Front End Board - inside the cryostat) from the discrete solution.
  FEC (Front End Controller – outside the cryostat) based on the prototype coming from the ASIC solution.
FEB
  2 channel 1 tiles
  4024 tiles
  Total 8048 channels
  Analog signals from FEB will be transferred to FEC via differential pairs, 3-4 m inside the SS tank, ~10 m outside the tank
FEC
  ADC is on FEC, used to digitize analog signals from FEB
  FPGA & Power boards in MicroTCA.4 crate
  Q/T information is extracted with FPGA (waveform analysis)

## Slide 3
FEB PRR Results

## Slide 4
TAO Experiment
FEBs mock-up

## Slide 5
FEB – planar version

## Slide 6
FEB – planar version
FR4 FEB prototype
Aramid (low background) FEB PCBs
100 pre-production batch

## Slide 7
FEB Interfaces
HV supply
3 mt cable inside cryostat
+ 10-12 mt outside
LV supply
3 mt cable inside cryostat
+ 10-12 mt outside
Tile
Board-to-board mating
Signal output
Differential pairs,
2-3 mt cable inside crysostat
+10-12 mt outside
LV supply
10 cm cable for the next FEB
(see next slide)
0R resistor
All the cable jackets will be made in PTFE to resolve the LAB compatibility issues

## Slide 8
FEB reliability
Longest run @ -50° : 8 days
Integrated days @ -50° : 65 days
Number of cycles +20°/ -50° : ~100
Number of FEB prototypes : ~20
Number of failures : 0
N.B. ~4000 FEB with 10k years MTBF  1 failure every 2,5 years
from
Analog Design
FEB MTBF (Mean Time Between Failure)

## Slide 9
FEB Dynamic Range
Single P.E. maximum amplitude ~ 20 mV
0V-2V output linearity range (compatible with ADC input range)
P.E. Channel dynamic range 1 – 100 p.e. (1-200 p.e. on the whole tile)
IBD occupancy from MC simulations (75 cm radius)

## Slide 10
Test result – recovery time after muon scintillation
We reduced the gain (1/20) of 1 FEB channel to have a «self calibration» of the light pulse.
The «rebounce» happens after the TIA saturation.
In the circled area the photon counting can be executed but it needs a careful evaluation of the baseline.
500 p.e.
1000 p.e.
2000 p.e.
2u 4u 6u
Standard
gain
Low
gain
Standard
gain
Low
gain
Standard
gain
Low
gain

## Slide 11
Peak maximum [mV]
Test result – improving trigger performance
All the plots are taken with the same laser configuration.
In the first plot the trigger comes from the laser itself.
On the others the trigger is setted on the FEB signal («self-trigger»).
The self trigger performance can be enhanched by adding a simple «time over threshold» constrain
Counts [au]
Counts [au]
Counts [au]
Counts [au]
Laser trigger
½ p.e. threshold
¼ p.e. threshold
¼ p.e. threshold +
150 ns time over threshold

## Slide 12
Compatibility issue
Connector alone is ok
A first test with Expoxy paint improves the performance of the whole board
Better results with sprayng methods may be achieved, test in progress.
From Zhangquan compatibility test results

## Slide 13
Each FEC can manage 96 channel  4 FEC can manage all Veto channels
Veto Electronic Readout Architecture
VETO dedicated FEB
(4 SiPM on the bottom
side, directly coupled
to the optical fiber output)
Crystal – FEB adapter
FEB shape and mechanics
discussed with the VETO group

## Slide 14
Flange design
8048 differential pair for signal readout
8 flanges, 1006 differential pairs each
6 «signal feedthrough» PCB on each flange (84 pairs each)
2 «power feedthrough» PCB on each flange (50 A + 50 A each)
External side
Internal side
Power
153 mm
10 simplified Signal PCB delivered last year (8 in IHEP, 2 in Roma)
Signal PCB
Power PCB
Power

## Slide 15
FEB Power Supply – Inside Tank
Flange
25 pin
DSUB
connectors
Signal cabling (1 cable/FEBs)
Signal cabling (1 cable/FEBs)
To FEC/ADC boards
To FEC/ADC boards
Signal cabling (1 cable/ 6 FEBs)
Power cables from
LV power supplies
(standard DSUB connectors)
9 pin DSUB connector
for V/I monitoring

## Slide 16
Cabling
Cable FEP/FEP (600V) Liv6YC6Y 3x0,25 mm2 – Ø 3,40 mm
Cond. 19x0,127 mm solered copper – isol. d. 1,05 mm
Shield: soldered copper twist
Cable FEP/FEP (600V) Liv6YC6Y-TP 2x2x0,25 mm2 – Ø 5,00 mm
Cond. 19x0,127 mm solered copper – isol. d. 1,05 mm
Shield: soldered copper twist
Shippment time: 25 weeks, imported from China.
Asked for a sample (few meters) to do compatibility test in IHEP.

## Slide 17
Cabling
Cavi interni
Cavi esterni

## Slide 18
FEB Power Supply – Inside Tank
Each FEB needs 80 mA at +2V and 80 mA at -2V
Each flange have to supplies ~500 FEBs.
FEBs can be suddivided in ~86 rows with 6 FEBs each.
86 long pairs come from the flange to the first FEB of each FEB rows
Shorter cables can be used to bring supply to the other FEBs on the row
Flange
FEB
FEB
FEB
FEB
FEB
FEB
FEB
FEB
FEB
FEB
FEB
FEB
FEB
FEB
FEB
FEB
FEB
FEB
FEB
FEB
Inside cryostat
LV Supply
Connectors
FEB
FEB
FEB
FEB
FEB
FEB
FEB
FEB

## Slide 19
Isolated +2V
Isolated +2V
GND
FEB Supply Modules
+2V
-2V
+12V
from uTCA connector
+1.8V ADC
2.5V, 1.8V, 1,2V, 1V FPGA
FEC
+12V
from uTCA connector
FPGA
SiPM
SiPM
Supply
+70V / +110V
GND
FEC Supply
ADC
ADC
ADC
GND
Differential Output
FEB
Power Distribution

## Slide 20
Flange PCB
(analog signal)
Flange PCB
(low Voltages)
+2
GND
-2
+2V
GND
-2V
Flange PCB
(SiPM bias)
+Vbias
GND
FEB
0R / NM
FEC
GND
LV Supply
+2.2V
GND
-2.2V
HV Supply
GND
GND
+Vbias
GND
SiPM
+Vbias
GND
NM / 0R
Copper
Grounded
Clean
Ground
shielding
Grounding
Different grounding schematics are feasible with the current hardware.
We are going to setup lab to study the different grounding connection.
The results of this study will be tested in the 1-to-1 prototype setup in IHEP.
Inside Outside

## Slide 21
ADC Prototype
FPGA
250 Msample/sec ADC,125 MHz bandwidth
16 channels
Differential input
2 Vpp input
Dedicated Input for shared clock between different board
Prototype system:
ADC board with AD9083
FPGA controller (Kintex7)
Ethernet connection with PC
100 kHz laser
ADC Board:
Two AD9083 – 16 Channels
Dedicated PLL
Working on the prototype,
still managing the external
connector.

## Slide 22
RTM-v2
u4FCP-v1
FEC and ADC Board
ADC Board - Roma3
FEC Board - IHEP

## Slide 23
Production schedule
Part Number | Board | Lead Time
LTC6269 (8400 pcs) | FEB | Scheduled for May
LTC6405 (8400 pcs) | FEB | Received
AD9083 (550 pcs) | ADC | Included in the bid
LTM8078 (290 pcs) | ADC | Included in the bid
AD9528 (290 pcs) | ADC | Received
100 FEB pre-production received and tested. 50 of them shipped to China.
4200 FEBs
should be ready
during the summer
275 ADC boards
should be ready
for September
Components for the full production
Board | PCB status | Lead Time
Signal Flange | Received | 2° week of April
Power Flange | Received | 2° week of April
Test system for FEB | 1° week of April | 3° week of April
VETO FEB | Received | 1° week of April
The full production probably
will be done in China