## Slide 1
Student: Giovanni Celotto
Supervisor: Vadim Rusu
Final term presentation
19 September 2022
Measurement of the gas gain and understanding the gas flow in the Mu2e Tracker

## Slide 2: Overview
Mu2e experiment and tracker
Aim of the project
Experimental setup
Measurement of the current
Results:
  Effect of water
  Effect of humidity, temperature, pressure
Conclusions and Perspectives

## Slide 3: Mu2e Experiment
Neutrino-less conversion of muons into electrons in the Coulomb field of an Al nucleus
PS
TS
DS

## Slide 4: Mu2e Tracker
Low mass straw tube detector
Position and the momentum of the electron
Still under construction

## Slide 5: Mu2e Tracker’s panel
96 Mylar straws
5 mm diameter and 15 µm walls
25 µm gold plated tungsten sense wire

## Slide 6: Mu2e Tracker’s panel
96 Mylar straws
5 mm diameter and 15 µm walls
25 µm gold plated tungsten sense wire

## Slide 7: What’s the problem?
Higher Voltage
Higher gain
Measurement with no sources
No ionization products
Small currents

## Slide 8: What’s the problem?
Measurement with no sources
No ionization products
Small currents
Higher Voltage
Higher gain

## Slide 9: What’s the problem?
Measurement with no sources
No ionization products
Small currents
Higher Voltage
Higher gain

## Slide 10: The Experimental Apparatus
Gas Flowmeter
Ar- 0.14 SCFH
H20 0.1 SCFH
HV supplier
1450 V
Slow Amplifiers
24 (2 circuits each)
Panel MN084
Gas Flowmeter
HV supplier
Slow Amplifiers
Raspberry Pi

## Slide 11
Measuring the current
A
A
C
ANODE
CATHODE
ANODE
ANODE
CATHODE
Two rows of straws
Measurement is performed on the cathode at low potential

## Slide 12
Measuring the current
Two rows of straws
Measurement is performed on the cathode at low potential
ANODE
CATHODE
ANODE
ANODE
CATHODE

## Slide 13: New data
Conditioning effect
The first time HV is applied high currents are measured

## Slide 14: Currents before the water

## Slide 15: Old Data vs New Data before flushing water

## Slide 16: Old Data vs New Data before flushing water
After 1 month
Reduction in currents occurred before the water vapour was introduced
Some currents increased of few nA

## Slide 17: Effect of the water

## Slide 18: Effect of the water
After flushing water vapour
Further decrease in the highest currents without going to zero
Daily periodic oscillations
Secondary oscillation with 15 min-period

## Slide 19: Effect of the water
Daily periodic oscillations
Secondary oscillation with 15 min-period
After flushing water vapour
Further decrease in the highest currents without going to zero

## Slide 20: Daily oscillations
B (%) | A (%) | τ (hours)
20.05 ± 0.03 | 26.606 ± 0.001 | 5.541 ± 0.003
B (%) | A (%) | τ (hours)
20.05 ± 0.03 | 26.606 ± 0.001 | 5.541 ± 0.003

## Slide 21: Why 20% of humidity in the panel?
We have 4 possibilities
The sensor does not go lower than 20%
The sensor is working. There is a leak
The sensor is broken. The humidity is actually 0%
The humidity is actually 20%

## Slide 22: Why 20% of humidity in the panel?
We have 4 possibilities
The sensor does not go lower than 20%
The sensor is working. There is a leak in the panel
The sensor is broken. The humidity is actually 0%
The humidity is actually 20%

## Slide 23: Why 20% of humidity in the panel?
We have 4 possibilities
The sensor does not go lower than 20%
The sensor is working. There is a leak in the panel
The sensor is broken. The humidity is actually 0%
The humidity is actually 20%

## Slide 24: Why 20% of humidity in the panel?
We have 4 possibilities
The sensor does not go lower than 20%
The sensor is working. There is a leak
The sensor is broken. The humidity is actually 0%
The humidity is actually 20%

## Slide 25: Why 20% of humidity in the panel?
We have 4 possibilities
The sensor does not go lower than 20%
The sensor is working. There is a leak
The sensor is broken. The humidity is actually 0%
The humidity is actually 20%

## Slide 26: 15-min period oscillations

## Slide 27: Conclusions
Conditioning effect has played by far the most dominant role on the reduction of the currents
Water has had a little effect on reducing small currents. More measurements are needed
Environmental conditions strongly affect the currents of the panel

## Slide 28: Conclusions
Conditioning effect has played by far the most dominant role on the reduction of the currents
Water has had a little effect on reducing small currents. More measurements are needed
Environmental conditions strongly affect the currents of the panel

## Slide 29: Conclusions
Conditioning effect has played by far the most dominant role on the reduction of the currents
Water has had a little effect on reducing small currents. More measurements are needed
Environmental conditions strongly affect the currents of the panel

## Slide 30: Thank you for your attention

## Slide 31: Calibration of the humidity sensor
 | Humidity 1 (%) | Humidity 2 (%)
Outside | 55 | 52
Inside | 58 | 100
Humidity of the cleanroom is 60 %

## Slide 32: Pedestal Measurement

## Slide 33: Environmental conditions

## Slide 34: Environmental conditions
Pressure seems to have no clear effects on the currents