## Slide 1
Anna Vignati
Dip. Fisica Università degli Studi di Torino e INFN Torino
The challenges of beam monitoring from particle therapy
to FLASH radiotherapy.

## Slide 2
Anna Vignati – UniTo and INFN
Beam
TARGET
…is the continuous check of beam parameters to correct or interrupt the irradiation in case of values out of clinical tolerances. [Giordanengo, Palmans Med Phys. 2018]
Beam
The role of beam monitors…

## Slide 3
Ultra-high dose rate radiotherapy (FLASH)
The role of beam monitors…in the near future?
Anna Vignati – UniTo and INFN
Technology innovation in charged particle therapy

## Slide 4
Anna Vignati – UniTo and INFN
Radiation delivered in < 200 ms
at Ultra-High Dose Rate
(average dose rate > 40 Gy/s)
The role of beam monitors…in UHDR RT
The roles of dose-per-pulse, instantaneous dose-per-pulse, pulse duration and frequency still remain to be understood.
few flash cards on FLASH RT
[Di Martino et al 2020 Frontiers]
Because of the heavy saturation (i.e. non linearity) of conventional transmission chambers in UHDR,
new beam monitoring devices are needed
[Oesterle et al. 2021 J App Clin Med Phys]
Decreasing toxicity maintaining tumor control: FLASH effect

## Slide 5
Anna Vignati – UniTo and INFN
F Romano et al. Ultra-high dose rate dosimetry: Challenges and opportunities for FLASH radiation therapy. Medical Physics, 2022
Beam temporal structure

## Slide 6
Anna Vignati – UniTo and INFN
All you need is…
High temporal resolution
  ideally: intra-pulse monitoring;
  reaction time: inter-pulse;
High spatial resolution
  Flatness, symmetry
  Maximum acceptable dose uncertainty
Beam transparency
Large response dynamic range
Large sensitive area
(Beam energy)
Radiation hardness

## Slide 7
Anna Vignati – UniTo and INFN
optimization of existing technologies,
investigation of novel radiation detection methods.
Ionization chambers for proton UHDR beams
The role of beam monitors…in UHDR RT

## Slide 8
Anna Vignati – UniTo and INFN
1Department of Radiation Oncology, Hospital of the University of Pennsylvania, Philadelphia, USA.
2Pyramid Technical Consultants, Inc, Boston, USA.
3IBA PT Inc, Louvain-La-Neuve, Belgium.
Wei Zou, Eric S Diffenderfer et al.
High-resolution position sensing transmission ionization chamber with a purpose-built multichannel electrometer for both conventional and FLASH dose rate proton RT.
Polyimide film as electrode substrate
spot positioning accuracy 0.10 ± 0.08 mm
ks <1,005 with <0,5% uncertainty
good linearity (compared with Faraday cup measurements)
64x64 mm2 sensitive area
3 active chambers [1 with thin 1 mm gap integral electrode; two with two sets of 64 parallel strip electrodes of 1 mm width].
Air (Helium or Argon)
Bias voltage >200 V (lower than 1000 V/mm)

## Slide 9
Anna Vignati – UniTo and INFN
The Heidelberg Ion-Beam Therapy Center (HIT) synchrotron extracts ≥ 5·108 12C ions in 150 ms.
The beam monitors were filled with helium.
Tinganelli et al.
GSI Helmholtzzentrum fuer Schwerionenforschung, Darmstadt, Germany
Heidelberg Ion-Beam Therapy Center, Heidelberg, Germany
20x20 mm2 sensitive area
2 active gaps [10 mm each].
Bias voltage >500 V

## Slide 10
Anna Vignati – UniTo and INFN
optimization of existing technologies,
investigation of novel radiation detection methods.
Ionization chambers for charged particle UHDR beams
Beam current transformers for electron UHDR beams
The role of beam monitors…in UHDR RT

## Slide 11
Anna Vignati – UniTo and INFN
Jorge et al. - Lausanne, Switzerland
Oriatron eRT6
Limitation: inability to monitor the beam's spatial profile (flatness and symmetry)
Do we still need to check the beam profile?
Very High Electron Energy (VHEE)…pencil beams?

## Slide 12
Anna Vignati – UniTo and INFN
Jorge et al. - Lausanne, Switzerland
Oriatron eRT6
Limitation: inability to monitor the beam's spatial profile (flatness and symmetry)
Do we still need to check the beam profile?
Very High Electron Energy (VHEE)…pencil beams?
Probably, only the beam energy can change
In a radiofrequency powered linac, the power generated by the RF source can be written as:
Energy variation
[Di Martino et al. Front. Phys.2020]

## Slide 13
Anna Vignati – UniTo and INFN
Jorge et al. - Lausanne, Switzerland
Oriatron eRT6
Limitation: inability to monitor the beam's spatial profile (flatness and symmetry)
Exit window
Primary scattering foil
Secondary scattering foil
Ion chambers
Upper BCT
Lower BCT
Optics
Volume24, Issue2,
February 2023
Liu et al. – MD Anderson Cancer Center
Do we still need to check the beam profile?
Mobetron
Very High Electron Energy (VHEE)…pencil beams?
Probably, only the beam energy can change

## Slide 14
Anna Vignati – UniTo and INFN
optimization of existing technologies,
investigation of novel radiation detection methods.
Scintillators, air fluorescence, solid state detectors (et al…)
The role of beam monitors…in UHDR RT
FRIDA INFN project
?

## Slide 15
Anna Vignati – UniTo and INFN
large area coverage (30 cm × 30 cm);
beam transmissive (<1 mm water equivalent);
a linear response over a broad dynamic range;
radiation tolerance;
real-time analysis IEC-compliant fast beam-interrupt signal;
two-dimension beam imaging with spatial resolution < 0,5 mm.
University of Michigan, Loma Linda University Hospital,
Integrated Sensors LLC

## Slide 16
Anna Vignati – UniTo and INFN
large area coverage (30 cm × 30 cm);
beam transmissive (<1 mm water equivalent);
a linear response over a broad dynamic range;
radiation tolerance;
real-time analysis IEC-compliant fast beam-interrupt signal;
two-dimension beam imaging with spatial resolution < 0,5 mm.
University of Michigan, Loma Linda University Hospital,
Integrated Sensors LLC

## Slide 17
Anna Vignati – UniTo and INFN
Dipartimento di Fisica, Sapienza Università di Roma
INFN sezione Roma I
Active volume of 2x2x60 cm3 air enclosed by a 5 um thick layer of black Tedlar
light sensors on the two opposite edges
Beam position
Number of electrons
9 MeV UHDR pulsed electron beams accelerated from the ElectronFLASH linac at the Centro Pisano for FLASH Radiotherapy

## Slide 18
Anna Vignati – UniTo and INFN
Università degli Studi di Torino INFN Torino, Italy
Vignati, Giordanengo, Fausti, Marti Villarreal, Mas Milian, Mazza, Shakarami, Cirio, Monaco, Sacchi
INFN Catania, SIT Sordina, CPRF Pisa, INFN e Università Torino, STLab et al.
Romano, Milluzzo, Di Martino, D’Oca, Felici, Galante, Gasparini, Mariani, Marrale, Medina, Pacitti, Sangregorio, Vanreusel, Verellen, Vignati, Camarda
UNPUBLISHED RESULTS!
TERA08 chip
30/45𝜇m thickness
RC
2.2 mm2
FPGA
Oscilloscope
PiN planar silicon sensor
Signal integration
9 MeV e- UHDR beams @ CPFR EF linac

## Slide 19
Anna Vignati – UniTo and INFN
Università degli Studi di Torino INFN Torino, Italy
Vignati, Giordanengo, Fausti, Marti Villarreal, Mas Milian, Mazza, Shakarami, Cirio, Monaco, Sacchi
INFN Catania, SIT Sordina, CPRF Pisa, INFN e Università Torino, STLab et al.
Romano, Milluzzo, Di Martino, D’Oca, Felici, Galante, Gasparini, Mariani, Marrale, Medina, Pacitti, Sangregorio, Vanreusel, Verellen, Vignati, Camarda
UNPUBLISHED RESULTS!
30/45𝜇m thickness
2.2 mm2
Signal integration
TERA08 chip
RC
FPGA
Oscilloscope
PiN planar silicon sensor

## Slide 20
Anna Vignati – UniTo and INFN
30/45𝜇m thickness
Abacus chip
2.2 mm2
Digitizer
Ultra-Fast Silicon Detector (UFSD) (LGAD technology)
Signal integration
Thin p+ gain layer implanted under the n++ cathode
Controlled low gain (~ 10-30)
Gain increases with bias voltage
High Signal/Noise ratio
Small thickness of active volume 50 µm
Reduced beam perturbation, signal steepness
Loss of signal due to reduced thickness compensanted by internal gain
227 MeV protons
Threshold
time [ns]
Signal [ADC counts]
FPGA

## Slide 21
Anna Vignati – UniTo and INFN
30/45𝜇m thickness
2.2 mm2
Ultra-Fast Silicon Detector (UFSD) (LGAD technology)
Signal integration
Thin p+ gain layer implanted under the n++ cathode
Controlled low gain (~ 10-30)
Gain increases with bias voltage
High Signal/Noise ratio
Small thickness of active volume 50 µm
Reduced beam perturbation, signal steepness
Loss of signal due to reduced thickness compensanted by internal gain
Short signal duration (< 2 ns) → particle counting
Excellent time resolution (tens of ps)
→ beam energy measurement from time-of-flight
Abacus chip
Digitizer
FPGA

## Slide 22
Anna Vignati – UniTo and INFN
30/45𝜇m thickness
2.2 mm2
Ultra-Fast Silicon Detector (UFSD) (LGAD technology)
Signal integration
Thin p+ gain layer implanted under the n++ cathode
Controlled low gain (~ 10-30)
Gain increases with bias voltage
High Signal/Noise ratio
Small thickness of active volume 50 µm
Reduced beam perturbation, signal steepness
Loss of signal due to reduced thickness compensanted by internal gain
Abacus chip
Digitizer
FPGA
single
particle
counting
time
resolution
excellent
Short signal duration (< 2 ns) → particle counting
Excellent time resolution (tens of ps)
→ beam energy measurement from time-of-flight

## Slide 23
The role of beam monitors…and online range verification in PT
Annihilation photons from positron emitters
prompt photons from nuclear de-excitation
secondary charged fragments
Zhu X, Fakhri GE. Theranostics. 2013;3(10):731-740.
Charged particles
𝛄 511 keV
𝛄 511 keV
neutrons
𝛄 prompt
beam
2.2 mm2
UFSD
single
particle
counting
time
resolution
excellent
courtesy of E. Fiorina, F. Pennazio, V. Ferrero

## Slide 24
Annihilation photons from positron emitters
prompt photons from nuclear de-excitation
secondary charged fragments
Charged particles
𝛄 511 keV
𝛄 511 keV
neutrons
𝛄 prompt
beam
PI: Francesco Pennazio
INFN CSN5 Young Researcher grant
I3PET
Prompt
gamma
timing
2.2 mm2
time
resolution
excellent
single
particle
counting
UFSD
PI: Veronica Ferrero
INFN CSN5 Young Researcher grant
courtesy of E. Fiorina, F. Pennazio, V. Ferrero
The role of beam monitors…and online range verification in PT

## Slide 25
Annihilation photons from positron emitters
prompt photons from nuclear de-excitation
secondary charged fragments
Charged particles
𝛄 511 keV
𝛄 511 keV
neutrons
𝛄 prompt
beam
PI: Francesco Pennazio
INFN CSN5 Young Researcher grant
I3PET
Prompt
gamma
timing
2.2 mm2
time
resolution
excellent
single
particle
counting
UFSD
PI: Veronica Ferrero
INFN CSN5 Young Researcher grant
First PGT measurements
Primary beam
Secondary radiation
courtesy of E. Fiorina, F. Pennazio, V. Ferrero
Prompt g
The role of beam monitors…and online range verification in PT

## Slide 26
Anna Vignati – UniTo and INFN
Integrate Dose Delivery System and Range Verification System
Combine different signals
Improve clinical performance
PGT TOF distribution
227 MeV protons on PMMA
Beam average rate 107 pps
beam
beam
courtesy of E. Fiorina, F. Pennazio, V. Ferrero
The role of beam monitors…and online range verification in PT

## Slide 27
Anna Vignati – UniTo and INFN
There are not yet validated and standardized protocols for UHDR beam dosimetry for FLASH RT and detectors for dosimetry and beam monitoring have not been fully established.
[F Romano et al. Medical Physics, 2022]
The future of proton and heavy ion therapy therefore depends on technological improvements to make particle therapy cheaper, more conformal and faster.
[C Graeff, L Volz, M Durante. Progress in Particle and Nuclear Physics, 2023]
…to reduce treatment margins
Maybe…
a combination of different technologies would be a viable solution

## Slide 28
Anna Vignati – UniTo and INFN
anna.vignati@unito.it
Thank you!