## Slide 1: ULTRA HIGH VACUUM TECHNIQUES
Frascati, LNF – INFN, 17 May 2018 D. Alesini, S. Bini, F. Cioeta

## Slide 2
Acknowledgements
This presentation contains the importants concepts taken from:
Vacuum technology advanced lessons - Paolo Michelato INFN Sezione di Milano Laboratorio Acceleratori e Superconduttività Applicata;
Pfeiffer Vacuum Know-How;
Agilent Vacuum Products Catalog .

## Slide 3: Index
INTRODUCTION
Vacuum Definition and Applications
Pressure measurements and Vacuum classification
Vacuum in particle accelerators
BASIC CONCEPTS
Mean free path and Gas Flow Regimes ( Transition, Molecular and Viscous)
Gas flow rate and pumping speed
Desorption, Outgassing and Degassing
Leaks
Throughput Q=p*S
Pumping Speed
Residual Gas Composition
Vacuum Conductance, Series and Parallel
Electrical Analogy and Examples
PUMPING TECHNOLOGY
Primary pumping systems: Scroll Pump, Turbomolecular Pump,
UHV pumping system: Ion Pump, Getter Pump,
Vacuum Gauge
Leak Detectors
VACUUM SYSTEM DESIGN
Materials
Clean Process
Laboratory experience description

## Slide 4
Vacuum Definition and Applications
Pressure measurements and Vacuum classification
Vacuum in particle accelerators
    INTRODUCTION

## Slide 5
What is vacuum? How we define it?
• Ideal
Classical metaphysics: a space that contains nothing
• Real
Whatever has a pressure below atmospheric pressure
Vacuum science studies behavior of rarefied gases, interactions between gas and solid surfaces (adsorption and desorption), etc.
Vacuum technology covers wide range of vacuum pumping, instrumentations, materials engineering, and surface engineering,…
• In practice
Any volume that has a number of gas molecules per unit of volume less than that of the atmosphere that surrounds.

## Slide 6
a) To Avoid chemical and physical processes caused by atmospheric gases (e.g: during the fusion of particular reactive metals, like Ti,…)
APPLICATIONS
b) To increase the mean free path of molecules, atoms and ions avoid the impacts with residual gas molecules (e.g: Metalization processes under vacuum, particles accelerators, ion implantation,…)
c) To increase thermal insulation ( e.g: in the Dewars, criogenics systems)
d) To simulate some particulars physical situations (e.g. chamber of space simulation for test on satellites or space stations)
e) Food and packaging, brazing in furnaces, sputtering processes..

## Slide 7
Pressure measurements and Vacuum Classification
High Vacuum (HV)
10-3 to 10-6 mbar
Extreme UltraHigh Vacuum (XHV)
~10-12 mbar
Low Vacuum
(LV)
30 to 103 mbar
Medium Vacuum (MV)
10-3 to 30 mbar
Industrials Scope
Ultra High Vacuum (UHV)
10-9 to 10-12 mbar
Very High Vacuum (VHV)
10-6 to 10-9 mbar
Particles Accelerator -P~10-8-10-10 mbar
NB: P=10-10 mbar  106 molecules/cm3 !!!
1 cm3 of air at atmospheric pressure contains about 1019 molecules

## Slide 8
1) Circular machines like synchrotrons (multi-passage, high current)
The interaction between the residual gas and the particles beam can have several effetcs:
 reduction of beam lifetime (because of scattering and energy lost by bremsstrahlung). The lifetime is proportional to 1/P where the P is the residual gas pressure.
 instability of the stored particles beam (ion trapping, fast ion instability)
 betatron tune variation
 Increase in beam emittance
Vacuum in Particle Accelerators
2) Linear accelerators (single-passage, low current)
In LINAC the vacuum requirements are less demanding because of the single passage (no cumulative effects) and less average current. The vacuum can still have impact on:
 Increase in beam emittance
 discharges in high gradient (10-100 MV/m) accelerating structures
 Contaminations of targets, …
Typical vacuum pressures in synchrotons are 10-8-10-11 mbar
Typical vacuum Pressures are 10-8-10-9 mbar

## Slide 9
Mean free path and Gas Flow Regimes ( Transition, Molecular and Viscous)
Gas flow rate and pumping speed
Desorption, Outgassing and Degassing
Leaks
Throughput Q=p*S
Pumping Speed
Exercise
Residual Gas Composition
Vacuum Conductance, Series and Parallel
Electrical Analogy and Examples
BASIC CONCEPTS

## Slide 10
The mean free path is the average distance that a gas molecule can travel before colliding with another molecule and is determined by:
- Size of molecule (2r)
- Pressure (p)
- Temperature (T)
a =
K
T
(2r)2 p
 2
.
Gas Flow Regimes
The gas in a vacuum system can be in a viscous state, in a molecular state (or in a transition state) depending on the dimension-less parameter know as the Knudsen number (Kn) that is the ratio between the mean free path and the characteristic dimension of the flow channel.
Viscous Flow
(momentum transfer
between molecules)
Molecular Flow
(molecules move
independently)
P> 1 mbar
P<10-3 mbar
Typical pressure values
Notes: Flusso molecolare:Per gas rarefatti la probabilita’ di urto fra
le molecole del gas e’ << della proba8
bilità di collisione con le pareti del
condotto. E’ il regime normalmente
valido in vuoto

## Slide 11
P(mbar) | 1000 | 1,33 | 0,00133 | 1,33x 10-6 | 1,33*10-9
λ (cm) | 6,6x10-6 | 5,1x10-3 | 5,1 | 5100 | 5,1x106
mean free path Air @ 22°C

## Slide 12
In a vacuum system (to the 1° order) the total gas load is the sum of several contributions.
The main important (for our typical applications) are:
Desorption or outgassing ( Qdes);
Leaks (Real or Virtual , QL)
SYSTEM
PUMP
Desorption or Outgassing (Qdes)
Leaks (QL)
Pumping speed (S)
P(t)
Gas flow rate and pumping speed
The final pressure in the system is given by the equilibrium between the total gas load (Qtot) and the gas flowing into the pump (pS).
The pressure p(t) can be obtained solving this equation:
Qtot = Qdes+ QL
QpV

## Slide 13
Gas molecules, (primarily water) are bound to the interior surfaces of the vacuum chamber and gradually desorb again under vacuum. The desorption rate of the metal and glass surfaces in the vacuum system produces a gas yield that decreases over time.
Qdes | Desorption rate | [mbar l s-1]
qdes | Desorption rate density or specific desorption rate | [mbar l s-1 m-2]
A | Area | [m2]
Desorption, outgassing, degassing
Qdes(t)=A*qdes(t)
mbar
mbar
mbar
mbar
t
Qdes
s-1cm-2
s-1cm-2

## Slide 14
QL describes the leak rate, i.e. a gas flow, which enters the vacuum system through leaks.
The leakage rate is defined as the pressure rise over time in a given volume:
QL | Leak rate | [mbar l s-1]
Δp | Pressure change during measurement period | [mbar]
V | Volume of the system | [l]
Δt | Measurement period | [s]
Leaks

## Slide 15
Gas flow rate is the volume of gas, at a known pressure, that passes through a plane per unit time
The throughput of a vacuum pump is the gas flow rate that a pump is able to absorb and is related to the pressure at the pump inlet:
This equation is directly obtained from the main ideal gas equation: PV=nRT
= ( )
PS
Q is define as the quantity of gas that leaves the pipe in the unit time.
Gas flow rate and throughput of a vacuum pump
P = Pressure of gas
V = Volume
n= number of molecules
R= Boltzman Constant (8,314472 J/mol K)
T= Temperature
P
dV
dt
Pump Inlet
PUMP
S is called pumping speed
The lower the pressure, «Better» or «Higher» is the vacuum
[mbar*l/s]
Notes: Q rappresenta la quantità di materia che deve essere portata via dal condotto nell'unità di tempo se si vuole mantenere la pressione bassa. Tale grandezza ha le dimensioni di una pressione per un volume diviso un tempo.
Con maggiore rigore la quantità di materia da portare via nell'unità di tempo dovrebbe essere il numero di molecole trascinate diviso il tempo in cui avviene il fenomeno. Quindi dall'equazione dei gas perfetti, dovrebbe essere pari alla derivata rispetto al tempo di pV/T. Se però le temperature del sistema da vuoto sono costanti nel tempo come spesso accade, risulta più semplice definire nella maniera anzidetta.
La grandezza dV, è la derivata temporale del volume che nel linguaggio del vuoto viene spesso denominata velocità di aspirazione, viene spesso indicata con il simbolo S (dall'inglese speed).

## Slide 17
Pumping Speed
(l/s, m3/h, cm3/s)
The pumping speed depends on the system pressure and gas type
All pumps have both high
and low applicable
pressure limit
Nitrogen
Argon
Is the volumetric flow (volume of gas per unit of time) through the pump’s intake port, measured at the pressure p. Determination of the pumping speed is described in base standard ISO 21360-1
[l/s]

## Slide 18
EXAMPLE

## Slide 19
When working in ultra-high vacuum, it can be important to know the composition of the residual gas.
A residual gas analyzer (RGA) is a small and usually rugged mass spectrometer(*), typically designed for process control and contamination monitoring in vacuum systems.
As example the percentages of water (m/e = 18) and its fragment OH (m/e = 17) will be large in the case of vacuum chambers that are not clean or well baked.
Leaks can be identified by the peaks of nitrogen (m/e = 28) and oxygen (m/e = 32) in the ratio N2/O2 of approx. 4 to 1. Hydrogen (m/e = 2), water (m/e = 17 and 18), carbon monoxide (m/e = 28) and carbon dioxide (m/e = 44) will be found in well-baked chambers.
No hydrocarbons have to be present in well cleaned chambers.
Typical residual gas spectrum of a vessel evacuated by a turbomolecular pump
Residual Gas Composition

## Slide 20
Section A
Pipe
Pump
P1
P2
Conductance is an abstract concept used to describe the behavior of gas in a vacuum system.
• Conductance is specific to a particular geometrical configuration.
• Conductance is specific to the actual gas species and temperature.
C= 1/Z=Q/P1-P2
Defines the pressure
drop in a pipe
Impedance
Z=( P1-P2 )/ Q
Gases moving through elements (pipes, tubes, vessels, and orifices) in a vacuum system encounter resistance to their motion. We can define the impedance of a tube as:
The Condutance is the capability to let through a particular gas volume in a known time
Vacuum Conductance
Notes: Prima di introdurla dare un esempio semplice. Manca la definizione di conduttanza o impedenza.
Consideriamo un recipiente ideale a pareti rigide di volume V
contenente N molecole di gas alla temperatura T ed alla pressione P
􀂄 Il recipiente è collegato ad una pompa da vuoto tramite un condotto di
sezione A (A è la sezione del foro che collega il recipiente al condotto).
􀂄 All’istante t nel recipiente ci sono N molecole alla pressione P
􀂄 All’istante t+dt nel recipiente ci sono N′ molecole alla pressione P′
La conduttanza è la capacità di un’apertura di lasciar fluire un determinato volume di gas in un’unità di tempo

## Slide 22
The conductance of pipes and pipe bends will differ in the various flow regimes.
In viscous flow they are proportional to the mean pressure p and in molecular flow they are independent of pressure. Knudsen flow represents a transition between the two types of flow, and the conductivities vary with the Knudsen number.
Conductance Properties
10-2 mbar
1 mbar

## Slide 23
Under molecular flow conditions, doubling the pipe diameter increases the conductance eight times. The conductance is INVERSELY related to the pipe length.
Example:
Conductance in Molecular Flow of a Long Round Tube

## Slide 24
Conductances
in Series
Conductances
in Parallel
A vacuum system can be analized/designed using an equivalent electrical model.
In this model:
1-pressures at a given point is the voltage
2-gas flow rate Q is the current
3-the conductances are electrical resistors
4-the pumping system are voltage generator
5- leacks as resistors that connect a given point to the mass.
Electrical Analogy of a Vacuum System (1/2)
I=ΔV/R= I=ΔV*G  Q=ΔP*C

## Slide 25
Pumping Technology
Primary pumping systems: Scroll Pump, Turbomolecular Pump,
UHV pumping system: Ion Pump, Getter Pump, Titanium Sublimation Pump
Vacuum Gauge and Leak Diagnostic

## Slide 26
Overview of vacuum pumps
Pumping Technology

## Slide 27
Primary pumping systems are mechanical pumps that work to decrease the pressure from atmospheric pressure to the pressure (10-6-10-8) to start the ion pump or other UHV pumping systems. We have:
Scroll Pump ( atmospheric pressure to about 10-3 mbar)
Turbomolecular Pump ( from 10-2 mbar to about 10-8 mbar )
UHV pumping system are the pumps that work at low pressure or in ultra high vacuum. The typical pumps are :
Ion Pump (from 10-6 mbar to 10-11 mbar)
Getter Pump
Titanium Sublimation Pump
In particles accelerator ion, Ti Sublimation and NEG pumps are in general used.
Different pumps are more effective for different chemical species.
In a vacuum system we find these typical gases:
Nitrogen N2, CO, CO2, methane , Argon , Oxygen , Hydrogen, Helium, Water ..
Pumping Technology

## Slide 28
Scroll Pump
A scroll compressor (also called spiral compressor, scroll pump and scroll vacuum pump) is a device for compressing air or refrigerant. It is used in air conditioning equipment, as an automobile supercharger (where it is known as a scroll-type supercharger) and as a vacuum pump.
A turbomolecular pump is a type of vacuum pump, used to obtain and maintain high vacuum.
These pumps work on the principle that gas molecules can be given momentum in a desired direction by repeated collision with a moving solid surface.
In a turbomolecular pump, a rapidly spinning fan rotor (50000-100000 rpm) 'hits' gas molecules from the inlet of the pump towards the exhaust in order to create or maintain a vacuum.
Turbomolecular Pump
Schematic of a turbomolecular pump.
Interior view of a turbomolecular
pump
A scroll compressor uses two interleaving spirals that allow to physically remove the gas from The system. It allows to reach pressure of the order 10-2-10-3 mbar .
These devices are known for operating more smoothly, quietly, and reliably than conventional compressors in some applications
Their design is similar to that of a turbine. A multi-stage, turbine-like rotor with bladed disks rotates in a housing.
These pumps can be a very versatile pump. It can operate from intermediate vacuum (~10−2 mbar) up to ultra-high vacuum levels (~10−8 mbar).
Notes: Most turbomolecular pumps employ multiple stages consisting of Rotor turbine / stator pairs mounted in series. Gas captured by the upper stages is pushed into the lower stages and successively compressed to the level of the fore-vacuum (backing pump) pressure. As the gas molecules enter through the inlet, the rotor, which has a number of angled blades, hits the molecules. Thus the mechanical energy of the blades is transferred to the gas molecules. With this newly acquired momentum, the gas molecules enter into the gas transfer holes in the stator. This leads them to the next stage where they again collide with the rotor surface, and this process is continued, finally leading them outwards through the exhaust.

## Slide 29
An ion pump (also referred to as a sputter ion pump) is a type of vacuum pump capable of reaching pressures as low as 10−11 mbar under ideal conditions. An ion pump ionizes gas within the vessel appling a strong electrical voltage, typically 3–7 kV, which allows the ions to accelerate and be captured by a solid electrode.
Ion Pump (1/2)
Ion pumps and are available in four types:
StarCell
Triode
Noble Diode
Diode

## Slide 30
The ion pump is composed of the following components:
a parallel array of short stainless steel tubes (anode)
two titanium plates (cathodes) which are spaced a short distance from the open ends of the tubes
and a strong magnetic field parallel to the tubes’ axes.
Ion Pump (2/2)
When inert gases are ionized, they are accelerated toward the cathode. Upon impact, they may penetrate several atomic layers and become trapped within the cathode lattice structure. They may also reflect as energetic neutrals and become embedded and trapped in the pump surfaces that see little or no sputtering such as the anode surfaces.
In summary, the pumping efficiency depends on the electron “cloud” density (which determines the number of ions produced) and on the sputtering yield (which determines the quantity of active getter material produced).
The electron cloud density mainly depends on the Penning cell geometry and on the electric and magnetic field
Strengths.
Ion pumps generally have very fast pumping speeds for reactive gases but poorer pumping for noble gases.
High voltage (<10kV) is applied between the anode and cathode. Electrons from the plates are generated by cold cathode emission and are accelerated toward the anode but are caused to oscillate along a helical trajectory in the anode space by the magnetic field, B.
Notes: The general pumping speed characteristics of sputter-ion pumps are similar to other vacuum pumps. There is a more-or-less constant-speed region at intermediate pressures, the net pumping speed is reduced to zero at low pressure, and it decreases to very low values at higher pressures. Due to practical engineering considerations sputter-ion pumps are not designed for high continuous-gas-throughput conditions. Compared to other high-vacuum pumps of equal pumping speed, sputter-ion pumps have approximately 100 times lower maximum throughput values. The pumping speed begins to decrease near 1×10-5 torr rather than near 1×10-3 torr
Sputter ion pumps operate by ionizing gas within a magnetically confined cold cathode discharge.
The events that combine to enable pumping of gases under vacuum are:
Entrapment of electrons in orbit by a magnetic field.
Ionization of gas by collision with electrons.
Sputtering of titanium by ion bombardment.
Titanium gettering of active gases.
Pumping of heavy noble gases by ion burial.
Diffusion of hydrogen and helium into titanium.
Dissociation of complex molecules into simple ones for pumping ease, e.g., CH4 breaks down into C and 2H2. Hydrogen is pumped separately. Carbon is no longer part of the residual gas and resides in solid form.
3 -7 kV range. 5 kV
No moving parts or oil: no maintenance or vibration

## Slide 31
Evaporable Getters
Non-Evaporable Getters - NEG
The active Ti surface is obtained under vacuum with subsequent depositions of a metal film of Ti in the system
Also in this case, a particular material alloy (NEG coating) has the property to absorb the molecules of gas. To activate the surface it is necessary to simply heat it at 250-300 C.
TSP
The titanium is heated to the sublimation temperature (about 1100 ° C).
The gas particles which collide on the layer of titanium are linked via chemi-absorption.
Main Getter Elements are:
Barium, zirconium, tantalum, molybdenum, vanadium, titanium, niobium
Support: plated iron or nickel, tickeness 0,2 mm
NEG Coating : i.e. 84% Zr; 16% Al or Zr, Fe, V
About 0,07 mm
Getter Pump

## Slide 32
Thermal Conductivity VacuumGauge
(Pirani)
Two platinum filaments are used as two arms of a Wheatstone bridge:
The filament in the reference tube is immersed in a gas at a fixed pressure in the high vacuum regime;
 The measurement filament is exposed to the vacuum system environment.
Both filaments are heated to a constant temperature by the current through the bridge.
Advantages:
Stable measurements within a wide temperature range
Highly resistant to overpressure
The resistance change define the vacuum pressure
This measurement principle utilizes the thermal conductivity of gases for the purpose of pressure measure in the range from 10-4mbar to atmospheric pressure.
When gas molecules in the vacuum system hit the filament, thermal energy is conducted away. This loss in thermal energy is detected and replaced by the feedback circuit to the power supply. The amount of electrical current needed to restore the temperature of the filament is then converted to a pressure readout.
Pirani gauges have inherent errors because the thermal conductivity and viscosity for each specific gas is different and varies non-linearly with pressure. They are therefore not used for measuring absolute pressures.
Notes: Scelta del vacuometro
Campo di pressione
Precisione necessaria
Condizioni dell’ambiente circostante
Compatibilita’ con altri strumenti
Integrazione/automazione richiesta
il Pirani misura la resistenza del filamento riscaldato per indicarne la temperatura, talvolta attraverso l’uso di un ponte Wheatstone calibrato.
Campo: da 1000 mbar a 10-3 mbar
Tempo di risposta: lento - 100 msec
Applicazioni: qualsiasi camera per basso vuoto
Precisione: +/- 10% a 10-3 mbar

## Slide 33
Penning gauge: more stable, less precise
cold-cathode gauge
2 electrodes: anode, cathode+ permanent magnetic field
invented 1937 by Penning
precursor of sputter-ion pump
nonlinear dependence
Bayard-Alpert gauge: less stable, more precise
hot-cathode gauge,
3 electrodes: filament, collector, grid
invented 1950, revolution in vacuum technology linear dependence
Here the pressure is measured through a gas discharge in the gauge head where the gas discharge is obtained by applying a high voltage.
The pressure range from 10-4 to 1 x 10-9 mbar.
Advantages:
is rugged enough
is resistant to sudden variations of pressure.
Low tendency for contamination (also during argon operation) due to high voltage reduction after ignition of the plasma and due to the titanium cathodes
Advantages:
The sensitivity of the device is more different for each gas;
is necessary to degas the head of measure to avoid outgassing
Vacuum Gauge based on Ionization Probability
Molecules are ionized and collected.
Pressure reading is determined by the electronics from the collector current.

## Slide 34
What is a Helium Mass Spectrometer Leak Detector?
It is a Helium-specific partial pressure analyzer
It detects Helium applied as a tracer or probe gas
It consists of:
the mass spectrometer tube tuned on He
it’s own vacuum system capable of 10-5 mbar in the spectrometer tube
a sensitive and stable amplifier valves, and auxiliary pumps for interfacing to vacuum system
a display for monitoring leak rate
Sensitivity is 10-10 mbar or better
A helium leak detector permits the localization of leaks and the quantitative determination of the leak rate, i.e. the gas flow through the leak. Such a leak detector is therefore a helium flow meter.
In practice the leak detector performs this task by firstly evacuating the part which is to be tested, so that gas from the outside may enter through an existing leak due to the pressure difference present. If there is a leak, helium can enter in the system from the leak (for example by using a spray gun). This helium flows into the leak detector and is detected.
Helium Leak Detectors
System
He
Notes: Principle of Detection
The leaked helium gas is ionized by the electron beam from the filament within the ion chamber of the analyzing tube. The ions are accelerated and move out through a slit and then pass through the magnetic field generated by the analyzer. Since the circular trajectories of the ions depend on their mass, the collector can catch only the helium ions and detect helium.

## Slide 35
Vacuum Materials
Clean Process
Laboratory Experience
VACUUM SYSTEM DESIGN

## Slide 36
Vacuum Materials

## Slide 38
STAINLESS STEEL
  Most common choice in HV and UHV systems
  304 – common,
  304L – Low carbon variant of 304 especially in UHV systems
  321 – for when low magnetic permeability is required
BUT…. Avoid 303 grade – contains sulphur and tends to outgas
COPPER (Oxygen-free) C10100 & C10200
  ‘Oxygen-free’ type is widely used
  Easy to machine
  Impermeable to hydrogen and helium
  Low sensitivity to water vapour
TUNGSTEN
  Can be used at high temperatures
  Can be used for filaments
BUT…Becomes brittle when work-hardened
ALUMINIUM & AL ALLOYS
  Low outgassing
  Easy to machine
  Low weight and lower cost
BUT…Some alloys contain a high proportion of Zinc ; Must NOT be anodised; Poor strength at high temperatures ;
Not easy to weld
Materials to use: Metals

## Slide 39
PORCELAIN AND ALUMINA
  Excellent electrical insulation
  Non-porous if fully vitrified
  Low coefficient of thermal expansion – usable to 1500oC
BOROSILICATE GLASS
  Used for viewports
  Can be machined and joined with metals
  Low coefficient of thermal expansion – resistant to thermal shock
Materials to use: Polymers
PTFE-TEFLON
  Good electrical insulator
  Tolerant to high temperatures
  Low outgassing
BUT…Cannot be used as a barrier between vacuum and atmosphere as it is permeable to gases
KAPTON
  Good electrical insulator
  Tolerant to high temperatures
  Very low outgassing
  Available in tape and film form
PEEK – Polyether ether ketone
  Excellent mechanical & chemical resistance
  Suitable for UHV applications
  Very low outgassing
BUT…Has a melting point of 343oC
VITON
  Used for demountable seals (‘O’ rings etc.)
  Can also be used as a seating face in valves
  Good electrical insulator
  Good chemical resistance
  Bakeable to 200oC
Materials to use: Ceramics

## Slide 40
CADMIUM
Often present in the form of plating (fasteners etc.) or in some brazing alloys
ZINC
Is a problem in high vacuum and high temperatures. Present in some alloys like brass (some electrical fittings)
MAGNESIUM
  Low melting point (650oC at atmosphere). Contains free hydrogen gas
PVC
  Often found in wire insulation, dust caps etc.
POLYMERS
  Many have an affinity to water
  Especially plastic tapes. Mould release residue can be an issue too. Polymers may generate a static charge attracting dust
  Nylon has a high outgassing rate
Materials To Avoid
because of the high vapor pressure

## Slide 41
To achieve pressures in the ultra-high vacuum range (<10-8 hPa) the following conditions must be met:
The base pressure of the vacuum pump should be a factor of 10 lower than the required ultimate pressure.
The materials used for the vacuum chamber and components must be optimized for minimum outgassing and have an appropriate surface finish grade.
Metallic seals (e. g. CF flange connections or Helicoflex seals for ISO flange standards) should be used.
Clean work is a must for ultra-high vacuum, i. e. all parts must be thoroughly cleaned before installation and must be installed with grease-free gloves.
The equipment and high vacuum pump must be baked out.
Leaks must be avoided and eliminated prior to activating the heater.
A helium leak detectors or mass spectrometer must be used for this purpose.
Bake-out significantly increases desorption and diffusion rates, and this produces significantly shorter pumping times. As one of the last steps in the manufacturing process, chambers for UHV use can be annealed at temperatures of up to 900 °C.
N.B: If stainless steel vessels with an appropriate surface finish grade and metal seals are used, bake-out temperatures of 120°C and heating times of approximately 48 hours are sufficient for advancing into the pressure range of 10-10 mbar.
Bake-out