## Slide 1
Experimental activity RF#2 - Pulse compressors
Why do we need pulse compressors?
Real devices widely used in electron linacs to increase the available peak power
High accelerating gradients require very high RF power ()
RF peak power available from klystrons is typically limited to ≈ 50-60 MW The duration of the klystron RF pulses ( few ) may largely exceed the typical filling time of an accelerating structure ()
Based on RF response of resonant cavities and other connection devices (hybrid couplers or circulators)
Also very instructive, suitable for an RF lab experience!

## Slide 2
The Idea: SLED (Stanford Linac Energy Doubler)
A factor 4 in power corresponds to a factor 2 in accelerating gradient, i.e. in beam energy!!
This idea has been implemented for the first time at SLAC in the 70’s [1]
Q: How can a large RF power flowing in a waveguide be manipulated to be compressed?
A: Compression is obtained capturing the RF power reflected by high-Q resonant cavities.
The energy of the RF pulse is compressed in about 1 filling time to increase the peak power to the maximum!
[1] Farkas et al. - SLED A method of doubling SLAC energy
4 µs
50 MW
Klystron typical RF pulse
Compressed pulse
200 MW

## Slide 3
The extra-power that flows out through the cavity couplers may significantly change the characteristics of the resonance. This effect is known as “cavity loading”. The loaded cavity Q-factor is lowered by the power coupled out through the cavity ports and results to be:
where is the usual quality factor of the resonant mode, related only to the dissipation on the cavity walls
Recap: coupling coefficients of a resonant cavity
«Real» resonant cavities are NOT perfectly closed volumes, at least 1 opening (port) must be present to allow the RF power to flow in (and out)
The coupling strength of a port can be measured as the amount of power extracted from the cavity through the port itself for a given level of the mode fields inside
This leads to the definitions of the coupling coefficient (β) of a port and external-Q (), in analogy with the definition of the resonance quality factor :

## Slide 4
Wave reflection in a guide terminated on a resonant load
A resonant cavity is a non-matched load for a waveguide or a transmission line, and can be modeled as an RLC parallel circuit.
An RF wave travelling from the generator toward the cavity is always partially reflected at the cavity input. The reflected wave can be calculated according to:

## Slide 5
Cavity reflected wave and input coupling coefficient
Lorentzian network
transfer function
Depending on the value of , 3 practical situations can occur:
under-coupling
critical coupling
over-coupling
Calculate and plot the Lorentzian network response when VIN :
Calculate and plot the cavity reflected wave when VFWD :
In the previous equation, if we substitute the expression for , the reflected wave becomes function of the coupling coefficient:
Step pulse
Rectangular pulse
Step pulse
Rectangular pulse
What are we going to do in the following:
Coupl. coeff.

## Slide 8
Lorentzian network response: RF step pulse
with
The response of a Lorentzian network to an RF step pulse is:
1
Tstart

## Slide 9
Lorentzian network response: RF rectangular pulse
with
The response of a Lorentzian network to a RF rectangular pulse is:
1
Tstart
Tend

## Slide 10
RF step pulse reflected by a resonant cavity
;
Under-coupled:
Critically coupled:
Over-coupled:
In time domain:
The reflected wave of a resonant cavity to an RF step pulse has the following transfer function:
180° phase inversion
1
Tstart

## Slide 11
RF pulse reflected by an over-coupled resonant cavity
The RFL peak field is maximum at the end of the FWD pulse (the negative step sums in phase with the reflected wave) -> its value is larger than the peak of the FWD itself.
The wave reflected by an over-coupled cavity excited by a rectangular RF pulse is:
1
0
Tp

## Slide 12
RF pulse reflected by an over-coupled resonant cavity
This effect can be enhanced by forcing a 180° phase jump in the driving pulse at a certain time (Ts) before the end of the pulse.
Thus we induce an equivalent step in the driving signal of double amplitude which adds in phase with the signal originated by the first rising front of the driving RF pulse.
1
0
Tp
Ts
Max theoretical amplitude gain = 3
-1

## Slide 13
Circulators are non-reciprocal (typically) 3-ports devices whose scattering matrix is ideally given by:
How to route the power flow towards the accelerator
90° hybrids are reciprocal 4-ports devices whose scattering matrix is ideally given by:
1
2
3
4
(and not towards the klystron…)

## Slide 14
How to route the power flow towards the accelerator
Reflected power from cavity 2:
No power emerging from port 3 S34=0
Half emerges from port 1 in phase S14= 1/√2
Half emerges from port 2 with 90° phase shift S24= j/√2
Reflected power from cavity 1:
No power emerges from port 4 S43=0
Half emerges from port 1 with another 90° phase shift S13= j/√2
Half emerges from port 2 in phase S23= 1/√2
Tot. power towards klystron (port 1):
Tot. power towards accel. structure (port 2):
-1/2
+ 1/2 = 0
j/2
+ j/2
(and not towards the klystrons…)
If we feed RF power from port 1 (klystron):
No power emerges from port 2 S21=0
Half power emerges from port 3 (90° phase shift) S31= j/√2
Half power emerges from port 4 (in phase) S41= 1/√2
RF source
(klystron)
Accelerating structure
RF cavity 1
RF cavity 2
1
2
3
4

## Slide 15: Experimental activity
Introduction:
  Measure environmental conditions and calculate SLED tuning frequency (working temperature Top = 40 °C, and nominal frequency fRF = 2856 MHz)
Time domain: set-up of the RF pulse
  CW signal at fLP from signal generator
  Stanford digital delay setup (5 us, 100 Hz, TTL HighZ)
  RF Switch setup
  Oscilloscope cross-check
Frequency domain: realization of the 180° phase jump
  Design of a 180° phase shifting circuit
  Functionality cross-check of the 180° phase shifter provided
  Actual phase shift measurement
  Generation and timing regulation of the trigger signals
  Oscilloscope cross check (30 dB att. 500 mV/div)
Time domain: SLED tuning with rectangular input pulse (no 180° phase jump)
  Cavity n.1 optimization
  Cavity n.2 optimization
  Tuning optimizing SLED output
Time domain: visualization of SLED output pulse with 180° phase jump
  Fine frequency tuning of the SLED pulse
  Qualitative measurement of energy/power gain
  Scan Tjump vs energy gain
Input port
90° hybrid
High Q cavities
Output port

## Slide 16: Hint: SLED tuning
WHY?
If the resonant frequencies of the 2 cavities are not the same and/or equal to the operating frequency (2856 MHz for S-band linacs) the SLED can produce high reflections at the input port, causing breakdowns in the input waveguide or near the klystron window.
HOW?
Generally frequency tuning is done in air and with low RF power. For this reason the measurement needs to account for several environmental variation with respect to nominal conditions: room temperature, pressure, humidity and air (instead of vacuum).
Where:
is the SLED nominal working frequency = 2856 MHz
is the SLED resonant frequency to be tuned at low power, room temperature and in air
is the frequency correction for temperature
is the frequency correction for pressure and humidity

## Slide 17: Hint: SLED tuning (finding fLP)
Where:
is the SLED nominal working frequency = 2856 MHz
is the SLED resonant frequency to be tuned at low power, room temperature and in air
is the frequency correction for temperature
is the frequency correction for pressure and humidity
EMPIRICAL FORMULAE (from SLAC & IHEP experience)
Where:
is the temperature (in °C) for high power operation;
is the room temperature (in °C) during low power measurement;
is the relative humidity (in %) during low power measurement.

## Slide 18: Hint: SLED tuning (single cavities)
Connect the output waveguide to a matched load;
Set-up the RF pulse:
  CW signal at fLP
  Pulsed signal at fLP (5 us length, 100 Hz rep. rate)
Display the reflection signal from each cavity on an oscilloscope (use a flat pulse without phase reversal). The connections should look like the scheme below:
Rotate the tuning screw to tune the resonant frequency of the cavity to have a minimum (dashed line). Repeat the procedure a couple of times for fine tuning.
Check the input VSWR: it should not be higher than 1.2
Oscilloscope Screen

## Slide 19: Hint: SLED tuning (sum of 2 cavities)
Once the cavities are individually tuned, the output signal of the SLED (without phase jump) should look like this:
Rotate the tuning screw to fine-tune the SLED output
If the cavities do not have the same response (different Q0 or different β), the optimum might not be where one expects it…
Oscilloscope screen