## Slide 1
Front-End Electronics Workshop 2023
Readout ASICs for photon science at LCLS-II-HE
Lorenzo Rota – TID-ID Integrated Circuits, SLAC lorenzor@slac.stanford.edu
Contributions from: C. Bakalis, S. Block, D. Doering, A. Gupta, A. Habib, J. Hasi, M. Hammer, C. Hansson, R. Herbst, C. Kenney, H. Kim, P. King, G. Liu, B. Markovic, F. Mele, A. Miceli, A. P. Perez, B. Reese, L. Ruckman, J. Segal, H. Shi, S. Strempfer, J. Weizeorick, A. Dragone

## Slide 2: LCLS-II & LCLS-II-HE: revolutionary tools for X-ray science
LCLS-II will be the first XFEL to be based on continuous-wave superconducting accelerator technology
Continuous repetition rate of 1 MHz, with photon energies between 250 eV and 12 keV

## Slide 3: LCLS-II & LCLS-II-HE: revolutionary tools for X-ray science
LCLS-II will be the first XFEL to be based on continuous-wave superconducting accelerator technology
Continuous repetition rate of 1 MHz, with photon energies between 250 eV and 12 keV

## Slide 4: SLAC long-term X-ray detector development plan
With goals built into projects progressively meeting science priorities and requirements
  Bigger, Faster, Higher resolution and Higher Energies
SparkPix (2020-2027)
ePixUHR (2019-2029)
Full frame 100kHz-1MHz
“General purpose” cameras
Efficient information extraction
(Advanced Calibration / Edge-ML)
ePixUHR35kHz (2019-2027)
Full frame 35kHz for HE first light
Sensors for hard X-rays
ePixHR5kHz (2016-2023)
5kHz for LCLS-II first light
ePixHR10k for tender Imaging
ePixHR250M for soft-ray Imaging
120Hz
5kHz
25kHz
100kHz-1MHz
1MHz-1GHz
ePix (2013-2018)
In use at LCLS
Revolutionary Experiment Specific X-ray Cameras
Information extraction at full rate (1MHz CW, 3GHz burst)
Advanced features (triggering, sparsification, zero suppression etc.)

## Slide 5: SLAC long-term X-ray detector development plan
With goals built into projects progressively meeting science priorities and requirements
  Bigger, Faster, Higher resolution and Higher Energies
SparkPix (2020-2027)
ePixUHR (2019-2029)
Full frame 100kHz-1MHz
“General pourpose” cameras
Efficient information extraction
(Advanced Calibration / Edge-ML)
ePixUHR35kHz (2019-2027)
Full frame 35kHz for HE first light
Sensors for hard X-rays
ePixHR5kHz (2016-2023)
5kHz for LCLS-II first light
ePixHR10k for tender Imaging
ePixHR250M for soft-ray Imaging
120Hz
5kHz
25kHz
100kHz-1MHz
1MHz-1GHz
ePix (2013-2018)
In use at LCLS
Revolutionary Experiment Specific X-ray Cameras
Information extraction at full rate (1MHz CW, 3GHz burst)
Advanced features (triggering, sparsification, zero suppression etc.)

## Slide 6: ePixUHR-35 kHz: overview
Summary of main specifications
Noise: <100 e-, single-photon resolution at 4 keV
Dynamic range: 104 @ 8keV
Size: 192x168 pixels with 100 µm pitch, ~2x2cm2
Frame-rate: 35 kHz (pixel and ADC meet designed for 100kHz operation)
Goals of 1st prototype:
Demonstrate path towards >25kHz detector on single ASIC on 130 nm
Retire risks associated with TSV and 3D stacking (ePixHR2.5D)
Evaluate performance at system level & at LCLS-II beamline at 35 kHz
Goals of 2nd prototype:
Optimize pixel front-end circuitry for optimal performance
Pixel matrix: 192 (H) x 168 (V) - Clusters
Readout
x16 LVDS outputs @ 1.1 Gb/s
Slow control & analog bias
ePixUHR ASIC
Readout
Readout
Si Sensor 100 µm pitch
x16
Serializers & clock spine

## Slide 7: Pixel matrix architecture
Operates at 100 kHz – 1 MHz
Si sensor: 100x100 µm2
ASIC: 50x100 µm2
Pixel
Cluster
600x1200 µm
72 pixels  1 ADC @ 8 MSPS
Digital logic for pixel configuration and readout
ADC
Digital
logic
Chart: 
| Pixels | ADC Ref | ADC | Digital
Consumption | 1.38 | 0.58 | 0.1 | 0.15
Power consumption of different blocks in matrix
(power density: 0.94 W/cm2)
Pixel analog front-end block diagram

## Slide 8: 12b Analog-to-Digital Converter
Differential Successive Approximation Register (SAR) architecture :
Optimized switching scheme to reduce both area and power consumption (on ADC local reference buffers)
Novel circuit to perform built-in single-ended to differential conversion
Based on custom-designed MOM capacitors with differential metal strips [Harpe, JSSC’19]
Total of 384 ADCs + local reference buffers operating synchronously in pixel matrix
180 um
260 um

## Slide 9: Moving data to the ASIC periphery
A/D conversion is synchronous in all clusters in pixel matrix
Data is shifted to cluster below
Matrix acts as a large shift register

## Slide 10: Results – prototype system
Pixel matrix test-pattern obtained at a clock frequency of 35 kfps
Prototyping carrier
Prototyping camera
First prototype readout architecture is fully functional up to the target frame-rate of 35kHz
ASIC has been assembled on prototype camera system with cooling block (no Si sensor)

## Slide 11: Results – ADC matrix
Preliminary
ADC: no missing codes in any of the 384 ADCs, linearity better than 0.2% (limited by test setup)
Hard to evaluate ADC performance because of limitations in ASIC/test setup:
  Can only test all ADCs simoultaneously
  Calibration voltage distributed across all matrix
  Noise <1.2 ADUs (limited by test setup)
  Non-linearity due to kick-back on calibration line?
Range ~650mV
1 ADU = 150uV
Noise on calibration line

## Slide 12: Results – pixel response
‘Correlated pre-charging’ circuit reduces noise after gain-switching [1]
‘CompEn’ feature = pixel is not allowed switch near the end of the integration phase
LG
MG
Preliminary
Noise map of pixels in High-Gain mode [ADUs]
Response of full-readout chain with pixel injection circuitry, pixel configured in auto-gain switching mode:

## Slide 13: Towards ePixUHR-100kHz
Matrix
100 kHz
Slow control
1 Gb/s I/Os
Matrix
100 kHz
Slow control
5 Gb/s I/Os
ePixHR
ePixUHR
ASIC
PCB / FPGA
ePixUHR 35kHz
Upgrade with 5 Gb/s links
Bottleneck
New readout
architecture
ePixUHR 100kHz
Same
Current design will be upgraded with 5 Gb/s links (CML/SST) (2024)
New FPGA readout architecture is currently being designed (2024)

## Slide 14: SLAC long-term X-ray detector development plan
With goals built into projects progressively meeting science priorities and requirements
  Bigger, Faster, Higher resolution and Higher Energies
SparkPix (2020-2027)
ePixUHR (2019-2029)
Full frame 100kHz-1MHz
“General pourpose” cameras
Efficient information extraction
(Advanced Calibration / Edge-ML)
ePixUHR35kHz (2019-2027)
Full frame 35kHz for HE first light
Sensors for hard X-rays
ePixHR5kHz (2016-2023)
5kHz for LCLS-II first light
ePixHR10k for tender Imaging
ePixHR250M for soft-ray Imaging
120Hz
5kHz
25kHz
100kHz-1MHz
1MHz-1GHz
ePix (2013-2018)
In use at LCLS
Revolutionary Experiment Specific X-ray Cameras
Information extraction at full rate (1MHz CW, 3GHz burst)
Advanced features (triggering, sparsification, zero suppression etc.)

## Slide 15: The science case for detectors: high-rate and versatility
“Rare Events”
e.g. Dynamic function of enzymes
Require 1 MHz detectors
Event Driven
Imaging non-identical objects
e.g. Heterogeneity of functioning nano-catalysts
Stochastic Dynamics
e.g. Glass transition,
charge fluctuations in quantum materials
XPCS, XSVS
Require 1 MHz detectors Sparse Data
Require full frame 100 kHz detectors

## Slide 16: SparkPix family
Four different pixel detectors in SparkPix family (currently):
I | Front-end | Information extraction | Frame-rate
SparkPix-ED | energy | triggering | 1 MHz / 100 kHz CW
SparkPix-RT | energy | data compression | 100 kHz
SparkPix-S | energy | sparse readout | 1 MHz
SparkPix-T | timing | sparse readout | 1 MHz  see tomorrow’s talk

## Slide 17: Energy + triggering
SparkPix-ED
vent
r
iven

## Slide 18: SparkPix-ED: science driver & concept
Rare events in X-ray scattering experiments
Science case: capture interesting events happening at random, stochastic times that make their observation difficult
Requirements: record N high-resolution images at closely spaced times around the rare event
  Operation:
  Low-resolution images are continuously sent to the EDGE computing layer @ 1 MHz Meanwhile, N images are continuously recorded & stored in the detector in a ring-buffer
  EDGE computing layer continuously analyzes data for rare events, then generated a trigger
  The trigger starts the readout of the high-res N images recorded around the event of interest
Detector
1 MHz
Stream
low-res
EDGE computing
Memories:
N frames hi-res
Trigger: read memories
Trigger generation
Data analysis
This talk
L. Rota et al., “SparkPix-ED: a readout ASIC with 1 MHz frame-rate for rare event experiments at LCLS-II”, IWORID 2021

## Slide 19: MHz mode
+
Fast read-out to triggering layer: 1 MHz stream
Sum signal of 9 pixels, resulting in a SuperPixel (SP) of 300*300 µm2
Sum operation reduces SNR by √9
High-resolution read-out
The ring-buffers in each pixel are read-out
Read-out rate: 1 MHz / (N * 9)
Read-out rate in 1st prototype: 25 kHz (N=4)
Can be run in “CW imager” mode: 100 kHz

## Slide 20: SparkPix-ED: overview of results from 1st prototype
SparkPix-ED ASIC with 48x48 prototype Si sensor
Spectrum of Cd109 source (preliminary, raw data).
Illuminated with red laser from right side
Pixel response with charge injection
Gain switching
Sum: 1 MHz
Hi-res: 0.1 MHz

## Slide 21: SparkPix-ED: fast trigger mode
ASIC
FPGA
DAQ
Trigger condition:
>20 pixels in low-gain
…
1 MHz stream
…
Inject at frame 8
Switch to
memories
readout
Triggered on frame 8
Frame 6
Frame 7
Frame 8
Frame 9
Fast images not stored
(data reduction)

## Slide 22: Energy + data compression
SparkPix-RT
eal
ime

## Slide 23: SparkPix-RT: motivation
  Next-gen photon science sources are facing a data deluge problem
  LCLS–II: MHz readout with on-line data reduction
  APS-U: Increases in brightness motivate the development of high-speed scanning microscopy instruments with rates approaching 1 MHz
  Collaboration between Argonne & SLAC detector groups:
  SLAC is developing ePixUHR & SparkPix family: front-end, FPGA, etc…
  ANL studying techniques for data compression on-ASIC
  Technical scope of SparkPix-RT project:
  Evaluate real-time compression in algorithms, in-ASIC
  Leverages modularity: front-end circuitry of ePixUHR/SparkPix-ED
Examples of techniques that will benefit form MHz rate detector. Top: diffraction pattern at LCLS Bottom: ptychography image at APS

## Slide 24: Compression algorithms
Initial Work – Compression performance on photon counting detector data
Lossless – Zero suppression
Pre-processing of data depends on the type of detector
Photon counting: information is already quantized
Charge integrating: pre-processing is needed before compression  pedestal correction, thresholding, gain correction
We are exploring techniques for on-ASIC, real-time data calibration

## Slide 25: Compression algorithms
Initial Work – Compression performance on photon counting detector data
Lossless
Lossy
On photon counting detector data, the lossless compression performs similar to more complicated compression methods such as PNG and JP2

## Slide 26: SparkPix-RT: status and next steps
  Digital logic implemented in periphery of ASIC
  1st prototype (MPW) is now in-fab, expected delivery in Aug 2023
  Working towards a 100kHz, full-scale (~2x2 cm2) demonstrator
5 mm
7.5 mm
SRAM
SRAM
SRAM
SRAM
SRAM
SRAM
SRAM
SRAM
Layout of (half) compressor digital logic

## Slide 27: Energy + sparse readout
SparkPix-S
parse

## Slide 28: SparkPix-S: science driver and detector needs
  X-ray Photon Correlation Spectroscopy (XPCS) and Speckle Visibility Spectroscopy (XSVS) will open new areas of science at LCLS-II:
    Study spontaneous fluctuations at the atomic scale from microseconds down to the femtosecond timescale
    Access stochastic fluctuations of matter on ultrafast timescales
  XPCS and XSVS experiments will benefit from a 2D integrating detector with:
    Fine spatial resolution
    Operating at 1 MHz, the max rate of LCLS-II
    Discriminating between 0, 1, 2, 3…. photons/pixel/frame
    On-line data reduction through sparsified readout
[1] W. Roseker et al., Towards ultrafast dynamics with split-pulse X-ray Photon Correlation Spectroscopy at Free Electron Laser sources
Signal of photon events in a small region in a speckle pattern [1]
  XPCS speckle pattern

## Slide 29: SparkPix-S: overview
  Block diagram of SparkPix-S prototype
  Goal is to meet the main requirements of XPCS/XSVS experiments in a single camera solution:
    High-spatial resolution: 50um pixel pitch
    High frame-rate: 1 MHz
    Low-noise performance < 60 e-
  Designed to be compatible with different sensors:
    4 - 12 keV  Si sensor
    0.25 - 2 keV  investigated in future phase of the project
  Moved away from “cluster” approach due to several reasons:
    sparse readout, small pixel pitch (area, IR drop, routing)
  Column-parallel ADCs in periphery (adapted from ePixUHR)
  Sparsification digital engine from SparkPix-T (adapted)
Pixel matrix & ADCs
Sparsification Engine
x4 LVDS output
@ 1 Gbps
Slow control & analog bias
Sparsification Engine
SparkPix-S ASIC
Sensor
(different variants for soft and hard X-rays)
x48 column SAR-ADC @ 10b

## Slide 30: SparkPix-S: pixel overview
Each pixel negotiates access to a bus shared every two columns.
10b SAR ADC (100um pitch) serving two columns.
Integration and readout are pipelined using 2 analog memories.
Filter is programmable: trade frame-rate for better noise performance.
Fixed gain, dual supply, synchronous reset CSA.
Threshold generation strategy based on a matched buffers with tunable offset + auto-zeroed comparator.
Asynchronous sparse logic

## Slide 31: SparkPix-S: status and next steps
  First MPW has been designed between July 2022 – May 2023
  Chip is now in fab, expected delivery in August 2023
  Simulated performance of first prototype:
    Noise: 50-55 e-
    Dynamic range: 50keV
  Full-scale design will feature x8 5Gb/s output links to meet the 1% occupancy spec @ 1 MHz
  Targeting tape-out of 1st full ASIC in mid 2024
5 mm
7 mm

## Slide 32
Summary

## Slide 33: Summary
Full-frame detectors:
ePixUHR-35kHz: first large-scale ASIC functional, characterization on-going
ePixUHR-100kHz: upgrade clocking & increase output throughput, 1st prototype in 2024
Detectors with on-ASIC information extraction:
SparkPix-T: 1st full-scale prototype is alive, characterization on-going, beam-time at TMO beamline
SparkPix-ED: 1st prototype demonstrated low-latency feedback, re-assessing science specifications
SparkPix-RT: 1st MPW in fab, currently scaling the design to full-scale for late 2023 submission
SparkPix-S: 1st MPW in fab, characterization will start in fall, full-scale ASIC in mid-late 2024
SLAC is developing a new family of ultra high-rate X-ray detectors to meet the science goals of LCLS-II / LCLS-II-HE

## Slide 34: Future research directions
Push computing as close as possible to the source
What is the path forward?
Distributed processing for efficiency (energy, information)
Leverage AI/ML for real time data extraction with low latency (see Ryan’s talk)
Develop adaptive data driven readout architectures
Develop supporting technology for high bandwidth on detector communication
Adopt beyond CMOS technologies with heterogenous integration

## Slide 35
Backup slides

## Slide 36: Methodology towards higher rates
Large number of detectors to be designed & commissioned in next years at LCLS-II
Strategy:
  Top-down approach: focus is on delivering detectors systems to meet LCLS-II science program
  Modularity, following the success story of the ePix family
  Sensors:
    Sensors compatible with different detectors (pitch, polarity, etc.)
  ASICs:
    CMOS 130 nm node selected: cost-effective, availability & good analog performance
    “Overhead” of IP blocks development & prototyping (communication, biasing, control, etc.) is spread over multiple projects / fabrication runs
  FPGA & readout cards:
    Common data protocols & slow-control interface
    Common readout cards, mechanics, cooling strategies etc..
    Firmware & software are modular (and open-source!)
https://github.com/slaclab

## Slide 37: 1 to 1,000 Hz

## Slide 38: ePix X-ray Detectors 1-120 Hz
Energy integrating hybrid
Common platform: off-chip ADCs, near identical balconies
Deployed at:
  Broad science impact: Photo-chemistry, HIV, inertial fusion, etc.
  LCLS and LCLS II
  EuXFEL, etc.
ePix100
  50 um pitch, 120 Hz
  rms noise ~45 e-, full well ~ 800 keV
EMP hardened
Vacuum compatible ~1e-7 Torr
250nm CMOS
ePix100
ePix100
Jungfrau
Characterizing the ionization potential depression in Dense Carbon Plasmas with High-Precision Spectrally Resolved X-ray Scattering (June 2019) PI: D. Kraus (HZDR) et al.

## Slide 39: ePix10K 120 – 1000 Hz
ePix10K
  100 um pitch, 1 kHz
  rms noise ~100 e-, full well ~80 MeV
  Auto-ranging per pixel per frame
Deployed at:
  Broad science impact: Photo-chemistry, HIV, inertial fusion, etc.
  LCLS and LCLS II
  Ultra-fast electron diffraction,
  Strong Field QED at FACET-II, APS, EuXFEL, etc.
Near-Physiological-Temperature Serial Femtosecond X-ray Crystallography Reveals Novel Conformations of SARS-CoV-2 Main Protease Active Site for Improved Drug Repurposing
https://www.biorxiv.org/content/10.1101/2020.09.09.287987v1.full.pdf

## Slide 40: 5 to 7 kHz

## Slide 41: ePixHR @ 5-7 kHz
X-ray Pump and X-ray Probe TXI instrument @ LCLS
Photon energies between 2 and 5 keV
Hybrid, synchronous, integrating detector
100 um pixels
Auto-ranging gain in pixel – up to 64 MeV/pix/frame
r.m.s noise < 100 e-
Move ADCs from board to the Read Out Chip
Thin entrance window sensors for tender x-rays – codeveloped with SINTEF
144
32
6
ΣΔ
ΣΔ
ΣΔ
DEC
DEC
DEC
ΣΔ
DEC
Temp
Row selector
Serializer
LVDS Output Buffers
CMD unit
SACI
Bias
Global Registers
Sigma-delta ADC per column

## Slide 42: ePixHR @ 5-7 kHz
Shingled for better fill factor
Moves 0.5m along beam and 2-jaw beam aperture
CO2 two-phase cooling system due to high areal power density
CO2 phase change cooling is the default technology for future LCLS high rate detectors

## Slide 43: ePixHR @ 5-7 kHz
First beam at LCLS of prototype
Used LCLS-II DAQ
Soft X-rays only
Photon beam pulse rate at 120 Hz (LCLS-I)
Operated Detector at ~ 5kHz
LCLS-II timing generating 4920 triggers/sec with 120Hz x-ray beam. 1 x-ray event every 41 events 5kfps
Soft FEL X-ray pulses

## Slide 44: ePixM – Monolithic / Hybrid Soft X-ray Imager @ 5kKz
Combined monolithic CIS and Hybridized to Back-End ASIC
Monolithic CMOS Image Sensor
50 um pitch
Auto-ranging with 2 gains
15 e- noise in High Gain
Full well ~ 500 keV
Multiplex 384 FE pixels to ADC input
LFoundry 150nm
Back-End ASIC
Array of 192 ADCs
4 ADC rows
5 kHz frame rate
X-rays Eγ = 250 eV – 2keV
Back-illuminated Monolithic CMOS Imaging Sensor
Back-end, ASIC with array of 192 ADCs
Solder bumps

## Slide 45: ePixM – Monolithic / Hybrid Soft X-ray Imager @ 5kKz
Solder Bump Array (IZM and SLAC in house)
Sparse array of signal bumps
Double column of power & control bumps along one edge for redundancy and improved yield
Encountered significant problems when initially bump bonding
Added dense array of mechanical bumps
Power/Control
Bumps
Mechanical
Bumps
ADC connecting bumps

## Slide 46: ePixM – Performance
Auto-ranging demonstrated
Gain switching seen with single 22 keV photons
Good QE
Science sensors fabricated and ready for IZM bonding
BE ADC ASICs fabricated
Quantum Efficiency
ALS beamline 6.3.2
Eric Gullikson
QE @ 277 eV ~84%
LG
HG
Electrical
Test Pulses
1.0
0.6
0.7
0.8
0.9
Back thinned to 115 microns thick
Implanted
Microwave annealed

## Slide 47: ePixM – Performance
15.7 e- noise in High Gain
Prototype monolithic CMOS sensor
Front illuminated
700um thick
Fe-55
Back illuminated
Fully depleted
Post thinning, implant, and microwave anneal
Fe-55
Photon Energy (eV)

## Slide 48: Cluster & pixel layout
Pixels
Pixels
ADC
ADC Refs & Decaps
Timing & Control
1200 um
Cluster
AP
PS
M8
M7
M6
ADC Cap
M5
Si sensor
Distributed ADCs and dig. logic in matrix  pixel circuitry displaced
Need to re-route from bump pad to pixel input
Can be done on-ASIC  do not need special Si sensor
Routing and shielding implemented in SparkPix-ED

## Slide 49: Cluster & pixel layout
Pixels
Pixels
ADC
ADC Refs & Decaps
Timing & Control
Noise across pixel in a cluster is “uniform”: no visible difference between pixels above ADC and pixel above “analog” front-end
Shielding + RDL of bump connections on ASIC demonstrated
Cluster Noise Map
Cluster Layout
[11,0]
[0,0]
[11,5]
[0,5]

## Slide 50: ePixUHR digital readout logic

## Slide 51: Clock distribution and serializer
Pad Frame
Serial
DigRO
Serial
DigRO
…
Serial
DigRO
Serial
DigRO
…
CLK_SER (from FPGA)
Data_out<8>
Data_out<15>
Data_out<0>
Data_out<7>
Digital balcony with datapath split into multiple identical blocks (DigRO)
Clock is generated by FPGA or external PLL
CMOS clock spine with duty-cycle re-generation circuit
Serializers generate local lock for each DigRO, skew is handled in FPGA
Fully functional up to >1.1 Gbps with forced air cooling

## Slide 52: ASIC-FPGA codesign
Having a behavioral model of the ASIC allows us to speed up the firmware/software design of the camera.
  Ease the testing and verification of the ASIC logic before the tape-out.
  Development of FW/SW and integration with DAQ can now run in parallel to the ASIC fabrication.
We target the emulation of the ASIC inside the FPGA (done for Fabulous, and under progress for ePixUHR).
Simulation
Emulation
EpixUhrTb.vhd
EpixUhr (Top Level)
U_Core
U_App
ePixUHR_ASIC_behav
Xilinx Kintex UltraScale
EpixUhrEmulation (Top Level)
U_Core
U_App
ePixUHR_ASIC_behav

## Slide 53: Methodology: co-design
App
Core
FPGA
Emulated
Data reduction on-ASIC  must ensure data quality and integrity after reconstruction
Industry-standard verification frameworks (UVM) are very powerful, but…
We developed an agile verification frame-work:
  Development of ASIC and FPGA digital logic is done in parallel (not sequentially)
  Simulate full ASIC + FPGA + Software stack
  Verification code = Deployment code
  Enable emulation of ASIC within FPGA
Additional benefits:
Integration with LCLS-II DAQ starts while ASIC in fab
Enables design space exploration of heterogeneous architectures: co-design of detector
Software / DAQ

## Slide 54: Methodology: co-design

## Slide 55: SparkPix-ED: specifications
Detector | SparkPix-ED | 
Mode | High-res mode | Fast-trigger mode
Pixel Pitch [um] | 100 | 300
Frame rate [kHz] | 100 | 1000
Matrix size | > 192 x 144 | > 64 x 48
Read Noise [e- rms] | 100 | 500
Well depth [4keV photons] | 104 | 104-105 [*]
Power consumption [W/cm2] | 1.2 | 
Data rate [Gb/s] | 44 | 50
CMOS tech node [nm] | 130 nm | 
Main design challenges:
Pixel operation from ~1 MHz
Maintain similar noise performance, linearity, dynamic range, power as ePixHR10kt (5 kfps)
Implement ADCs and digital logic distributed across pixel matrix, in the same die
[*] 105 assuming photons are equally distributed over 9 pixels

## Slide 56: SparkPix-ED: overall architecture
Operates at 35 kHz -100 kHz
Si sensor: 100x100 µm2
ASIC: 50x100 µm2
Pixel
Cluster
600x1200 µm
72 pixels  1 ADC @ 8 MSPS, 12b
Digital logic for pixel configuration and readout
ADC
Digital
logic
Balcony includes:
Analog biasing
Digital read-out logic
Slow control/configuration
x8 LVDS out @ 1 Gb/s
ASIC: 5x6 mm

## Slide 57: SparkPix-S: specifications
  Charge-integrating, synchronous front-end architecture with sparse readout engine
  Sparse readout needed: 1 MPix @ 1 MHz with 8b/pixel = 1 TB/s
Parameter | Hard X-ray | Soft X-Ray
Energy sensitivity | 4-12 keV | 250 eV-2 keV
QE at max energy | >70% @ 12keV | >70% @ 250eV
Occupancy (pix hit/pulse) | <1% | 
Frame-rate | 1 MHz | 
Pixel full well | 5 ph/pulse | 
Read-out noise | <60 e- (30e- long term) | <15 e-
Sensitive Area | >0.5 MPix - 40x40 mm2 | 
Pixel size | 50 x 50 µm2 | 
Data rate | ~40 Gb/s (full-size ASIC) | 

## Slide 58: ePixUHR ADCs: DAC testing line
ADC
Pixel
External DAC
16-bit DAC8812
FPGA
ePixUHR
DAC
There is a single DAC line (Vcal_p) that is fed to all the ADC following this distribution line.
AdcCalibEn