## Slide 1
CIF 07/06/2023
Gianmario Cesarini
PP1 Heaters

## Slide 2
PP1 Heaters

## Slide 4
W = 0.5 mm
T = 35 μm
Total length ~50 m
R = 48 Ω

## Slide 5
Heater with copper layer to homogenize the surface temperature and minimize the thermal contact resistance
TCu = 35 μm
Top
Bottom
* Thermal diffusivity of aromatic polyimide thin films by temperature wave analysis
Journal of Applied Physics 105, 113506 (2009)
Thermal Diffusivity (·10-6) [m2/s] | 
Cu | 111
AA6061 | 70
Kapton* | 0.12
FR4 | 0.14

## Slide 6
FR4
Copper Strips
NTC 10 kΩ
Top layer pad
Bottom layer pad
Short Circuit
SMD NTC 10kΩ
F – Feedback
I – Interlock
M – Monitoring
M
I
I
F
F
F
M
I
F
I
0.4 mm

## Slide 7
Calibration carried out in the small climatic chamber of the Electronic Service at three temperatures:
-40°C, 0 °C and 40 °C
R [Ω]
NTC 10 kΩ
PT100

## Slide 8
Climatic Chamber
Dry Air Inlet
Dry Air Exhaust
Insulating Material
Heater
Test Box
(Dew Point <-30 °C)
Climatic Chamber Dew Point -60 °C
Cold Dry Air -55 °C
Aluminum panel

## Slide 9
Dry Air Exhaust
Dry Air Inlet
Sensor cables
2 mm thick aluminum sheet

## Slide 10
NTC_1
NTC_2
NTC_3
NTC_4
NTC_5
Top Left
Top Right
Bottom Center
Box Wall
Box Air

## Slide 11
Armaflex
Dew Point sensor pipe
Top Left
Top Right
Bottom Center

## Slide 12
flow reduction in the box
zero flow in the box
Bottom resistor
Both resistors
Top resistor

## Slide 13
Top
Bottom
Both

## Slide 14
15/05/2023
14

## Slide 15
Time to reach -30 °C and saturation temperature when the climatic chamber temperature is -55 °C and the test box temperature is -50 °C
 | Top (min) |  | Bottom (min) |  | Both (min) | 
 | T = -30 °C | Tsat | T = -30 °C | Tsat | T = -30 °C | Tsat
Heater without tinned copper foil | --- | --- | 3 | 19 | 1 | 17
Heater with tinned copper foil | 3 | 8 | 3 | 14 | 1 | 19

## Slide 16
Test with IR camera to see the surface temperature distribution
test performed only on the heater without tinned copper foil
FR4 emissivity 0.9
Reflected temperature ~30 °C (reflector method)
Temperature measurement only via NTC sensors

## Slide 17
Top
Bottom
Top Resistor
(image obtained near the peak of saturation temperature)
Bottom Resistor
(image obtained near the peak of saturation temperature)

## Slide 18
(Only Bottom Resistor)

## Slide 19
Climatic Chamber

## Slide 20
Comparison of dew point sensors in the functional test
Dew Point Sensor Vaisala DMT143
Dew Point Sensor Rotronic Hygroflex 5

## Slide 21
Initial conditions:
Internal temperature of the climatic chamber 21.5 °C
Initial Dryer Dew Point -73 °C
Internal initial Dew Point Chamber 12.5 °C
External line pressure reducer 3 bar
Overpressure 1.8 mbar
Dew Point reading with chamber instrument
(External line pressure reducer 2.8 bar - Overpressure 1.7 mbar)
Dew Point [°C]
Sensore Dew Point Sensor Vaisala DMT143 Data
Time [h]
2
4
6
0
-10
-20
-30
-40
-50
-60
The dew point sensor of the climatic chamber differs considerably from the Vaisala (after 90 minutes the difference is already over 20 °C)
Autocalibrations
Drying of the climatic chamber environment and comparison between dew point sensors

## Slide 22
Initial conditions:
Chamber internal temperature 21 °C
Dryer Dew Point -74 °C
Internal Dew Point Chamber -20.8 °C
External line pressure reducer 2.8 bar
Overpressure 1.65 mbar (after one hour 1.88 mbar)
Dew point variation test with overpressure
Dew point variation test with refrigeration and subsequent heating
Dew Point [°C]
Overpressure
test
Temperature
test
Time [h]
2
4
6
-20
-30
-40
-50
-60

## Slide 23
Dew Point as a function of the overpressure by regulating the pressure reducer
(20/10/2022)
TCC = 20.4 °C |  | 
Pressure (bar) | Overpressure (mbar) | Dew Point (°C)
2,8 | 2,20 | -58.6
2.4 | 1.87 | -58.3
2.0 | 1.64 | -58.2
1.6 | 1.40 | -58.3
1,2 | 1.21 | -58.2
0.8 | 1.0 | -57.7
0.4 | 0.9 | -56.0
1 | 1.2 | -56.9
2 | 1.75 | -58.2
2.4 | 2.0 | -58.4
Overpressure reference:
First results from the upgrade of the Extreme Energy Events experiment, M. Abbrescia et al 2019 JINST 14 C08005
The dependence on the overpressure is approximately 2.5 °C from 0.9 mbar to 2.20 mbar (reducer pressure), but since the requirement of -60 °C is stringent, it is advisable to stay between 1 and 2 mbar (upper limit up to which work is considered safe of the climatic chamber)
Chart: Overpressure [mbar]
| 2.2000000000000002 | 1.87 | 1.64 | 1.4 | 1.21 | 1 | 0.9
 | -58.6 | -58.3 | -58.2 | -58.3 | -58.2 | -57.7 | -56

## Slide 24
Thermal Cycle for Type I Cable Test
(20/03/2023)