## Slide 1: Report misure TWPA (terza gen) Aprile 2023 @ LNF
Alessio Rettaroli X

## Slide 2
Cryostat setup
The chip

## Slide 3
A freddo si notano due zone di attenuazione attorno a 7 GHz e a 7.8 GHz. Inoltre su tutta la banda ci sono ripple di almeno 10 dB.
Zoom tra 6 e 9 GHz:
S21

## Slide 4
A freddo si notano due zone di attenuazione attorno a 7 GHz e a 7.8 GHz. Inoltre su tutta la banda ci sono ripple di almeno 10 dB.
S21
E. Enrico,

## Slide 5
Single tone spectroscopy vs Ibias
Ibias da -50 μA a +50 μA,
centro span VNA: 8 GHz
P_VNA = -30 dBm
Pump off
Period of about 46 A

## Slide 6
Single vs double tone spectroscopy
Ibias da 0 μA a +50 μA,
centro span VNA: 16.5 GHz
P_VNA = -20 dBm
Pump off
Ibias da 0 μA a +50 μA,
centro span VNA: 16.5 GHz
P_VNA = -20 dBm
fPump: 7.5 GHz
P_pump = -30 dBm
Pump off
Pump on

## Slide 7
Single vs double tone spectroscopy
Ibias da a ,
centro span VNA: 16.4 GHz
P_VNA = -20 dBm
fPump: 8.2 GHz
P_pump = -4 dBm

## Slide 8
Parametric amplification
Trovato un punto di amplificazione parametrica (circa 8 dB)
Ibias = 0
F_pump = 7.74727 GHz
P_pump = 4.5 dBm
F_vna = 7.6863 GHz
P_vna = -30 dBm

## Slide 9
Acquisiti signal e idler in funzione della frequenza di segnale (svolto da signal generator), con:
Ibias =
P_sig = -20 dBm
F_p = 8.2 GHz (svolta da VNA in CW)
P_p = -9.5 dBm
Signal + idler (4WM)

## Slide 10
Modulation with coil magnetic field
Acquisita modulazione degli rf-SQUID in funzione della bobina sul portacampioni
Nota: Dopo i 20 mA la temperatura del piatto a 10 mK comincia a salire rapidamente

## Slide 11
Conclusions
Seems that a stop band is right inside the signal region (7 GHz and 7.8 GHz)
Ripples: mismatches in our setup?
No particular bandwidth or signal gain observed
Modulation of 3WM with Ibias verified
Parametric amplification of single frequency points observed