## Slide 1: Development of a pixel detector for the European XFEL
Dr. Davide Mezza :: Chip Designer :: Paul Scherrer Institut
International School On Particle Detectors, Cogne 06/2023

## Slide 2: Outline
Introduction: the X-rays
X-ray detector: basic principle and components
The European XFEL (Eu.XFEL)
Detector requirements/How do we cope with them?
The AGIPD concept
How do we assess the performance of the ASIC?
Calibration
Overview of the final system (1Mpixel camera)

## Slide 3: Outline
Introduction: the X-rays
X-ray detector: basic principle and components
The European XFEL (Eu.XFEL)
Detector requirements/How do we cope with them?
The AGIPD concept
How do we assess the performance of the ASIC?
Calibration
Overview of the final system (1Mpixel camera)

## Slide 4
A brief introduction about light…

## Slide 5: Usage of light
Image credits: Delcreations, 123RF Ltd
VISION

## Slide 6: Usage of light
Medicine (surgery)
Read data from storage media
Transmit information
Many others…

## Slide 7: Usage of light
Medicine (surgery)
Read data from storage media
Transmit information
Many others…

## Slide 9
What are the applications?
(Why?)

## Slide 10: X-ray absorption
Air
no absorption
Lungs and soft tissues low absorption
Bones
high absorption

## Slide 12: X-rays diffraction
X-rays
Diffraction image

## Slide 13: X-rays applications
Reconstruction
Crystals
Reconstruction
Proteins

## Slide 14: Applications
  Physics
  Material science
  Environmental science
  Chemistry
  Biology
  Medicine
  …
Better understanding of
biological systems
Develop and engineer better drugs

## Slide 15
How do we produce X-rays?

## Slide 16: How do we generate X-Rays?
Three sources:
Synchrotron
X-ray Free Electron Laser
X-ray tube

## Slide 17: X-ray sources comparison
Light quality
www.lightsources.org
Low intensity
Unfocused
Continuous
Polychromatic
High intensity
Highly collimated
Quasi continuous (divided into bunches)
Partially coherent
Very high intensity
Highly collimated
Pulsed time structure
Very short pulses (10 fs)
Fully coherent
End 19th century
50’s ÷ Today
Since 2004

## Slide 18: X-ray sources
SLS @PSI (2001)
d = 138 m
SwissFel @PSI (2017)
l = 750 m
33 cm

## Slide 19: X-ray tubes in “every day” life
Hospital
Dentist
Airport

## Slide 20: Synchrotrons around the world (2018)
Image from: “The complexity of thermoelectric materials: why we need powerful and brilliant synchrotron radiation sources?”, X. Wu et al., Material Physics Today, Volume 6, 2018
> 50 worldwide

## Slide 21: FELs around the world (02/2023)

## Slide 22
How do we detect X-rays?

## Slide 23: The detector
The “EYE” of the machine

## Slide 24: Outline
Introduction: the X-rays
X-ray detector: basic principle and components
The European XFEL (Eu.XFEL)
Detector requirements/How do we cope with them?
The AGIPD concept
How do we assess the performance of the ASIC?
Calibration
Overview of the final system (1Mpixel camera)

## Slide 25
X-rays
Measurable quantity
Electromagnetic Radiation
Charge signal
(Analog)
Electrical signal
(Analog/Digital)
Sensor
X-ray (“light”) sensor
Readout Chip
ASIC
Backend/Real world
Data storage
Software
Data & Control <=> Digital signals
(Ethernet)
Detector System
Medium
(Air/Vacuum)
Bump-bonds
Wire-bonds
Network/Power Cables
Control/Readout Boards
Firmware
Software
FPGA
PCBs
Components
Sensor
Readout chip (ASIC)
Readout & control boards
Mechanics & cooling
Interconnections
Firmware
Software

## Slide 26
X-rays
Measurable quantity
Electromagnetic Radiation
Charge signal
(Analog)
Electrical signal
(Analog/Digital)
Sensor
X-ray (“light”) sensor
Readout Chip
ASIC
Backend/Real world
Data storage
Software
Data & Control <=> Digital signals
(Ethernet)
Detector System
Medium
(Air/Vacuum)
Bump-bonds
Wire-bonds
Network/Power Cables
Control/Readout Boards
Firmware
Software
FPGA
PCBs

## Slide 27: Hybrid pixel detectors
Bump Bonds
Connection between sensor and ASIC
Sensor HV
CSA
Signal treatment
ASIC
n+
p+

## Slide 28: Hybrid pixel detectors
SENSOR
Conversion of the radiation (X-rays) into and electrical signal (e--h+ pairs) by means of the photoelectric effect.
Electrons and holes are separated by means of the electric field and drift to the back side (e-) and to the readout anode (h+)

n+
p+
CSA
Signal treatment
Sensor HV
hν
E
ASIC

## Slide 29: Hybrid pixel detectors
n+
p+
CSA
Signal treatment
Sensor HV
E
ASIC
Readout ASIC
The “heart” of the detector:
Charge signal amplified
Signal treatment
  Amplification
  Filtering
  Digitization
Data sent to the outside world

## Slide 30
Single photons
Single Photon Counting
Noise highly suppressed
Readout of digital info
(Almost) unlimited dynamic
range
Discrimination of photons
below threshold
Minimum detectable energy
Count-rate limitation (pile-up)
Δtph > tshape
Δtph < tshape
Ph. separated
Ph. same time
Preamplifier
(charge sensitive)
Comparator
Preamplifier
(charge sensitive)
Comparator
Preamplifier
(charge sensitive)
Preamplifier
(charge sensitive)
Charge Integrating
Multiple photons
Energy for each photon
(Low flux, Polychromatic)
Measurement of charge sharing
(Energy reconstruction, Position
interpolation)
Practically no count-rate limitation
(High flux, Dynamic Gain, Monochromatic)
Integration of leakage current
Readout of analog information
Calibration procedure challenging
@FELs

## Slide 31: Take home messages
X-rays are useful and there are several ways to produce them.
Highest quality of X-rays are produced by FREE ELECTRON LASERS.
Page 31
Two basic working principle of ‘how to detect photons’:
Single photon counting;
Charge integrating.

## Slide 32: Outline
Introduction: the X-rays
X-ray detector: basic principle and components
The European XFEL (Eu.XFEL)
Detector requirements/How do we cope with them?
The AGIPD concept
How do we assess the performance of the ASIC?
Calibration
Overview of the final system (1Mpixel camera)

## Slide 33: European XFEL
Injector
LINAC (2 km, 17.5 GeV)
Undulators
Experimental
Hall
A 3.4 km long facility in Hamburg (DE)

## Slide 34: European XFEL

## Slide 35
First beam of the European XFEL (2017)

## Slide 36
EuXFEL
Great scientific opportunities
Big challenges for the detector system

## Slide 37: EuXFEL call
2006: the European X-ray Free Electron Laser launched a call to
develop new detectors capable to cope with the unique features
of this new machine…

## Slide 38: 3 consortia
LPD
Science & Technology
Facilities Council

## Slide 39: The AGIPD collaboration
Deutsches
Elektronen
Synchrotron
Put together different ideas and competences

## Slide 40: What’s special about the European XFEL?
109
1035
Brilliance
Time structure of the beam…

## Slide 41: What’s special about the European XFEL?
t
Eu.XFEL
PULSE
100 ms (10 Hz)
100 ms (10 Hz)
600 µs
99.4 ms
2700 ph. pulses
t
Up to 1012 X-ray ph. per pulse

## Slide 42: Outline
Introduction: the X-rays
X-ray detector: basic principle and components
The European XFEL (Eu.XFEL)
Detector requirements/How do we cope with them?
The AGIPD concept
How do we assess the performance of the ASIC?
Calibration
Overview of the final system (1Mpixel camera)

## Slide 43: Detector requirements
Eu.XFEL
Unprecedented peak brilliance
Pulse duration (< 100 fs)
Science
(Potentially) High number of photons impinging at the same time on the detector.
CHARGE INTEGRATING DEVICE
1 pulse every 222 ns
FRAME RATE
4.5 MHz
2700 ph. pulses
Readout of all pulses
@4.5 MHz not feasible
ON-CHIP STORAGE
OF THE FRAMES
Single photon resolution (@12.4 keV)
Up to 104 ph. per pixel (@12.4 keV)
HIGH (104)
DYNAMIC RANGE
Linear response
1%
(OVER ENTIRE DR)
DETECTOR

## Slide 44: Understanding the requirements
120 g
My pen
My car
± 10 g
Next obj: 120 g
x 10000
= 1200 kg !!!!
1200 kg
± 1 kg
This is NOT the entire system but
1 PIXEL

## Slide 45
Q·δ(t)
How do we make the readout electronics?
Ch. Amp.
Cf,pre
Vout = -
Vdd = 1.5 V
0V
Technology: IBM 130nm
ΔVout,max = 500 mV
Working point
(‘Baseline’)
Amplifier saturation
+
<Vout,n> ~ 1 mV rms
Linear range
Page 45
Sensor
Let’s look from the point of view of the minimum signal (1 x 12.4keV ph.)?
Our “scale” would break!!!

## Slide 46
Page 46
How do we make the readout electronics?
Q·δ(t)
Ch. Amp.
Cf,pre
Vout = -
Vdd = 1.5 V
0V
Technology: IBM 130nm
ΔVout,max = 500 mV
Working point
(‘Baseline’)
Amplifier saturation
+
<Vout,n> ~ 1 mV rms
Linear range
Sensor
How do we handle the maximum signal (104 x 12.4keV ph.)?
We would not see any signal… Our “scale” is not sensitive enough for small signals…

## Slide 47
Noise and gain can be optimized but we would never be able to ‘fill the gap’ of 2 order of magnitude… We have to introduce a new concept
Page 47
How do we make the readout electronics?
Q·δ(t)
Ch. Amp.
Cf,pre
Vout = -
Vdd = 1.5 V
0V
Technology: IBM 130nm
ΔVout,max = 500 mV
Working point
(‘Baseline’)
Amplifier saturation
+
<Vout,n> ~ 1 mV rms
Linear range
Sensor
Main limitation: only one parameter (Cf,pre) to play with…

## Slide 48: Dynamic gain switching
High Gain Stage
Switching Point: High  Medium
Medium Gain Stage
Low Gain Stage
Idea
Insert a mechanism that is able to recognize when the input of the charge amplifier is too ‘high’ for the given capacitor and it adapts dynamically the sensitivity by changing the value of the feedback capacitance Cf
Page 48

## Slide 49: Dynamic gain switching
Low Gain Stage
Key features
Single photon resolution in high gain stage
Dynamic range up to 104 x 12.4 keV ph. (Linearity <1%)
Poisson limited over whole dynamic range
Three gain stages
High: 1 … 80 x 12.4 keV ph.
Medium: 81 … 2000 x 12.4 keV ph.
Low: 2001 … <10000 x 12.4 keV ph.
Page 49

## Slide 50
The MINIMUM signal (charge) that we can detect is limited by the NOISE
The MAXIMUM signal (charge) that we can detect is limited by the SATURATION OF THE READOUT
Solving the problem of the high dynamic range with a single feedback capacitor CSA is not feasible
Page 50
NEW CONCEPT: the DYNAMIC GAIN SWITCHING
Take home messages

## Slide 51: Outline
Introduction: the X-rays
X-ray detector: basic principle and components
The European XFEL (Eu.XFEL)
Detector requirements/How do we cope with them?
The AGIPD concept
How do we assess the performance of the ASIC?
Importance of calibration
Overview of the final system (1Mpixel camera)

## Slide 52: AGIPD collaboration
Deutsches
Elektronen
Synchrotron
Adaptive Gain Integrating Pixel Detector
Readout ASIC

## Slide 53: Prototyping: a bit of history
01.2009
05.2010
11.2010
11.2011
10.2013
AGIPD 0.1
AGIPD 0.2
AGIPD 0.3
AGIPD 0.4
AGIPD 0.5
No pixels yet
Analog chain test
(Preamplifier + CDS)
Test of the gain switching
16 x 16 pixels
Full readout chain
(100 storage cells)
Different preamps and storage cell architectures
16 x 16 pixels
(200 storage cells)
On chip digital control interface
Radiation hard storage cell design
16 x 16 pixels
(352 storage cells)
DR improvement (smaller reset switches)
New off-chip buffer
16 x 16 pixels
(352 storage cells)
Precharge Scheme to enhance DR
AGIPD1.0 pixel structure included

## Slide 54: Full scale ASIC (11/2013)
AGIPD1.0 | 
ASIC technology | IBM 130 nm
Pixel | 64 x 64
Pixel size | 200 x 200 μm2
Storage cells | 352
Chip size | 13.1 x 14.8 mm2
Non-linearity | <1 %
Protection measure | Diode
Frame Rate | >4.5 MHz (burst)
Vetoing scheme | yes
AGIPD1.0
 |  | Standard Gain Mode | Low Noise Mode
Dynamic range |  | >1.104 x 12.4 keV | >4.103 x 12.4 keV
Noise (rms) |  | ~320 e- ENC | ~240 e- ENC
SNR @ 12.4 keV |  | >11 | >16
Min. energy (keV) |  | <6 keV | <4 keV

## Slide 55: Final ASIC Structure
3-fold dynamically gain switching preamplifier
CDS stage
(2 selectable gains)
Pixel buffer
(charge sensitive)
Column buffer
(interleaved,
precharging signal to bus)
Offchip driver
(Fully differential Amplifier)
Dynamic gain switching logic
Writing Section
Used during the
600µs bunch train
Reading Section
Used during the
gap of 99.4ms
Storage cells
(On-pixel frame storage )
Input protection diode
Page 55

## Slide 56: Outline
Introduction: the X-rays
X-ray detector: basic principle and components
The European XFEL (Eu.XFEL)
Detector requirements
The AGIPD concept
How do we assess the performance of the ASIC?
Calibration
Overview of the final system (1Mpixel camera)

## Slide 57: IR pulsed laser setup
λ = 1030 nm
Spot size <10 µm
Filter #1
Filter #2
Filter #3
Laser beam
AGIPD

## Slide 58: Dynamic range scan – IR pulsed laser
HIGH gain range (x12.4 keV): 65
Vth,comp
LOW gain range
(x12.4 keV): 8000
MED gain range
(x12.4 keV): 1200
Mind the log scale!
3

## Slide 59: Noise over dynamic range
ENC 300/320e- RMS
1 %
Shot-to-Shot fluctuation

## Slide 60: Outline
Introduction: the X-rays
X-ray detector: basic principle and components
The European XFEL (Eu.XFEL)
Detector requirements
The AGIPD concept
How do we assess the performance of the ASIC?
Calibration
Overview of the final system (1Mpixel camera)

## Slide 61: Calibration
Eu.XFEL main beam
Output (ADU)
Incoming ph.
Vth
Any pixel can be in any gain
Users need the number of photons impinging every pixel
To recover this information, we need 3 gains and 3 offset
Gain bit info
Image taken with the final 1M camera

## Slide 62
IDEAL CASE
1 pixel
Nominal Values
Eph=0
Page 62

## Slide 63
REAL CASE
1 pixel
Cpar,pre
~ 50 fF
Cpar,mem
~ 20 fF
Page 63

## Slide 64
The real value of the parasitics cannot be extracted with a 100% precision therefore the real value is unknown.
2. THE REAL VALUES OF BASELINE AND GAIN HAVE TO BE MEASURED EXPERIMENTALLY (FOR EVERY PIXEL)
(for this particular detector also for every storage cell)
The parasitics are different from pixel to pixel. The baseline and the sensitivity (gain) to incoming radiation are different from pixel to pixel.
1. EVERY PIXEL IS “UNIQUE”
(for this particular detector also every storage cell is unique)
Page 64
Take home messages

## Slide 65: “Raw data” image…
ADU… not energy!!

## Slide 66: Calibration
Fluorescence samples
X-Ray Tube
AGIPD
Fluorescence
photons
(Monochromatic)
In
Mo
Where:
qHG is the offset in ADU
men,HG is the gain in
X-ray tube photons
(Polychromatic)
0ph
1ph
2ph
3ph
The beam is monochromatic
0ph
1ph
2ph

## Slide 67: Gain distribution
Gain = (12.15 ± 0.3) ADU/keV (wrt the ideal 27.3!!!)
The gain (and offset) has to be evaluated for every pixel (and every storage cell)

## Slide 68: Dynamic range scan – Current source
Calibration Circuitry
Current Source: injection of a constant current into the v.g. of the preamp.
Iinj
(#ck)
Highly configurable
Requirements
The current has to be stable over the entire acquisition time
The ASIC (e.g. all the voltages) has to be stable over the entire acquisition time

## Slide 69: Calibration
In order to calibrate the final system:
  3 gains and 3 offsets per pixel;
  352 storage cells per pixel;
  1 Mpixel

## Slide 70: Corrected image
“RAW” IMAGE
ADU
CALIBRATED IMAGE
E (keV)

## Slide 71: Outline
Introduction: the X-rays
X-ray detector: basic principle and components
The European XFEL (Eu.XFEL)
Detector requirements
The AGIPD concept
How do we assess the performance of the ASIC?
Calibration
Overview of the final system (1Mpixel camera)

## Slide 72: AGIPD 1M System
AGIPD Module
Delivered to Eu.XFEL in July 2017
1 Mpixel detector system | 
Quadrant | 4 modules =256 kpixel
Overall pixel count | 1 Mpixel
Active area | ~22 x 22 cm2
Central hole | Adjustable
Heat dissipation | 1 kW / Mpixel system
Cooling | Liquid cooling (silicone oil)
Cooling power | 1.5 kW (-60 °C) 0.5 kW (-80°C)
Operating Temp. | -20 °C
Module system | 
Module | 8 x 2 chips =65 kpixel
Area | 102.4 x 25.6 mm2

## Slide 73: The final 1M system
1M system = 500 kg + 2.2 tons of copper cables

## Slide 74: Experimental results
Megapixels @ Megahertz The AGIPD High-Speed Cameras for the European XFEL
U. Trunk et al., Nucl. Instr. Meth. Phys. A 942, 162324 (2019).
Reconstruction of protein structures

## Slide 75: Our group @PSI
Group members (left to right)
Back: Jiaguo Zhang, Aldo Mozzanica, Thomas King, Davide Mezza, Anna Bergamaschi, Julian B. D. Heymes Middle: Erik L. Fröjdh, Carlos Lopez, Martin Brückner, Christian P. Ruder, Bernd Schmitt, Konstantinos Moustakas, Dominic Greiffenberg Front: Viktoria Hinger, Dhanya Thattil, Roberto Dinapoli, Shqipe Hasanaj, Maria Carulla, Simon Ebner
Missing: Rebecca Barten, Pawel Kozlowski, Filippo Baruffaldi, Kirsty A. Paton, Xiangyu Xie

## Slide 76: Thank you for your attention!
Contact:
davide.mezza@psi.ch

## Slide 77: The Paul Scherrer Institut (PSI)
Paul Scherrer Institut (PSI) is the largest research center for natural and engineering sciences in Switzerland (~2200 employees)
PSI is running Switzerland‘s large research facilities (neutrons, protons, muons, X-Rays):
 Synchrotron Swiss Light Source (SLS)
Free-Electron Laser SwissFEL
We develop detectors for the needs of the SLS and SwissFEL
Past detector developments:
PILATUS / MYTHEN (Commercialized by DECTRIS)
SLS
SwissFEL

## Slide 78: Single Photon Counting Detectors (since 1999)
 | MYTHEN | PILATUS | EIGER | MATTERHORN
Technology | UMC 110 nm | UMC 250 nm | UMC 250 nm | UMC 110 nm
Dimensionality | 1-D (Strip) | 2-D (Pixel) | 2-D (Pixel) | 2-D (Pixel)
Status | Modules available | Modules available | Modules available | Prototyping
Strip/Pixel size | 50 µm | 172 x 172 µm2 | 75 x 75 µm2 | 75 x 75 µm2
Maximum system size | 120º (=48 modules) | 6M (=42 x 43 cm2) | 9M (=23 x 23 cm2) | 
Minimum threshold | < 5 keV | < 2 keV | < 2.5 keV | 
Count rate capability | 0.6 MHz/Strip (10% deviation, Standard) | 0.5-1.0 MHz/Pixel (10% deviation) | 0.2-0.7 MHz/Pixel (10% deviation) | 
Maximum frame rate | 1 kHz/Module | 300 Hz/Module | 23 kHz (4-bit) | 
Applications (Examples) | Powder Diffraction Energy dispersive Spectrometer Beam Position Monitors | Protein Crystallography Time-resolved experiments Small and wide-angle X-Ray Scattering (SAXS/WAXS) | Protein Crystallography XPCS Coherent X-Ray Imaging Photoelectron detection | 
NEW!!!!

## Slide 79: Charge integrating detectors (since 2008)
 | GOTTHARD | AGIPD1 | JUNGFRAU | MÖNCH
Technology | UMC 110 nm | IBM 130 nm | UMC 110 nm | UMC 110 nm
Dimensionality | 1-D (Strip) | 2-D (Pixel) | 2-D (Pixel) | 2-D (Pixel)
Status | Modules available | Modules available | Modules available | (Advanced) Prototyping
Pixel size | 50/25 µm (Strips) | 200 x 200 µm2 | 75 x 75 µm2 | 25 x 25 µm2
Maximum system size | Modules (=10/20 ASICs) | 1Mpixel (=16 Modules) | 16Mpixel (=32 Modules) | Single Chip (=1x1 cm2)
Noise (r.m.s.) | <270 e- ENC | < 322 e- ENC < 214 e- ENC (HG) | < 100 e- ENC (G0) < 34 e- ENC (HG0) | <35 e- ENC
Dynamic range | < 1.104 x 12.4 keV (3 gain stages) | < 1.104 x 12.4 keV (3 gain stages) | < 1.104 x 12.4 keV (3 gain stages) | < 500 x 12.4 keV (2 gain stages)
Maximum frame rate | 600 kHz (cont.) 4.5 MHz (burst*) * 2720 frames | 4.5 MHz (burst*) * 352 frames | 2.4 kHz (cont.) < 1 MHz (burst*) *16 frames | 6-8 kHz (cont.)

## Slide 80: Life is not perfect…
AGIPD1.0: List of shortcomings @11/2014

## Slide 82: Non-monotonicity after gain switching
HG
MG
1st gain sw
Vdd,cds
1.5 V
1.5 V -
Performance Degradation
Power Networks improved in AGIPD1.1
AGIPD1.0
3-fold dynamically gain switching preamplifier
CDS stage
(2 selectable gains)
Pixel buffer
(charge sensitive)
Column buffer
(interleaved,
precharging signal to bus)
Offchip driver
(Fully differential Amplifier)
Dynamic gain switching logic
Input protection diode
Storage cells
(On-pixel frame storage )

## Slide 83: AGIPD1.0: Simulations
Power network
Behavior of the writing section for different Vdd,cds, Simulations in Cadence 5®

## Slide 84: After power network improvement
Non monothonicity after gain switching completely gone…!!
Vth,comp
Power network Improvements!
AGIPD1.1

## Slide 85: AGIPD1.0: ‘Ghosting’
Ghost
Ghost
Ghost
Laser

## Slide 86: AGIPD1.0: ‘Ghosting’
3-fold dynamically gain switching preamplifier
CDS stage
(2 selectable gains)
Pixel buffer
(charge sensitive)
Column buffer
(interleaved,
precharging signal to bus)
Offchip driver
(Fully differential Amplifier)
Dynamic gain switching logic
Input protection diode
Storage cells
(On-pixel frame storage )

## Slide 87: ‘Ghosting’ solution
100 Ω
(para)
Reason: Resistive network
100 Ω
(para)
100 Ω
(para)
(para)
(para)
(para)
(para)
Solution: Buffering of Vref,COB
(para)

## Slide 88: Facts
EuXFEL Call
2006
Ideas
2006/2007
Proposal
05/2007
Accepted
09/2007
Start
End 2007
Development
2007 - 2014
End
2014
End
2017
= 8.5 M€
= 10 M€

## Slide 89
Vout,i(t)
Gi
Vout,i-1(t)
Gi-1
Vout,i+1(t)
Gi+1
HV
Sensor
Cback
Cback
Cback
Ccoup,i+1
Ccoup,i-1
hν
Qtot
Qi
Qi-1
Qi+1
= Qtot
COUPLING
Notes: Starting from the coupling this is how it works… when we generate charge in the sensor, even if this is collected by one channel, because of the low dc gain of the preamplifier the virtual ground we induce signal through the coupling capacitors, in the neighbouring channels.
So whenever we want to calibrate our syst we have to take into account this effect otherwise the total charge will not be conserved.

## Slide 90: Pixel to pixel coupling
TOTAL COUPLING IN HIGH GAIN
J. Zhang et al., “Towards Gotthard-II: development of a silicon microstrip detector for the European X-ray Free-Electron Laser”

## Slide 91: Other approaches: the LPD & DSSC
Source
Clear gate
Clear
Gate
Drain