## Slide 1: Deep learning image denoising for a cost-effective walk-through pet design with sparse detector coverage
Maya Abi Akl, Florence Muller, Jens Maebe, Meysam Dadgar,
Othmane Bouhali & Stefaan Vandenberghe
MEDISIP, Department of Electronics and Information Systems, Ghent University, Belgium
Maya.AbiAkl@ugent.be

## Slide 2: Walk-Through PET concept Walk-Through PET system design deep learning model for image denoising Results conclusion and future work

## Slide 3: WT-PET concept: Vertical flat panels of monolithic detectors
106 cm
82 cm
50 cm
71 cm
106 cm
Long axial FOV
Short/lower-dose scans
High throughput (no bed – 30s scan)
Smaller footprint
(2 x 1.5 x 2.3 m3)
Lower cost (1.9x less detector surface) 3-4 MEuro
Superior spatial resolution + DOI (monolithic detectors)

## Slide 4: WT-PET: flat panels of monolithic detectors
12
20
1.14 mm FWHM spatial resolution + 6 virtual DOI layers (experimental results) [1]
LYSO 50 x 50 x 16 mm3

## Slide 5: Project timeline 2022-2025
Concept &
Early tests
Q3-Q4 2022
First prototype
Q1-Q2 2025
Simulations & Recon & Motion
Q1-Q4 2023
Detector & CT design
Q1-Q4 2024

## Slide 6: Walk-Through PET concept Walk-Through PET system design deep learning model for image denoising Results conclusion and future work

## Slide 7: System design: full and reduced sparse
57.5 cm
25 mm gap (1/2 detector)
Sparse design
Full Design
106 cm
71 cm
LYSO 50 x 50 x 16 mm3
20 rows 8 rows

## Slide 8: Sparse design motion to cover the afov
Panels upward motion
Stationary panels
Notes: AFOV motion: 1.62 cm/s
Gap motion: 0.125 cm/s (translating ½ detector module)

## Slide 9: Sensitivity comparison
110 cm line source inside a cylindrical water phantom of radius 10 cm
4.8x lower sensitivity
Full design
Sparse design with motion

## Slide 10: Walk-Through PET concept Walk-Through PET system design deep learning model for image denoising Results conclusion and future work

## Slide 11: Why deep learning denoising?
Reduced AFOV
Sparse detector coverage
Panel motion
Lower count statistics & more noise in the reconstructed images!
AIM: Images from the sparse design of similar quality as the full design

## Slide 12: Convolutional Neural Network: 4-layer 2D U-Net
64
3
64
128
128
256
256
512
512
256
256
128
128
64
64
Skip connection: Copy and Concatenate
1
1
⨁
Conv (3x3), ReLu
MaxPool2d (2)
ConvTrans2d (2), Pad
Conv (1x1)
ReLU
Residual connection: Add
Variable matrix size:
2D Image Slice
1
# trainable parameters: 12,260,353

## Slide 13: Training of the network
Model | Low Dose Challenge data (Quadra)
Train/Val/Test | 30/3/3
Target image | Full count: 6 & 10 min acquisitions
Input image | Count reduction factors: 2, 4, 10, 20

## Slide 14: Walk-Through PET concept Walk-Through PET system design deep learning model for image denoising Results conclusion and future work

## Slide 15: Low bmi patient
1/20
1/10
1/4
1/2
Full
Input image
DL denoised image
SUV

## Slide 16: High bmi patient
1/20
1/10
1/4
1/2
Full
Input image
DL denoised image
SUV

## Slide 17: Test XCAT – low bmi
60s simulation/reconstruction sparse design
180s simulation/reconstruction sparse design
To meet design scan requirement
To match lowest count level in the Quadra training data
Reconstructed using LM MLEM
TOF resolution 300 ps (LYSO)
kBq/mL

## Slide 18: Test XCAT – high bmi
60s simulation/reconstruction sparse design
180s simulation/reconstruction sparse design
kBq/mL

## Slide 19: Walk-Through PET concept Walk-Through PET system design deep learning model for image denoising Results conclusion and future work

## Slide 20: Conclusion
The first images obtained with the sparse WT-PET design show good image quality. Further improvement is possible by optimizing the motion of the panels.
The DL network performs well!
Main issues of the current denoising model(s):
The non-uniform counts statistics in the Walk-Through-PET XCAT images
The count level of the test XCAT data is lower than the lowest count level used for training
The distribution of the XCAT data is different than the distribution of the Quadra data

## Slide 21: Model: Trained on WT-PET XCAT data
Sparse design
Full design
DL denoised
kBq/mL
20 s simulation
40 s simulation

## Slide 22: future work
Train the DL model on realistic data but simulated with the WT-PET sparse/full designs
Optimize the motion of the panels for the sparse WT-PET design to get more statistics and reduce the noise in the brain and leg areas
Evaluate the effect of DL denoising on the spatial resolution of the denoised images
AIM: Use deep learning to achieve images of similar quality to the full WT-PET design with data acquired by the sparse design!
Reminder

## Slide 23: Look out for more presentations!
Performance characteristics of multi-mouse imaging on monolithic large flat panel PET
Energy based scatter correction for the Walk-Through PET system
Different Deep Learning Training Strategies for Attenuation and Scatter Correction in PET
Motion analysis of Subjects standing in walk-through total body PET using infrared based localization
Investigating the Influence of TOF and DOI on Spatial Resolution in Flat-Panel and Cylindrical Total-Body PET

## Slide 24: Thank you! THE Walk-through PET team
WT-PET.org

## Slide 25: Back-up slides

## Slide 26: CT component design
2 CNT source arrays
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
.
photon-counting
line detectors
WT-PET
rectangular CT
CT-meeting – August 2024
Bamberg, Germany

## Slide 27: Training details
Loss function: Mean Absolute Error
Optimiser: Adam
Learning rate: 1e-4
2D slices along three views

## Slide 28: Training of the network on xcat data
Model | XCAT simulated with the WT-PET design
Train/Val/Test | 9/2/2 BMI: 18 – 38
Target image | Full design (40s simulation)
Input image | Sparse design (20s simulation)