## Slide 1: On-Chip Analog Neural Network for In-Sensor Image Reconstruction Towards PET Scanners with Large Fields of View
Susanna Di Giacomo1,2, M. Ronchi1,2, M. Amadori 1,2, M. Carminati1,2, G. Borghi 1,2, P. Carra3,4, G. Sportelli3,4, N. Belcari3,4, C. Fiorini1,2
1Politecnico di Milano, Dipartimento di Elettronica, Informazione e Bioingegneria, Milano, Italy
2INFN Sezione di Milano, Milano, Italy
3Università di Pisa, Dipartimento di Fisica, Pisa, Italy
4INFN Sezione di Pisa, Pisa, Italy

## Slide 2: Outline
Introduction
  Motivations
  Neural network offline Training and Performance
Analog Neural Network ASIC (ANNA)
  Concept
  Architecture and Working Principle
ASIC Experimental Characterization
  Measurement Setup
  Chip-based Model for Training
  Results
Conclusions and Future Work

## Slide 3: Motivations
Motivations:
Analog operations directly on photodetectors analog signals.
No need for signal ADC and FPGA for embedded processing.
Reduction in number of interconnections and data to transmit.
Interaction coordinates directly at the ASIC output.
Objective: Localization of the radiation event from detector signals for nuclear medical imaging (e.g. Anger Camera for PET/SPECT).
Detection Module
Fig. from: S. Vandenberghe, et al. "State of the art in total body PET." EJNMMI physics 7 (2020): 1-33.

## Slide 4: Network Performance with Simulated Data
Network training in Matlab with simulated dataset (ANTS2) based on a 51mm 51mm 10mm LYSO monolithic scintillator readout by an 88 array of SiPMs.
64-20-20-2 Fully Connected architecture, 1762 parameters.
Quantization Aware Training (QAT): weights are quantized during training.
Performance tested with different number of bits
   5-bit best compromise between achievable resolution, chip complexity and occupied area.
LYSO
Scintillator
51x51x10 mm3
8x8 matrix
of 6.2x6.2 mm NUV-HD SiPM
511 keV Gamma rays

## Slide 5: ASIC concept and architecture
A
CF
S1
S2
S3
S4
S5
Ci,j
SRST
Vin,i
C
2C
4C
8C
4-bit weight
1-bit sign
Integrator characteristic as Clipped ReLU
Vout,j
Crossbar array of programmable switched capacitors.

## Slide 6: MAC Operation
1. Charging (multiplication)
2. Integration (accumulation + activation function)
weight
Clipped ReLU

## Slide 7: Setup for ASIC Characterization
Microcontroller Board
Carrier Board
Mother Board
X
Y
Layer 1
Layer 2
3
Digital (SPI)
Integrators
CMOS 0.35 µm, 24 mm2
Graphic interface:
ASIC programming
Inputs delivery
Data acquisition
Inner nodes

## Slide 8: Chip-based Model for Training
Network Training (Matlab)  Accurate model of the electrical neuron including non-idealities.
Model V2
Integrators offset
Injected charges
Parasitic capacitances
Model V1
Integrators offset
Build analytical model of the electrical neuron;
Measure all integrator offsets;
Model fitting with experimental data (neuron outputs) to estimate values of non-idealities (parasitic capacitances and injected charges). This is done only for Model V2.
Measured neurons outputs with different input/weights combinations
Neuron Model fitting
Training with Neuron Model

## Slide 9: Data Acquisition for Neuron Model Fitting
A dataset has been generated experimentally:
30 possible inputs
30 possible weight matrices
The model is based on experimental data measured from all 42 integrators (neurons).
We identified 6 free parameters (relative to parasitic capacitances and charges) to be fitted.
Analytical model of the electrical neuron
64 Analog Inputs
Weights
ANNA ASIC
µC
900 combinations x 42 integrators  37,800 neuron outputs.
Neuron Output
I/O
Inputs/weights generation
Data acquisition
Model fitting
Mother Board
Carrier Board
The measure is repeated for each neuron of the NN, and for each input/weights combination.

## Slide 10: Neuron Model Performance
Model V1
Model V2
Model outputs are compared with experimental outputs, for the same input/weights combinations.
Model V1 is unable to replicate electrical neuron behaviour.
Model V2 with optimized parameters matches more closely the electrical neuron.

## Slide 11: ASIC Performance Characterization
Training with the electrical model of the neuron.
Grid of 121 simulated points (ANTS2) is used to test ASIC performance.
8x8 SiPMs signal [V]
x [mm]
y [mm]
64
The two output voltages are the predicted X, Y.
The steady-state values are converted in mm values within the range ± 25 mm.

## Slide 12: Accuracy Measurement
Neural Network | Bias [mm]* | 
Matlab 5-bit |  | 1.7
ASIC – Model V1 |  | 11.0
ASIC – Model V2 |  | 3.9
* Average bias among all 121 points in the grid.
Bias = displacement between position reconstructed by the network and the real event position.
Test points
ASIC pred. – Model V1
ASIC pred. – Model V2
*
Figure: Example of a subset of grid points along diagonal reconstructed by the ASIC NN, comparing the two models used for training.
The average bias for these 9 points is 2.3 mm using Model V2.
It is fundamental, during training, to use a model that can accurately reproduce chip behaviour.
Notes: ERORRE SISTEMATICO DELLA RETE SUL CHIP SUI PUNTI

## Slide 13: Precision Measurement
100 inferences of the same input.
Std. dev. of predictions (X ± σx, Y ± σy) to evaluate effect of noise on ASIC NN positioning precision.
The predicted X, Y have standard deviation σx = 0.074 mm and σy = 0.067 mm.
The fluctuations produced by noise can be considered negligible if we target a resolution of ≈1.5 mm FWHM.
σx = 0.074 mm
σy = 0.067 mm
2D FWHM = 0.166 mm

## Slide 14: Optimization Study of NN Architecture
Objective: study of spatial and timing resolution of monolithic crystals (LYSO and LaBr3), optimizing neural network dimension compatible with ASIC implementation for a new chip version.
Neural network training
  Min: 500 parameters, 2 layers
  Max: 5000 parameters, 6 layers
  100,000 events across 2-mm spaced points
Geant4 simulation of LaBr3 and LYSO:
  50mm x 50mm x 12/16mm
  25mm x 25mm x 12/16mm
  6mm x 6mm and 3mm x 3mm Hamamatsu SiPM
Geometry | Neural Network | Res. (X,Y) [mm] | Res. Z [mm] | Res. T [ps]
50mm/16mm + 6mm SiPM | 2 layers 500 params | 1.91 | 3.50 | 580
50mm/16mm + 6mm SiPM | 4 layers 2000 params | 1.20 | 2.30 | 345
50mm/16mm + 6mm SiPM | 6 layers 5000 params | 0.75 | 1.84 | 287
25mm/12mm + 3mm SiPM | 6 layers 5000 params | 0.65 | 1.95 | 92
LaBr3 results
Geometry | Neural Network | Res. (X,Y) [mm] | Res. Z [mm] | Res. T [ps]
50mm/16mm + 6mm SiPM | 2 layers 500 params | 1.95 | 3.61 | 612
50mm/16mm + 6mm SiPM | 4 layers 2000 params | 1.25 | 2.39 | 383
50mm/16mm + 6mm SiPM | 6 layers 5000 params | 0.82 | 1.99 | 301
25mm/12mm + 3mm SiPM | 6 layers 5000 params | 0.77 | 2.11 | 122
LYSO results

## Slide 15: Ongoing and Future Work
Start the design of a new ASIC version in a more scaled technology (CMOS 110 nm):
  Revision of the crossbar structure and integrator.
  Avoid non-idealities found in the first version.
  More flexibility.
  Recurrent Neural Network (programmable number of inputs, outputs, layers).
  Improve speed, efficiency.
 4 layers/2000 parameters could be a good compromise for the next ASIC version.
Di Giacomo, S., et al. ”Implementing an Integrated Neural Network for Real-Time Position Reconstruction in Emission Tomography With Monolithic Scintillators.” IEEE Trans. Rad. Pl. Med. Sci., vol. 8, no. 5, pp. 501-510, May 2024.
Publication:
Figure of Merit | This Work @10 MHz
Technology | 350 nm
Area [mm2] | 24
I/O Latency [us] | 4.5
Energy per Inference [nJ] | 76
Energy Efficiency [GOP/J] | 46.8

## Slide 16
Thank you!
susanna.digiacomo@polimi.it