## Slide 1: Qweak Target Performance
Greg Smith, JLab
September 7, 2011 PAVI 11, Rome
Collaborators: S. Covrig, D. Meekins, J. Dunne, A. Subedi
At PAVI 10, discussed conceptual design of a high power, low noise PVES tgt
At PAVI 11, we present msrd performance of this tgt
  As built design, novel features
    Problems experienced
  Measured cooling power
  Measured target noise (boiling)
  Implications for next generation (MOLLER tgt)
Design document:
  http://qweak.jlab.org/doc-public/ShowDocument?docid=1041

## Slide 2: Design by CFD
beam
Heater
Cell
Heat Exchanger
H2 Release/
Safety
Dummy/Bkg Tgts
CFD calculations
by S. Covrig (Jlab)
Raster
Windows

## Slide 3: LH2 Target Specifications
T = 20 ± 0.01 K
  (TF = 13.8 K, TB = 23.6 K)
P~30-35 psia (20 psig)
Pbeam ~ 2.1 kW @ 180 µA
Ptotal ~ 2.8 kW @ 180 µA
  Reserve power ~ 150 W
Fan  = 30 Hz
3x3 ≤ raster ≤ 5x5 mm2
2 axis motion
Noise < 50 ppm
LH2 (35 cm cell):
Heaters must provide Pbeam + Preserve ~ 2.3 kW (@ 180 µA)
Coolant must provide Pbeam + Preserve + Ppump + Pother ~ 2.8 kW
Coolant mass flow (g/s)
Cooling Power (W) for a 20 K Target
4.5 K
15 K

## Slide 4: Hybrid 3 kW HX
Uses 15 K from ESR (has its own 20 K return)
Uses 4 K from CHL (20 K return on xfer line shield)
Danger: 4 K << 14 K. H2 freezes at 14 K!
    4 K supply JT valve PID on loop temperature
    4 K supply heater at HX inlet (500 W), on PID
    4 K supply bypass JT valve to the 4 K return
½” Cu fin tube,
16 fpi.
87.4 cm

## Slide 5: Freezing the HX is all too easy...
4K coolant T out of HX starts to drop
LH2 temp suddenly rises as HX ices up
4K coolant T out of HX suddenly drops well below freezing
TO throttles 4K JT, causing coolant supply temp to rise, defrosting HX
14 K = Tfreeze(H2)
Temperature (K)
TIME
(only a problem during cool-down)

## Slide 6: Cell
LH2 Transverse Flow:
1.1 kg/s
15 L/s
3-7 m/s
Electron Beam
180 μA, 4x4 mm2
Cell 35 cm long
Al entrance window ~0.1 mm thick, 22.2 mm Φ
Al exit window ~0.125 mm thick over 15 mm Φ,
0.635 mm thick over 173.5 mm
Scattered electron acceptance ±13.9°
13.9°
13.9°

## Slide 7: LH2 Cell & Windows (after 6 months running)
12”
7.5”
flow
beam
15 mm Φ, 0.127 mm thick
exit window vacuum side
exit window, LH2 side
entrance window 0.097 mm thick
exit window vacuum side
1”

## Slide 8: 1 hp Centrifugal Pump
Early bearing failure.
Mitigated by promoting H2 flow past bearing & across the pump.
Volute
Rotor
Stator
Piece of bearing retainer
All-ceramic bearing failed
Impeller
Motor housing cooled with 4K return

## Slide 9: Signs of pump wear (after 6 months)
Original bearings: ceramic balls & race, teflon retainer (failed)
New bearings: ceramic balls, SS race, with vespel retainer (graphite impregnated)
Retainer still wears:
  big gaps between balls
  visible debris
Rotor will be replaced
New bearings will be press fit onto a precision turned rotor shaft (outside)
scoring

## Slide 10: Target centering strategy works well
LH2 Exit window with thin central nipple
Aluminum background targets
2 mm
Dynamic 2 axis motion system (x,y).
Manual cell adjustment mechanism (pre-cooldown) allows adjustment of z, roll, pitch & yaw.
15 mm

## Slide 11: Lessons from First Test Cooldown
Added:
  Manual 4K bypass valve to pre-cool xfer lines
  4K PID to automatically close 4K supply if loop temp drops (deadman switch)
  4K heater on PID (500 W) to quickly & automatically reduce 4K cooling power
  Low power heater (50 W) to keep motion system at 300K (counteracts cold leaks)
Problems:
  Pump at top of vertical loop is hard to prime
  Use of 4K coolant makes it easy to freeze LH2 during cooldown (stable once liquified)
  Had to improve pump bearing cooling after they seized
  Pump motor heat load larger than expected

## Slide 12: Qweak Tgt Performance: Design vs Reality
Design mass flow 1.1 kg/s
  Measured mass flow 1.1 ± 0.1 kg/s. GOOD!
Design head 1.2 psi (@ 15 l/s = 1.1 kg/s)
  Measured head 1.1 psi. GOOD!
Heater: design goal 3 kW
  Measured up to 2.4 kW (didn’t need more) GOOD!
HX: design goal 3 kW
  Measured up to 2.9 kW so far! GOOD!
Target Noise: design goal < 50 ppm
  Measured 46 ppm at 180 µA, 4x4 mm2 raster, 28.5 Hz pump, 960 Hz helicity reversal GOOD!

## Slide 13: Coolant Scheme
Hybrid scheme working beyond expectations:
  Otherwise wasted enthalpy of returning 4K used to boost 15K plant’s performance!
    > 2 !
(4K)
(15K)
Xfer Line
15K In
4K In
15K Out
4K Out
  Novel use of existing xfer line infrastructure works!
    4K returned in LN2 supply (shield) space

## Slide 14: Cooling Power: Great!
4K ΔT = 15°
15 K ΔT = 4.9°
Coolant Temp (K)
0
10
20
15
5
Time
Time
15K flow:
26 g/s to Hall C
13 g/s to Hall A
4 K flow:
20 g/s to Hall C
0
10
20
m (g/s)
30
25
Qtot=880 W (15K) + 2050 W (4K) = 2930 W
Heater 2515 W. 415 W “missing”, but
Viscous heating = 6.9 * 17.7 l/s * 1.1 psi / ε
= 270 W @ 50% effi
•

## Slide 15: MD Yield (time)
20 µA
150 µA
LH2 @ 3x3 mm2 & 30 Hz
Target noise
(not near helicity reversal frequency,
so it’s not a problem)
FFT
FFT: Tgt noise/bubbles at ν < 40 Hz.
Fast helicity reversal avoids tgt noise!
~
~1%

## Slide 16: Target Noise (aka boiling)
Extracted in quadrature from msr’d main detector asymmetry widths by varying:
  LH2 pump speed (cleanest: only tgt noise changes)
  Raster area (bkg can potentially also change)
  Beam current (must account for changing statistical widths, BCM, etc.)
Repeated these for different conditions
  Results are consistent within a few ppm

## Slide 17: Standard Target Performance
DA, RS, BM Boiling tech note (2003)
Standard pivot targets
Standard
pivot tgt
(NOT Qweak!)
Pure Statistics
Tgt Boiling!
Msr’d width

## Slide 18: Beam Current Scan
4 parameter fit:
  Stats ~ 1/√Ibeam
  BCM ~ 1/Ibeam
  Tgt Noise ~ 1/I3beam
  Other ~ constant
Msrmnts obtained with relatively small raster, before shielding improvements
46 ppm
  (@ 169 µA, scaled to 4x4 mm2 raster)
No boiling
Tgt Boiling!

## Slide 19: Pump Scans
3 parameter fit
Msrmnts obtained after many bkg improvements (cleaner)
42 ppm (@169 µA, 4x4 mm2 raster, 28.5 Hz)
Exponent fixed at 2
Hypothesis: flow velocity falls off outside 5 mm vertically (flow is horizontal). Consistent with FLUENT guidance.
____ assumes 1/f
Exp. 2
XxY raster,
flow in X

## Slide 20: 182 μA Raster Scan
226 ppm floor
Only 46 ppm at
182 µA & canonical
4x4 mm2 raster!
Expected shape based on lower current behaviour

## Slide 21: Explicitly Measured Dependence on Helicity Reversal Frequency
Scaled 480 Hz MD results by √2 to remove trivial statistics effect
Tgt noise is worse at 480 Hz, but not by as much as simulation indicated
Do we still win at 480 Hz anyway?

## Slide 22: Comparison of Different Flip Rates
480 Hz Reversal:
Slower reversal:
  improved livetime
  better BCM width
  worse boiling
    but this term makes a larger relative contribtion
  Net loss (slight: 249 ppm vs 238 ppm)
    need 4.5 x 4.5 mm2 raster to break even
960 Hz Reversal:
Follows ν(-0.4) power law determined from early tests!

## Slide 23: Moller Target Specifications
150 cm LH2 (17.5% X0) at 20K, 35 psia
5x5 mm2 raster area
85 µA beam current
Total cooling power required 5 kW
2 kHz helicity reversal frequency
Target noise contribution to asymmetry width ΔA ~ 26 ppm
< ~ 5% contribution to ΔA
Minimize window bkg

## Slide 24: Film Boiling @ Windows
MOLLER looks promising: careful design may eliminate film boiling @ windows!
Convective part
Predicted by CFD
Total Heat Flux
(dE/dx) / Araster
Threshold for film boiling

## Slide 25: Scaling Qweak Target performance
Moller target needs tgt noise < 25 ppm (85 μA, 5x5 mm2, 2 kHz)
Msr’d Qweak tgt noise ~ 7 ppm at 85 μA, 3.5x3.5 mm2 raster.
With a 5x5 mm2 raster, expect only 3.4 ppm !
But Moller target ~ 4 times longer. If scaling with length ~ scaling with current, expect ~27 ppm for Moller target.
  But this is way too harsh! Scaling should be much less because window boiling (dominant effect) does not increase with tgt length!
    Moller also plans twice the helicity reversal frequency
Qweak
Conclusion: Scaling Qweak performance, it should be possible to more than meet requirements for MOLLER target.

## Slide 26: Summary
Qweak target is world’s highest power, lowest noise PVES target:
  46 ppm @ 182 μA, 4x4 mm2 raster
Design goals were achieved
  Validates CFD as a design tool for the future
Operated successfully for 6 months
  Late start. Initial pump bearing failure.
  Six months to go, after planned refurbishment
Successful prototype for next gen tgts
  MOLLER

## Slide 27
The Qweak Collaboration
A. Almasalha, D. Androic, D. Armstrong, A. Asaturyan, T. Averett, J. Balewski, R. Beminiwattha, F. Benmokhtar, J. Benesch, J. Birchall, C. Capuano, R. D. Carlini1 (Principal Investigator), G. Cates, J. Cornejo, S. Covrig, M. Dalton, C. A. Davis, W. Deconinck, J. Diefenbach, K. Dow, J. Dowd, J. Dunne, D. Dutta, R. Ent, J. Erler, W. Falk, J.M. Finn1*, T. A. Forest, M. Furic, D. Gaskell, M. Gericke, J. Grames, K. Grimm, D. Higinbotham, M. Holtrop, J.R. Hoskins, E. Ihloff, K. Johnston, D. Jones, M. Jones, R. Jones, K. Joo, J. Kelsey, C. Keppel, M. Khol, P. King, E. Korkmaz, S. Kowalski1, J. Leacock, J.P. Leckey, A. Lee, J.-H. Lee, L. Lee, N. Luwani, S. MacEwan, D. Mack, J. Magee, J. Mammei, J. Martin, M. McHugh, D. Meekins, J. Mei, R. Michaels, A. Micherdzinska, A. Mkrtchyan, H. Mkrtchyan, N. Morgan, K. E. Myers, A. Narayan, N. Nuruzzaman, A. K. Opper, SA Page1, J. Pan, K. Paschke, S. Phillips, M. Pitt, M. Poelker, W. D. Ramsay, M. Ramsey-Musolf, J. Roche, B. Sawatzky, T. Seva, R. Silwal, N. Simicevic, G. Smith2, T. Smith, P. Solvignon, P. Souder, D. Spayde, A. Subedi, R. Subedi, R. Suleiman, E. Tsentalovich, V. Tvaskis, W.T.H. van Oers, B. Waidyawansa, P. Wang, S. Wells, S. A. Wood, S. Yang, R. Young, S. Zhamkochyan
1Spokespersons *deceased 2Project Manager
Notes: 27