## Slide 1: Imaging methods for in-vivo Boron Neutron Capture Therapy
N. Ferrara, D. Ramos, G. Pugliese G. Iaselli, N. Amoroso, C. Gong, V. Pascali , S. Altieri, and N. Protti
nicola.ferrara@ba.infn.it

## Slide 2: Outline
BNCT introduction and dose measurement
Compton Camera Imaging
Imaging reconstruction with Maximum Likelihood Expectation Maximisation method
Innovative imaging method by using Deep Learning technic

## Slide 3: Boron neutron capture therapy
Boron Neutron Capture Therapy (BNCT) is an innovative hadrontherapy with high selectivity over cancer tissue based on the neutron capture reaction 10B(n, α)7Li
BNCT technique:
Tumour cells are loaded with 10B-enriched molecules and then irradiated with thermal neutrons
Neutron undergoes to capture reaction with boron-10
Energetic charged particles (α and 7Li*), released by reaction, produce damage to the tumour cells
The excited lithium nucleus decays emitting a prompt gamma ray of 478 keV

## Slide 4: Dosimetry within BNCT and project scope
Altieri S, Protti N. A brief review on reactor-based neutron sources for boron neutron capture therapy. Ther Radiol Oncol 2018;2:47.
In BNCT we have to take into account four separate radiation components:
Dtotal =
DB +
Dp +
Dn +
D𝛄
From 1H(n,𝛄)2H (E=2.2 MeV, 𝛔=0.33 b) & reactor background
Due to epithermal and fast neutron elastic scattering daily with H nuclei
From 14N(n,p)14C reactions (Ep=630 keV, 𝛔=1.9 b) due to thermal neutron
Therapeutic dose from 10B(n,⍺)7Li (𝛔=3837 b)
Nowadays dose estimation by blood test before, meanwhile and after irradiation
Project aims to develop an on-line dose estimation method by using Compton imaging
Dose distribution: dD(x,y,z) ≈ n10B(x,y,z) Φ(x,y,z) dV
The main contribution to the dose is coming from the Boron
Tumor-to-tissue boron concentration ratio (T/N) > 3
Notes: n is the boron concentration

## Slide 5: Compton method for imaging
Two sensitive layers are needed to evaluate the Compton angle
SCATTERER layer to promote the compton interaction of the incident gamma
ABSORBER layer to stop, by photoelectric effect, the scattered photon
Reconstruction is done by overlapping all cones, defined by the Compton angle, and interaction positions
(called back projection image)
Proposed detector is 3D CZT Compton camera
Room-temperature gamma-ray spectroscopic
Sub-millimetre spatial resolution and excellent energy resolution (<1% FWHM at 661.7 keV)
CZT single unit: 20x20x5 mm3, planar transversal field (PTF), orthogonal drift strip electrodes
Abbene l., et. al. High-resolution 3D CZT drift strip detector, J. Synchrotron Rad. (2020). 27, 1564-1576
Notes: CZT =

## Slide 6: Simulation layout description
Detector
CZT crystal stack (5mm thickness each), 60 mm apart from the source
Phantom material: Air or soft tissue
Simulated sources: 5-points like and spheric
Tomography: cube 120 mm side centered with source
Detector

## Slide 7: MLEM reconstruction method validation
Maximum Likelihood Expectation Maximisation (MLEM) → Iterative method to reconstruct the most probable source distribution
λnj = calculated amplitude of pixel j at the nth iteration
sj = sensitivity, i.e. the probability that a gamma ray originating from pixel j is detected anywhere
tij = imaging response matrix, i.e. the transition probabilities generated by the measured events(first estimation: based on back-projection, λ0)
Good resolution in x and y profiles, slighly worse in z direction
No image interference when material phantom is added
5-point like source
Notes: With only the
Back projection are not distinguished.
After 50 iterations …

## Slide 8: Boron ratio study
More realistic condition:
Spheric source to simulate the tumor
Two boron ratios: ideal case (T/N=5) and (T/N=2) extreme case (clinical values are T/N>3)
Both situations resoluted
More iterations needed to solve the image in z (≈250)
in Tissue phantom, T/N = 5.0
I = 250
I = 50
Back-projection
I = 250
I = 50
Back-projection
in Tissue phantom, T/N = 2.0

## Slide 9: Novel Deep Learning approach
Limitation: MLEM works only post-irradiation (not online), tomography takes few mins (≈ 20 min)
New method for an online reconstruction to go from the back-projection image to the tomography dose using ML procedure
Training Deep Learning model with back-projection and tomography image sets to make tomography reconstruction
Without availability of Compton images databases within BNCT, the approach was to identify the tumor by segmentation using Transfer Learning

## Slide 10: Deep learning model
Matlab pipe-line for segmentation of images
Use of Convolutional Neural Network model with a Residual U-net Architecture (ResNet), widely used for segmentation
Test metrics: accuracy and sensitivity
Resnet U-net architecture
Notes: We used a CNN model available in the Matlab
And we used as matrics (because a digital images)
True Positive and faulce negative

## Slide 11: Model performance results
Case 3 Spheric source in Tissue (T/N=5)
Case 4: 5 point-like source in air
Case 2: Spheric source in Tissue (T/N=2)
Case 1: Spheric source in Air
Normalized back projections
CNN segmentation

## Slide 12: Conclusions
                A Compton camera imaging approaches for dose tomography within Boron Neutron Capture Therapy was studied
                The iterative method (MLEM) allows the dose reconstruction, but it is not applicable in in-vivo therapy situations
                Deep Learning (DL) is promising as first approach for tumor monitoring with good accuracy/sensitivity compromise
                Next steps:
                Enrich Compton camera image database
                New DL model development for tomography reconstruction

## Slide 13: THANKS!
nicola.ferrara@ba.infn.it