## Slide 1
Motivations
n_TOF ideal facility for measurements of inelastic and (n,xn) reactions up to high energy.
Main applications: fission reactors, nuclear medicine, fusion reactors. Also important for refining theoretical models.
Technique
Measurements of (n,inel) and (n,xn) reactions typically performed by counting specific g-rays emitted in the deexcitation of the residue.
Need high efficiency and high resolution detectors -> ideally HPGe
Efforts have been going on at n_TOF since the beginning (2001) to operate HPGe in the EAR1.
However, the presence of the g-flash and the slow response of HPGe detectors, has hampered the use of these detectors so far, above a few hundred keV.
A fast switch circuit on the detector preamplifier has improved the situation, but the maximum neutron energy is still around a few MeV (too low considering the n_TOF potentialities).
Idea
LaBr3 crystal have high efficiency, reasonable resolution, and especially have fast signals.
They could allow to go much higher in energy then HPGe.
Using LaBr3 crystals for measurements of (n,inel) and (n,xn) reactions

## Slide 2
To investigate the use of LaBr3 crystals for measurements of inelastic and (n,xn) reactions, different tests have been performed and a LoI has been submitted (discussed in the INTC meeting few days ago).
First test in 2022: 1x1” crystal with PMT, both in EAR 1 and beam dump
Results
detectors recover fast from g-flash, but ringing in EAR1 prevented from getting a clear indication of the maximum energy that can be reached
very clean signal in beam dump, indication that hundreds of MeV neutron energy can be reached
resolution of 3% as expected
Second test in 2023: two larger crystals 1.5x1.5” and 1.5x2” with PMT, EAR1 only (capture position)
Test with all detectors at the end of this year campaign (Oct. 2023)
In the meanwhile:
used ENTE (Electronic Noise Testing Environment), setup by Nino Amato, to identify a possible source of ringing in the LaBr3
bad shielding of the Voltage Divider and connectors
fixed by wrapping the PMT+VD+connectors in an 26Al foil (laying on the floor …).
Using LaBr3 crystals for measurements of (n,inel) and (n,xn) reactions

## Slide 3
Three LaBr3 detectors looking at an Fe sample.
Capture position and supports.
Test performed by: Nikolas, Fran, Riccardo, Michele, Alice, Matt, Michi
Different geometries investigated:
3 angular positions (wrt beam): 135, 90, 45 degree
3 distances from the target: 11, 15 and 20 cm
Data collected over 2 days, mostly 20 minutes for each geometrical/sample configuration.
2023 test of LaBr3 crystals
Samples: Fe, Li, Au and empty

## Slide 4
The ringing problem has essentially disappeared !!
Clear (but not strong) dependence of the g-flash on the detector dimension.
Need a dedicated analysis for precise characterization of maximum energy reachable with the different sizes.
On-line plots and preliminary conclusions

## Slide 5
LOI: validation of the use of LaBr3 detectors for (n,inel) and (n,xn) reactions
We propose to use a combination of HPGe and LaBr3 detectors to investigate the possibility of measuring (n,inel) and (n,xn) reactions up to several tens of MeV (and even higher in some cases)
HPGe: high resolution, but limited neutron energy range (up to a few MeV, currently, due to g-flash response)
LaBr3: possible to go up to high neutron energy (up to hundreds of MeV), but with limited resolution.
A combination of the two might allow to extend the energy range well above current limits in data and evaluations.